Hemp Proficiency Testing

Summary All Labs Precision

Samples: HMNOV18-1 and HMNOV18-2

All Labs Precision

Issue Date:	1/4/2019

| Sam Num | Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Unit | Lab CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ∆9-THC |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| HM18NOV-1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 1 | ∆9-THC | 001.99 | Other | 114 |  |  |  | % |  |  |  |  |  | 1 |
| 1 | ∆9-THC | 001.10 | HPLC-UVVIS | 110 | <0.1600 | <0.1600 | <0.1600 | % |  |  |  |  |  | 1 |
| 1 | ∆9-THC | 001.10 | HPLC-UVVIS | 124 | .1780 | .1790 | .1800 | % | 0.56 | -1.35 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.10 | HPLC-UVVIS | 113 | 0.1278 | 0.1300 | 0.1279 | % | 0.97 | -1.25 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.10 | HPLC-UVVIS | 117 | 0.2900 | 0.2900 | 0.2800 | % | 2.01 | -0.99 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.10 | HPLC-UVVIS | 142 | 0.1284 | 0.1242 | 0.1227 | % | 2.36 | -0.9 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.10 | HPLC-UVVIS | 129 | 0.1072 | 0.1120 | 0.1118 | % | 2.46 | -0.88 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.10 | HPLC-UVVIS | 140 | 0.1436 | 0.1413 | 0.1488 | % | 2.66 | -0.83 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.10 | HPLC-UVVIS | 141 | 0.0971 | 0.1033 | 0.1028 | % | 3.41 | -0.65 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.99 | Other | 137 | 0.1270 | 0.1255 | 0.1343 | % | 3.65 | -0.59 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.99 | Other | 132 | .1756 | .1828 | .1889 | % | 3.65 | -0.59 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.99 | Other | 120 | 0.1607 | 0.1747 | 0.1683 | % | *4.17 | -0.46 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.99 | Other | 131 | 0.13 | 0.13 | 0.14 | % | *4.33 | -0.42 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.10 | HPLC-UVVIS | 101 | 0.1541 | 0.1415 | 0.1511 | % | 4.42 | -0.4 | 6.05 | 28 | 4.08 | 0 |

Population of Lab CVs

| Sam |  |  |  | Lab |  |  |  |  | Lab |  | Population of Lab CVs |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Num | Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Unit | CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
| 1 | ∆9-THC | 001.10 | HPLC-UVVIS | 111 | 0.1126 | 0.1177 | 0.1233 | % | 4.54 | -0.37 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.30 | HPLC-MS | 105 | 0.143 | 0.148 | 0.135 | % | *4.62 | -0.35 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.10 | HPLC-UVVIS | 108 | 0.1012 | 0.0929 | 0.1032 | % | 5.51 | -0.13 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.30 | HPLC-MS | 139 | 0.1662 | 0.1507 | 0.1491 | % | 6.08 | 0.01 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.30 | HPLC-MS | 119 | 0.132 | 0.1449 | 0.1505 | % | 6.66 | 0.15 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.10 | HPLC-UVVIS | 116 | 0.1300 | 0.1130 | 0.1240 | % | 7.05 | 0.25 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.10 | HPLC-UVVIS | 123 | 0.1502 | 0.1400 | 0.1301 | % | 7.17 | 0.28 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.30 | HPLC-MS | 102 | 0.1263 | 0.1459 | 0.1338 | % | 7.31 | 0.31 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.10 | HPLC-UVVIS | 109 | 0.1626 | 0.1499 | 0.1739 | % | *7.41 | 0.33 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.30 | HPLC-MS | 122 | 0.1307 | 0.1309 | 0.1511 | % | 8.52 | 0.61 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.10 | HPLC-UVVIS | 104 | 0.12 | 0.14 | 0.12 | % | *9.12 | 0.75 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.10 | HPLC-UVVIS | 125 | 0.1217 | 0.1549 | 0.1377 | % | 12 | 1.47 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.10 | HPLC-UVVIS | 118 | 0.121 | 0.115 | 0.089 | % | 15.7 | 2.37 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.10 | HPLC-UVVIS | 135 | 0.1528 | 0.1606 | 0.2118 | % | 18.3 | 3.01 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.10 | HPLC-UVVIS | 106 | 0.179 | 0.056 | 0.075 | % | 64.1 | 14.23 | 6.05 | 28 | 4.08 | 0 |
| 1 | ∆9-THC | 001.10 | HPLC-UVVIS | 136 | 0.01 | 0.02 | 0.06 | % | 88.2 | 20.14 | 6.05 | 28 | 4.08 | 0 |
| HM18NOV-2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 2 | ∆9-THC | 001.10 | HPLC-UVVIS | 124 | .2500 | .2460 | .2500 | % | 0.93 | -1.08 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.99 | Other | 137 | 0.2001 | 0.1939 | 0.1974 | % | 1.58 | -0.91 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.10 | HPLC-UVVIS | 116 | 0.1850 | 0.1910 | 0.1870 | % | 1.63 | -0.9 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.10 | HPLC-UVVIS | 117 | 0.3400 | 0.3500 | 0.3400 | % | 1.68 | -0.89 | 5.22 | 29 | 3.99 | 0 |

| Sam |  |  |  | Lab |  |  |  |  | Lab |  | Population of Lab CVs |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Num | Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Unit | CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
| 2 | ∆9-THC | 001.30 | HPLC-MS | 122 | 0.1769 | 0.1774 | 0.1830 | % | 1.89 | -0.84 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.10 | HPLC-UVVIS | 129 | 0.1614 | 0.1680 | 0.1641 | % | 2.02 | -0.8 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.99 | Other | 120 | 0.2301 | 0.2400 | 0.2372 | % | *2.16 | -0.77 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.99 | Other | 132 | .2432 | .2536 | .2522 | % | 2.26 | -0.74 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.10 | HPLC-UVVIS | 101 | 0.2177 | 0.2184 | 0.2275 | % | 2.47 | -0.69 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.10 | HPLC-UVVIS | 110 | 0.1810 | 0.1821 | 0.1732 | % | 2.71 | -0.63 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.10 | HPLC-UVVIS | 113 | 0.1914 | 0.1936 | 0.1826 | % | 3.08 | -0.54 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.10 | HPLC-UVVIS | 108 | 0.1611 | 0.1543 | 0.1507 | % | 3.4 | -0.46 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.10 | HPLC-UVVIS | 118 | 0.161 | 0.157 | 0.169 | % | 3.76 | -0.37 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.30 | HPLC-MS | 139 | 0.1918 | 0.2047 | 0.2055 | % | 3.83 | -0.35 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.10 | HPLC-UVVIS | 135 | 0.2338 | 0.2268 | 0.2163 | % | 3.9 | -0.33 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.10 | HPLC-UVVIS | 140 | 0.2194 | 0.2145 | 0.2029 | % | 3.99 | -0.31 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.30 | HPLC-MS | 105 | 0.186 | 0.177 | 0.192 | % | *4.08 | -0.29 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.10 | HPLC-UVVIS | 109 | 0.2288 | 0.2092 | 0.2203 | % | *4.48 | -0.19 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.10 | HPLC-UVVIS | 104 | 0.17 | 0.17 | 0.19 | % | *6.54 | 0.33 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.10 | HPLC-UVVIS | 123 | 0.2072 | 0.1851 | 0.2085 | % | 6.57 | 0.34 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.30 | HPLC-MS | 102 | 0.1935 | 0.2109 | 0.1840 | % | 6.96 | 0.43 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.99 | Other | 131 | 0.23 | 0.2 | 0.21 | % | *7.16 | 0.49 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.10 | HPLC-UVVIS | 111 | 0.1382 | 0.1464 | 0.1648 | % | 9.09 | 0.97 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.10 | HPLC-UVVIS | 141 | 0.1513 | 0.1841 | 0.1774 | % | 10.1 | 1.23 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.10 | HPLC-UVVIS | 125 | 0.1835 | 0.2255 | 0.2025 | % | 10.3 | 1.28 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.30 | HPLC-MS | 119 | 0.1813 | 0.2277 | 0.2195 | % | 11.8 | 1.66 | 5.22 | 29 | 3.99 | 0 |

| Sam Num | Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Unit | Lab CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 2 | ∆9-THC | 001.10 | HPLC-UVVIS | 142 | 0.2032 | 0.2121 | 0.1612 | % | 14.1 | 2.24 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.10 | HPLC-UVVIS | 136 | 0.08 | 0.12 | 0.09 | % | 21.5 | 4.09 | 5.22 | 29 | 3.99 | 0 |
| 2 | ∆9-THC | 001.10 | HPLC-UVVIS | 106 | 0.094 | 0.095 | 0.015 | % | 67.5 | 15.62 | 5.22 | 29 | 3.99 | 0 |

Population of Lab CVs

∆9-THCA

# HM18NOV-1

| 1 | ∆9-THCA | 002.10 | HPLC-UVVIS | 111 | <0.0050 | 0.0960 | <0.0050 | % |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | ∆9-THCA | 002.10 | HPLC-UVVIS | 110 | <0.1600 | <0.1600 | 0.2239 | % |  |  |  |  |  | 1 |
| 1 | ∆9-THCA | 002.30 | HPLC-MS | 122 | 0.1319 | 0.1312 | 0.1327 | % | 0.57 | -1.1 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.10 | HPLC-UVVIS | 101 | 0.1536 | 0.1552 | 0.1554 | % | 0.64 | -1.09 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.10 | HPLC-UVVIS | 113 | 0.0764 | 0.0776 | 0.0758 | % | 1.2 | -1.01 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.99 | Other | 132 | .0724 | .0747 | .0754 | % | 2.12 | -0.88 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.10 | HPLC-UVVIS | 140 | 0.0822 | 0.0805 | 0.0848 | % | 2.63 | -0.81 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.10 | HPLC-UVVIS | 136 | 0.21 | 0.2 | 0.21 | % | 2.79 | -0.79 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.10 | HPLC-UVVIS | 129 | 0.1063 | 0.1105 | 0.1038 | % | 3.17 | -0.73 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.99 | Other | 137 | 0.0820 | 0.0808 | 0.0859 | % | 3.22 | -0.73 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.10 | HPLC-UVVIS | 117 | 0.2700 | 0.2800 | 0.2900 | % | 3.57 | -0.68 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.30 | HPLC-MS | 105 | 0.0909 | 0.0931 | 0.0845 | % | *4.99 | -0.47 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.10 | HPLC-UVVIS | 109 | 0.1246 | 0.1366 | 0.1375 | % | *5.42 | -0.41 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.10 | HPLC-UVVIS | 135 | 0.1209 | 0.1236 | 0.1352 | % | 6 | -0.33 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.10 | HPLC-UVVIS | 123 | 0.1033 | 0.1085 | 0.0952 | % | 6.55 | -0.25 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.30 | HPLC-MS | 119 | 0.0961 | 0.1049 | 0.1095 | % | 6.58 | -0.25 | 8.35 | 26 | 7.07 | 0 |

| Sam |  |  |  | Lab |  |  |  |  | Lab |  | Population of Lab CVs |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Num | Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Unit | CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
| 1 | ∆9-THCA | 002.10 | HPLC-UVVIS | 116 | 0.1070 | 0.0950 | 0.1090 | % | 7.3 | -0.15 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.30 | HPLC-MS | 139 | 0.1260 | 0.1122 | 0.1087 | % | 7.91 | -0.06 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.99 | Other | 131 | 0.12 | 0.13 | 0.11 | % | *8.33 | 0 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.10 | HPLC-UVVIS | 108 | 0.1031 | 0.0860 | 0.0852 | % | 11.1 | 0.38 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.10 | HPLC-UVVIS | 124 | .0720 | .0920 | .0940 | % | 14.1 | 0.82 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.10 | HPLC-UVVIS | 142 | 0.0666 | 0.0537 | 0.0721 | % | 14.7 | 0.9 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.10 | HPLC-UVVIS | 141 | 0.0805 | 0.1043 | 0.0825 | % | 14.8 | 0.92 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.10 | HPLC-UVVIS | 106 | 0.101 | 0.129 | 0.135 | % | 14.9 | 0.93 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.30 | HPLC-MS | 102 | 0.1014 | 0.0831 | 0.0717 | % | 17.5 | 1.3 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.10 | HPLC-UVVIS | 118 | 0.089 | 0.091 | 0.061 | % | 20.9 | 1.77 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.10 | HPLC-UVVIS | 125 | 0.0922 | 0.1026 | 0.1399 | % | 22.5 | 2 | 8.35 | 26 | 7.07 | 0 |
| 1 | ∆9-THCA | 002.10 | HPLC-UVVIS | 104 | 0.07 | 0.12 | 0.13 | % | *30.1 | 3.08 | 8.35 | 26 | 7.07 | 0 |
| HM18NOV-2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 2 | ∆9-THCA | 002.10 | HPLC-UVVIS | 110 | 0.1700 | 0.1777 | <0.1600 | % |  |  |  |  |  | 1 |
| 2 | ∆9-THCA | 002.10 | HPLC-UVVIS | 117 | 0.3400 | 0.3400 | 0.3400 | % | 0 | -1.03 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.30 | HPLC-MS | 105 | 0.151 | 0.151 | 0.15 | % | *0.38 | -0.96 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.10 | HPLC-UVVIS | 124 | .2510 | .2500 | .2480 | % | 0.61 | -0.92 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.99 | Other | 132 | .1066 | .1049 | .1051 | % | 0.88 | -0.88 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.99 | Other | 137 | 0.1506 | 0.1465 | 0.1470 | % | 1.51 | -0.77 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.30 | HPLC-MS | 139 | 0.1726 | 0.1763 | 0.1782 | % | 1.62 | -0.75 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.10 | HPLC-UVVIS | 141 | 0.1585 | 0.1578 | 0.1527 | % | 2.03 | -0.68 | 6.07 | 27 | 5.92 | 0 |

| Sam |  |  |  | Lab |  |  |  |  | Lab |  | Population of Lab CVs |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Num | Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Unit | CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
| 2 | ∆9-THCA | 002.10 | HPLC-UVVIS | 136 | 0.25 | 0.26 | 0.25 | % | 2.28 | -0.64 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.99 | Other | 131 | 0.19 | 0.19 | 0.18 | % | *3.09 | -0.5 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.10 | HPLC-UVVIS | 113 | 0.1305 | 0.1314 | 0.1239 | % | 3.18 | -0.49 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.30 | HPLC-MS | 122 | 0.1880 | 0.2007 | 0.1976 | % | 3.39 | -0.45 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.10 | HPLC-UVVIS | 118 | 0.137 | 0.147 | 0.143 | % | 3.54 | -0.43 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.10 | HPLC-UVVIS | 109 | 0.2097 | 0.2 | 0.2169 | % | *4.06 | -0.34 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.10 | HPLC-UVVIS | 140 | 0.1442 | 0.1404 | 0.1329 | % | 4.13 | -0.33 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.10 | HPLC-UVVIS | 129 | 0.1806 | 0.1900 | 0.1741 | % | 4.4 | -0.28 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.10 | HPLC-UVVIS | 101 | 0.2163 | 0.2375 | 0.2268 | % | 4.67 | -0.24 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.10 | HPLC-UVVIS | 123 | 0.1914 | 0.1764 | 0.1768 | % | 4.71 | -0.23 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.10 | HPLC-UVVIS | 108 | 0.1546 | 0.1444 | 0.1359 | % | 6.46 | 0.07 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.10 | HPLC-UVVIS | 116 | 0.1600 | 0.1810 | 0.1770 | % | 6.46 | 0.07 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.10 | HPLC-UVVIS | 111 | 0.1239 | 0.1320 | 0.1440 | % | 7.59 | 0.26 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.30 | HPLC-MS | 119 | 0.149 | 0.1874 | 0.1821 | % | 12 | 1.01 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.10 | HPLC-UVVIS | 106 | 0.171 | 0.134 | 0.172 | % | 13.6 | 1.28 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.30 | HPLC-MS | 102 | 0.1684 | 0.1438 | 0.1283 | % | 13.8 | 1.3 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.10 | HPLC-UVVIS | 104 | 0.15 | 0.19 | 0.2 | % | *14.7 | 1.46 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.10 | HPLC-UVVIS | 135 | 0.2473 | 0.1754 | 0.198 | % | 17.8 | 1.98 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.10 | HPLC-UVVIS | 142 | 0.1152 | 0.0807 | 0.1148 | % | 19.1 | 2.21 | 6.07 | 27 | 5.92 | 0 |
| 2 | ∆9-THCA | 002.10 | HPLC-UVVIS | 125 | 0.1727 | 0.4202 | 0.1957 | % | 52 | 7.76 | 6.07 | 27 | 5.92 | 0 |

| Sam |  |  |  | Lab |  |  |  |  | Lab |  | Population of Lab CVs |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Num | Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Unit | CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
|  |  |  |  |  |  | CBD |  |  |  |  |  |  |  |  |
| HM18NOV-1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 1 | CBD | 003.10 | HPLC-UVVIS | 124 | 6.453 | 6.48 | 6.489 | % | 0.29 | -1.39 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.10 | HPLC-UVVIS | 117 | 1.5000 | 1.5100 | 1.5100 | % | 0.38 | -1.35 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.10 | HPLC-UVVIS | 113 | 1.4906 | 1.5415 | 1.5137 | % | 1.68 | -0.9 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.10 | HPLC-UVVIS | 141 | 1.4368 | 1.3848 | 1.4148 | % | 1.85 | -0.85 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.30 | HPLC-MS | 122 | 1.1843 | 1.1730 | 1.1405 | % | 1.95 | -0.81 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.99 | Other | 131 | 1.47 | 1.41 | 1.44 | % | *2.08 | -0.77 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.10 | HPLC-UVVIS | 135 | 1.4107 | 1.4405 | 1.4745 | % | 2.21 | -0.72 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.10 | HPLC-UVVIS | 108 | 1.3590 | 1.3480 | 1.298 | % | 2.44 | -0.64 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.10 | HPLC-UVVIS | 140 | 1.4920 | 1.4714 | 1.5448 | % | 2.52 | -0.61 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.10 | HPLC-UVVIS | 116 | 1.4120 | 1.3390 | 1.3930 | % | 2.74 | -0.54 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.10 | HPLC-UVVIS | 142 | 1.4218 | 1.3715 | 1.3417 | % | 2.94 | -0.47 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.10 | HPLC-UVVIS | 111 | 1.0993 | 1.1699 | 1.1371 | % | 3.11 | -0.41 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.99 | Other | 137 | 1.427 | 1.419 | 1.505 | % | 3.28 | -0.35 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.30 | HPLC-MS | 105 | 1.45 | 1.43 | 1.36 | % | *3.34 | -0.33 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.10 | HPLC-UVVIS | 110 | 1.2504 | 1.1709 | 1.2453 | % | 3.64 | -0.23 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.10 | HPLC-UVVIS | 136 | 1.28 | 1.38 | 1.4 | % | 4.75 | 0.16 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.10 | HPLC-UVVIS | 109 | 1.5929 | 1.5655 | 1.7107 | % | *4.75 | 0.16 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.99 | Other | 132 | 1.3839 | 1.4330 | 1.5210 | % | 4.8 | 0.18 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.10 | HPLC-UVVIS | 123 | 1.585 | 1.488 | 1.439 | % | 4.94 | 0.22 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.30 | HPLC-MS | 139 | 1.6106 | 1.5437 | 1.4496 | % | 5.27 | 0.34 | 4.3 | 28 | 2.89 | 0 |

| Sam |  |  |  | Lab |  |  |  |  | Lab |  | Population of Lab CVs |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Num | Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Unit | CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
| 1 | CBD | 003.10 | HPLC-UVVIS | 104 | 1.61 | 1.81 | 1.76 | % | *6.03 | 0.6 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.30 | HPLC-MS | 102 | 1.3268 | 1.3637 | 1.4913 | % | 6.19 | 0.66 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.10 | HPLC-UVVIS | 129 | 1.3423 | 1.4260 | 1.5374 | % | 6.82 | 0.87 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.10 | HPLC-UVVIS | 106 | 1.654 | 1.544 | 1.81 | % | 8.01 | 1.28 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.30 | HPLC-MS | 119 | 1.2165 | 1.4234 | 1.3969 | % | 8.37 | 1.41 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.10 | HPLC-UVVIS | 101 | 1.6495 | 1.8348 | 1.9631 | % | 8.68 | 1.52 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.10 | HPLC-UVVIS | 125 | 1.3470 | 1.7402 | 1.6676 | % | 13.2 | 3.08 | 4.3 | 28 | 2.89 | 0 |
| 1 | CBD | 003.10 | HPLC-UVVIS | 118 | 1.237 | 1.193 | 0.934 | % | 14.6 | 3.57 | 4.3 | 28 | 2.89 | 0 |
| HM18NOV-2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 2 | CBD | 003.10 | HPLC-UVVIS | 124 | 9.138 | 9.104 | 9.17 | % | 0.36 | -1.35 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.10 | HPLC-UVVIS | 106 | 2.124 | 2.127 | 2.152 | % | 0.72 | -1.19 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.10 | HPLC-UVVIS | 117 | 2.0200 | 1.9900 | 2.0200 | % | 0.86 | -1.13 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.30 | HPLC-MS | 122 | 1.5671 | 1.5654 | 1.5326 | % | 1.25 | -0.97 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.99 | Other | 137 | 2.095 | 2.038 | 2.075 | % | 1.4 | -0.9 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.99 | Other | 131 | 1.95 | 1.93 | 1.89 | % | *1.59 | -0.82 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.10 | HPLC-UVVIS | 113 | 2.0727 | 2.1350 | 2.0714 | % | 1.74 | -0.76 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.10 | HPLC-UVVIS | 118 | 1.647 | 1.589 | 1.637 | % | 1.91 | -0.69 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.10 | HPLC-UVVIS | 109 | 2.1291 | 2.1099 | 2.2024 | % | *2.27 | -0.53 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.30 | HPLC-MS | 139 | 1.8525 | 1.9283 | 1.9281 | % | 2.3 | -0.52 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.10 | HPLC-UVVIS | 141 | 1.9249 | 1.8388 | 1.8758 | % | 2.3 | -0.52 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.99 | Other | 132 | 1.7900 | 1.8404 | 1.8842 | % | 2.56 | -0.41 | 3.53 | 28 | 2.35 | 0 |

| Sam |  |  |  | Lab |  |  |  |  | Lab |  | Population of Lab CVs |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Num | Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Unit | CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
| 2 | CBD | 003.10 | HPLC-UVVIS | 108 | 1.7950 | 1.9040 | 1.8370 | % | 2.98 | -0.23 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.30 | HPLC-MS | 102 | 1.8338 | 1.8018 | 1.9167 | % | 3.2 | -0.14 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.10 | HPLC-UVVIS | 129 | 1.9154 | 2.0491 | 1.9942 | % | 3.38 | -0.06 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.10 | HPLC-UVVIS | 135 | 1.9365 | 2.0752 | 1.9886 | % | 3.5 | -0.01 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.10 | HPLC-UVVIS | 111 | 1.4031 | 1.4897 | 1.5043 | % | 3.73 | 0.09 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.10 | HPLC-UVVIS | 123 | 2.079 | 1.93 | 2.019 | % | 3.73 | 0.09 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.10 | HPLC-UVVIS | 104 | 2.28 | 2.16 | 2.36 | % | *4.44 | 0.39 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.10 | HPLC-UVVIS | 110 | 1.6190 | 1.6729 | 1.7752 | % | 4.7 | 0.5 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.10 | HPLC-UVVIS | 140 | 2.3068 | 2.2552 | 2.1049 | % | 4.72 | 0.51 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.10 | HPLC-UVVIS | 101 | 2.1886 | 2.2240 | 2.4117 | % | 5.27 | 0.74 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.10 | HPLC-UVVIS | 116 | 1.7230 | 1.8780 | 1.9350 | % | 5.95 | 1.03 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.30 | HPLC-MS | 105 | 1.8 | 1.93 | 2.04 | % | *6.25 | 1.16 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.10 | HPLC-UVVIS | 136 | 1.98 | 2.23 | 2.02 | % | 6.47 | 1.25 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.10 | HPLC-UVVIS | 125 | 1.9359 | 2.4297 | 2.3599 | % | 11.9 | 3.57 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.30 | HPLC-MS | 119 | 1.6341 | 2.0736 | 1.9667 | % | 12.1 | 3.65 | 3.53 | 28 | 2.35 | 0 |
| 2 | CBD | 003.10 | HPLC-UVVIS | 142 | 1.9459 | 2.2540 | 1.7372 | % | 13.1 | 4.08 | 3.53 | 28 | 2.35 | 0 |
|  |  |  |  |  |  | CBDA |  |  |  |  |  |  |  |  |
| HM18NOV-1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 1 | CBDA | 004.10 | HPLC-UVVIS | 117 | 4.4200 | 4.4400 | 4.4500 | % | 0.34 | -1.35 | 4.52 | 25 | 3.1 | 0 |
| 1 | CBDA | 004.10 | HPLC-UVVIS | 104 | 6.37 | 6.32 | 6.36 | % | *0.42 | -1.32 | 4.52 | 25 | 3.1 | 0 |
| 1 | CBDA | 004.10 | HPLC-UVVIS | 141 | 4.5822 | 4.6235 | 4.7103 | % | 1.41 | -1 | 4.52 | 25 | 3.1 | 0 |

| Sam |  |  |  | Lab |  |  |  |  | Lab |  | Population of Lab CVs |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Num | Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Unit | CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
| 1 | CBDA | 004.99 | Other | 131 | 4.63 | 4.55 | 4.48 | % | *1.65 | -0.93 | 4.52 | 25 | 3.1 | 0 |
| 1 | CBDA | 004.10 | HPLC-UVVIS | 116 | 3.7260 | 3.7470 | 3.8540 | % | 1.82 | -0.87 | 4.52 | 25 | 3.1 | 0 |
| 1 | CBDA | 004.10 | HPLC-UVVIS | 135 | 4.8818 | 5.0147 | 5.0614 | % | 1.87 | -0.85 | 4.52 | 25 | 3.1 | 0 |
| 1 | CBDA | 004.10 | HPLC-UVVIS | 108 | 4.0300 | 3.9200 | 3.8180 | % | 2.7 | -0.58 | 4.52 | 25 | 3.1 | 0 |
| 1 | CBDA | 004.10 | HPLC-UVVIS | 110 | 3.4948 | 3.3108 | 3.4432 | % | 2.78 | -0.56 | 4.52 | 25 | 3.1 | 0 |
| 1 | CBDA | 004.30 | HPLC-MS | 122 | 4.3595 | 4.4599 | 4.6314 | % | 3.07 | -0.47 | 4.52 | 25 | 3.1 | 0 |
| 1 | CBDA | 004.99 | Other | 132 | 3.2579 | 3.3077 | 3.4665 | % | 3.26 | -0.41 | 4.52 | 25 | 3.1 | 0 |
| 1 | CBDA | 004.99 | Other | 137 | 3.66 | 3.61 | 3.843 | % | 3.31 | -0.39 | 4.52 | 25 | 3.1 | 0 |
| 1 | CBDA | 004.10 | HPLC-UVVIS | 140 | 4.1332 | 4.0749 | 4.3498 | % | 3.46 | -0.34 | 4.52 | 25 | 3.1 | 0 |
| 1 | CBDA | 004.10 | HPLC-UVVIS | 129 | 3.9271 | 3.8492 | 4.1246 | % | 3.58 | -0.3 | 4.52 | 25 | 3.1 | 0 |
| 1 | CBDA | 004.10 | HPLC-UVVIS | 111 | 3.5724 | 3.9115 | 3.7033 | % | 4.59 | 0.02 | 4.52 | 25 | 3.1 | 0 |
| 1 | CBDA | 004.10 | HPLC-UVVIS | 125 | 4.1082 | 4.5228 | 4.4097 | % | 4.93 | 0.13 | 4.52 | 25 | 3.1 | 0 |
| 1 | CBDA | 004.10 | HPLC-UVVIS | 109 | 5.3161 | 5.3541 | 5.8714 | % | *5.63 | 0.36 | 4.52 | 25 | 3.1 | 0 |
| 1 | CBDA | 004.10 | HPLC-UVVIS | 101 | 4.8646 | 5.2082 | 5.4469 | % | 5.66 | 0.37 | 4.52 | 25 | 3.1 | 0 |
| 1 | CBDA | 004.10 | HPLC-UVVIS | 123 | 4.682 | 4.302 | 4.161 | % | 6.15 | 0.53 | 4.52 | 25 | 3.1 | 0 |
| 1 | CBDA | 004.10 | HPLC-UVVIS | 136 | 3.32 | 3.63 | 3.75 | % | 6.22 | 0.55 | 4.52 | 25 | 3.1 | 0 |
| 1 | CBDA | 004.10 | HPLC-UVVIS | 142 | 3.7674 | 3.3041 | 3.7112 | % | 7.03 | 0.81 | 4.52 | 25 | 3.1 | 0 |
| 1 | CBDA | 004.30 | HPLC-MS | 119 | 3.4915 | 4.033 | 3.9821 | % | 7.8 | 1.06 | 4.52 | 25 | 3.1 | 0 |
| 1 | CBDA | 004.10 | HPLC-UVVIS | 106 | 4.007 | 3.776 | 4.439 | % | 8.26 | 1.21 | 4.52 | 25 | 3.1 | 0 |
| 1 | CBDA | 004.30 | HPLC-MS | 139 | 3.7047 | 3.4483 | 3.1082 | % | 8.75 | 1.37 | 4.52 | 25 | 3.1 | 0 |
| 1 | CBDA | 004.30 | HPLC-MS | 105 | 4.15 | 4.72 | 3.98 | % | *9.05 | 1.46 | 4.52 | 25 | 3.1 | 0 |
| 1 | CBDA | 004.10 | HPLC-UVVIS | 118 | 4.08 | 3.978 | 3.029 | % | 15.7 | 3.6 | 4.52 | 25 | 3.1 | 0 |

| Sam |  |  |  | Lab |  |  |  |  | Lab |  | Population of Lab CVs |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Num | Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Unit | CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
| HM18NOV-2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 2 | CBDA | 004.10 | HPLC-UVVIS | 117 | 6.8100 | 6.7900 | 6.8500 | % | 0.45 | -1.26 | 2.81 | 25 | 1.87 | 0 |
| 2 | CBDA | 004.10 | HPLC-UVVIS | 106 | 5.863 | 5.846 | 5.933 | % | 0.78 | -1.08 | 2.81 | 25 | 1.87 | 0 |
| 2 | CBDA | 004.99 | Other | 132 | 5.0876 | 5.0157 | 5.0877 | % | 0.82 | -1.06 | 2.81 | 25 | 1.87 | 0 |
| 2 | CBDA | 004.10 | HPLC-UVVIS | 101 | 7.2105 | 7.1743 | 7.2989 | % | 0.89 | -1.03 | 2.81 | 25 | 1.87 | 0 |
| 2 | CBDA | 004.10 | HPLC-UVVIS | 141 | 6.9747 | 6.9645 | 7.0933 | % | 1.02 | -0.96 | 2.81 | 25 | 1.87 | 0 |
| 2 | CBDA | 004.30 | HPLC-MS | 122 | 5.9404 | 6.0149 | 6.1009 | % | 1.33 | -0.79 | 2.81 | 25 | 1.87 | 0 |
| 2 | CBDA | 004.99 | Other | 137 | 6.117 | 5.935 | 5.994 | % | 1.54 | -0.68 | 2.81 | 25 | 1.87 | 0 |
| 2 | CBDA | 004.99 | Other | 131 | 7.03 | 7 | 6.81 | % | *1.72 | -0.58 | 2.81 | 25 | 1.87 | 0 |
| 2 | CBDA | 004.10 | HPLC-UVVIS | 108 | 6.0190 | 6.2330 | 6.1090 | % | 1.76 | -0.56 | 2.81 | 25 | 1.87 | 0 |
| 2 | CBDA | 004.10 | HPLC-UVVIS | 110 | 5.2071 | 5.3446 | 5.4090 | % | 1.94 | -0.47 | 2.81 | 25 | 1.87 | 0 |
| 2 | CBDA | 004.10 | HPLC-UVVIS | 118 | 6 | 5.933 | 6.184 | % | 2.15 | -0.35 | 2.81 | 25 | 1.87 | 0 |
| 2 | CBDA | 004.10 | HPLC-UVVIS | 125 | 7.3829 | 7.0502 | 7.2299 | % | 2.31 | -0.27 | 2.81 | 25 | 1.87 | 0 |
| 2 | CBDA | 004.10 | HPLC-UVVIS | 109 | 8.1262 | 7.9283 | 8.3045 | % | *2.32 | -0.26 | 2.81 | 25 | 1.87 | 0 |
| 2 | CBDA | 004.10 | HPLC-UVVIS | 111 | 5.4668 | 5.7666 | 5.6759 | % | 2.73 | -0.04 | 2.81 | 25 | 1.87 | 0 |
| 2 | CBDA | 004.10 | HPLC-UVVIS | 129 | 6.3020 | 6.5855 | 6.2305 | % | 2.95 | 0.07 | 2.81 | 25 | 1.87 | 0 |
| 2 | CBDA | 004.10 | HPLC-UVVIS | 135 | 7.6264 | 8.1031 | 7.7563 | % | 3.15 | 0.18 | 2.81 | 25 | 1.87 | 0 |
| 2 | CBDA | 004.10 | HPLC-UVVIS | 142 | 5.9326 | 5.5948 | 5.5729 | % | 3.54 | 0.39 | 2.81 | 25 | 1.87 | 0 |
| 2 | CBDA | 004.10 | HPLC-UVVIS | 104 | 9.34 | 8.91 | 9.65 | % | *4 | 0.63 | 2.81 | 25 | 1.87 | 0 |
| 2 | CBDA | 004.10 | HPLC-UVVIS | 123 | 6.909 | 6.341 | 6.676 | % | 4.3 | 0.8 | 2.81 | 25 | 1.87 | 0 |
| 2 | CBDA | 004.30 | HPLC-MS | 105 | 6.85 | 7.43 | 6.97 | % | *4.32 | 0.81 | 2.81 | 25 | 1.87 | 0 |
| 2 | CBDA | 004.10 | HPLC-UVVIS | 116 | 5.7080 | 6.2640 | 5.9740 | % | 4.65 | 0.98 | 2.81 | 25 | 1.87 | 0 |

| Sam |  |  |  | Lab |  |  |  |  | Lab | Population of Lab CVs |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Num | Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Unit | CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
| 2 | CBDA | 004.10 | HPLC-UVVIS | 140 | 6.7856 | 6.6296 | 6.1449 | % | 5.12 | 1.24 | 2.81 | 25 | 1.87 | 0 |
| 2 | CBDA | 004.10 | HPLC-UVVIS | 136 | 5.87 | 6.44 | 5.91 | % | 5.24 | 1.3 | 2.81 | 25 | 1.87 | 0 |
| 2 | CBDA | 004.30 | HPLC-MS | 139 | 4.8867 | 5.2704 | 4.6687 | % | 6.16 | 1.79 | 2.81 | 25 | 1.87 | 0 |
| 2 | CBDA | 004.30 | HPLC-MS | 119 | 5.294 | 6.6423 | 6.4688 | % | 12 | 4.89 | 2.81 | 25 | 1.87 | 0 |
|  |  |  |  |  |  | CBN |  |  |  |  |  |  |  |  |
| HM18NOV-1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 1 | CBN | 005.10 | HPLC-UVVIS | 125 | <0.002 | <0.002 | <0.002 | % |  |  |  |  |  | 1 |
| 1 | CBN | 005.10 | HPLC-UVVIS | 111 | <0.0050 | <0.0050 | <0.0050 | % |  |  |  |  |  | 1 |
| 1 | CBN | 005.10 | HPLC-UVVIS | 142 | <0.0067 | <0.0067 | <0.0067 | % |  |  |  |  |  | 1 |
| 1 | CBN | 005.10 | HPLC-UVVIS | 140 | <0.01 | <0.01 | <0.01 | % |  |  |  |  |  | 1 |
| 1 | CBN | 005.10 | HPLC-UVVIS | 109 | <0.0100 | <0.0100 | <0.0100 | % |  |  |  |  |  | 1 |
| 1 | CBN | 005.99 | Other | 132 | <0.0100 | <0.0100 | <0.0100 | % |  |  |  |  |  | 1 |
| 1 | CBN | 005.10 | HPLC-UVVIS | 141 | <0.012 | <0.012 | <0.012 | % |  |  |  |  |  | 1 |
| 1 | CBN | 005.20 | HPLC-MS | 102 | <0.014 | <0.014 | <0.014 | % |  |  |  |  |  | 1 |
| 1 | CBN | 005.20 | HPLC-MS | 105 | <0.0144 | <0.0144 | <0.0144 | % |  |  |  |  |  | 1 |
| 1 | CBN | 005.10 | HPLC-UVVIS | 118 | <0.02 | <0.02 | <0.02 | % |  |  |  |  |  | 1 |
| 1 | CBN | 005.10 | HPLC-UVVIS | 104 | <0.02 | <0.02 | <0.02 | % |  |  |  |  |  | 1 |
| 1 | CBN | 005.30 | GC-FID | 112 | <0.0200 | <0.0200 | <0.0200 | % |  |  |  |  |  | 1 |
| 1 | CBN | 005.99 | Other | 137 | <0.0200 | <0.0200 | <0.0200 | % |  |  |  |  |  | 1 |
| 1 | CBN | 005.99 | Other | 131 | <0.03 | <0.03 | <0.03 | % |  |  |  |  |  | 1 |
| 1 | CBN | 005.10 | HPLC-UVVIS | 108 | <0.03 | <0.03 | <0.03 | % |  |  |  |  |  | 1 |

| Sam Num | Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Unit | Lab CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | CBN | 005.10 | HPLC-UVVIS | 123 | <0.0448 | <0.0453 | <0.0476 | % |  |  |  |  |  | 1 |
| 1 | CBN | 005.10 | HPLC-UVVIS | 117 | <0.0600 | <0.0600 | <0.0600 | % |  |  |  |  |  | 1 |
| 1 | CBN | 005.40 | GC-MS | 127 | <0.1500 | <0.1500 | <0.1500 | % |  |  |  |  |  | 1 |
| 1 | CBN | 005.10 | HPLC-UVVIS | 110 | <0.1600 | <0.1600 | <0.1600 | % |  |  |  |  |  | 1 |
| 1 | CBN | 005.10 | HPLC-UVVIS | 135 | 0 | 0 | 0 | % |  |  |  |  |  | 2 |
| 1 | CBN | 005.10 | HPLC-UVVIS | 116 | 0.0290 | <0.0250 | 0.0400 | % |  |  |  |  |  | 1 |
| 1 | CBN | 005.30 | GC-FID | 116 | 0.0385 | <0.0300 | 0.0354 | % |  |  |  |  |  | 1 |
| 1 | CBN | 005.30 | GC-FID | 106 | <0.062 | 0.0647 | <0.062 | % |  |  |  |  |  | 1 |
| 1 | CBN | 005.10 | HPLC-UVVIS | 101 | 0.0476 | 0.0485 | 0.0485 | % | 1.08 | -0.97 | 7.83 | 12 | 6.98 | 0 |
| 1 | CBN | 005.40 | GC-MS | 143 | 0.0179 | 0.0183 | 0.0182 | % | 1.15 | -0.96 | 7.83 | 12 | 6.98 | 0 |
| 1 | CBN | 005.30 | GC-FID | 107 | 0.0952 | 0.0922 | 0.0909 | % | 2.38 | -0.78 | 7.83 | 12 | 6.98 | 0 |
| 1 | CBN | 005.10 | HPLC-UVVIS | 136 | 0.22 | 0.22 | 0.23 | % | 2.59 | -0.75 | 7.83 | 12 | 6.98 | 0 |
| 1 | CBN | 005.20 | HPLC-MS | 122 | 0.0029 | 0.0029 | 0.0031 | % | 3.89 | -0.56 | 7.83 | 12 | 6.98 | 0 |
| 1 | CBN | 005.10 | HPLC-UVVIS | 106 | 0.109 | 0.1 | 0.111 | % | 5.49 | -0.33 | 7.83 | 12 | 6.98 | 0 |
| 1 | CBN | 005.10 | HPLC-UVVIS | 124 | .0100 | .0110 | .0120 | % | 9.09 | 0.18 | 7.83 | 12 | 6.98 | 0 |
| 1 | CBN | 005.30 | GC-FID | 114 | 0.1178 | 0.1040 | 0.0982 | % | 9.44 | 0.23 | 7.83 | 12 | 6.98 | 0 |
| 1 | CBN | 005.30 | GC-FID | 139 | 0.1186 | 0.1019 | 0.0991 | % | 9.9 | 0.3 | 7.83 | 12 | 6.98 | 0 |
| 1 | CBN | 005.10 | HPLC-UVVIS | 129 | 0.0012 | 0.0015 | 0.0015 | % | 12.4 | 0.65 | 7.83 | 12 | 6.98 | 0 |
| 1 | CBN | 005.30 | GC-FID | 138 | 0.0115 | 0.0081 | 0.0078 | % | 22.5 | 2.1 | 7.83 | 12 | 6.98 | 0 |
| 1 | CBN | 005.20 | HPLC-MS | 139 | 0.0019 | 0.0025 | 0.0011 | % | 38.3 | 4.37 | 7.83 | 12 | 6.98 | 0 |

Population of Lab CVs

| Sam |  |  |  | Lab |  |  |  |  | Lab | Population of Lab CVs |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Num | Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Unit | CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
| HM18NOV-2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 2 | CBN | 005.10 | HPLC-UVVIS | 125 | <0.002 | <0.002 | <0.002 | % |  |  |  |  |  | 1 |
| 2 | CBN | 005.10 | HPLC-UVVIS | 111 | <0.0050 | <0.0050 | <0.0050 | % |  |  |  |  |  | 1 |
| 2 | CBN | 005.10 | HPLC-UVVIS | 142 | <0.0067 | <0.0067 | <0.0067 | % |  |  |  |  |  | 1 |
| 2 | CBN | 005.10 | HPLC-UVVIS | 140 | <0.01 | <0.01 | <0.01 | % |  |  |  |  |  | 1 |
| 2 | CBN | 005.10 | HPLC-UVVIS | 141 | <0.012 | <0.012 | <0.012 | % |  |  |  |  |  | 1 |
| 2 | CBN | 005.20 | HPLC-MS | 102 | <0.014 | <0.014 | <0.014 | % |  |  |  |  |  | 1 |
| 2 | CBN | 005.20 | HPLC-MS | 105 | <0.0144 | <0.0144 | <0.0144 | % |  |  |  |  |  | 1 |
| 2 | CBN | 005.10 | HPLC-UVVIS | 118 | <0.02 | <0.02 | <0.02 | % |  |  |  |  |  | 1 |
| 2 | CBN | 005.10 | HPLC-UVVIS | 104 | <0.02 | <0.02 | <0.02 | % |  |  |  |  |  | 1 |
| 2 | CBN | 005.30 | GC-FID | 112 | <0.0200 | <0.0200 | <0.0200 | % |  |  |  |  |  | 1 |
| 2 | CBN | 005.99 | Other | 137 | <0.0200 | <0.0200 | <0.0200 | % |  |  |  |  |  | 1 |
| 2 | CBN | 005.99 | Other | 131 | <0.03 | <0.03 | <0.03 | % |  |  |  |  |  | 1 |
| 2 | CBN | 005.10 | HPLC-UVVIS | 108 | <0.03 | <0.03 | <0.03 | % |  |  |  |  |  | 1 |
| 2 | CBN | 005.10 | HPLC-UVVIS | 123 | <0.0442 | <0.0471 | <0.0476 | % |  |  |  |  |  | 1 |
| 2 | CBN | 005.10 | HPLC-UVVIS | 117 | <0.0600 | <0.0600 | <0.0600 | % |  |  |  |  |  | 1 |
| 2 | CBN | 005.40 | GC-MS | 127 | <0.1500 | <0.1500 | <0.1500 | % |  |  |  |  |  | 1 |
| 2 | CBN | 005.10 | HPLC-UVVIS | 110 | <0.1600 | <0.1600 | <0.1600 | % |  |  |  |  |  | 1 |
| 2 | CBN | 005.10 | HPLC-UVVIS | 135 | 0 | 0 | 0 | % |  |  |  |  |  | 2 |
| 2 | CBN | 005.10 | HPLC-UVVIS | 109 | 0.01 | <0.0100 | 0.0097 | % |  |  |  |  |  | 1 |
| 2 | CBN | 005.10 | HPLC-UVVIS | 116 | <0.0250 | <0.0250 | 0.0320 | % |  |  |  |  |  | 1 |
| 2 | CBN | 005.30 | GC-FID | 116 | 0.0400 | <0.0300 | 0.0363 | % |  |  |  |  |  | 1 |

| Sam Num | Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Unit | Lab CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 2 | CBN | 005.30 | GC-FID | 106 | <0.062 | 0.067 | 0.069 | % |  |  |  |  |  | 1 |
| 2 | CBN | 005.10 | HPLC-UVVIS | 124 | .0160 | .0160 | .0160 | % | 0 | -1.06 | 8.05 | 13 | 7.59 | 0 |
| 2 | CBN | 005.10 | HPLC-UVVIS | 106 | 0.165 | 0.165 | 0.163 | % | 0.7 | -0.97 | 8.05 | 13 | 7.59 | 0 |
| 2 | CBN | 005.40 | GC-MS | 143 | 0.0331 | 0.0333 | 0.0357 | % | 4.25 | -0.5 | 8.05 | 13 | 7.59 | 0 |
| 2 | CBN | 005.20 | HPLC-MS | 122 | 0.0026 | 0.0024 | 0.0026 | % | 4.56 | -0.46 | 8.05 | 13 | 7.59 | 0 |
| 2 | CBN | 005.10 | HPLC-UVVIS | 129 | 0.0026 | 0.0029 | 0.0028 | % | 5.52 | -0.33 | 8.05 | 13 | 7.59 | 0 |
| 2 | CBN | 005.10 | HPLC-UVVIS | 101 | 0.0657 | 0.0734 | 0.0697 | % | 5.53 | -0.33 | 8.05 | 13 | 7.59 | 0 |
| 2 | CBN | 005.30 | GC-FID | 107 | 0.1317 | 0.1477 | 0.1365 | % | 5.92 | -0.28 | 8.05 | 13 | 7.59 | 0 |
| 2 | CBN | 005.10 | HPLC-UVVIS | 136 | 0.22 | 0.25 | 0.23 | % | 6.55 | -0.2 | 8.05 | 13 | 7.59 | 0 |
| 2 | CBN | 005.30 | GC-FID | 139 | 0.1336 | 0.1390 | 0.1524 | % | 6.83 | -0.16 | 8.05 | 13 | 7.59 | 0 |
| 2 | CBN | 005.30 | GC-FID | 114 | 0.1611 | 0.1403 | 0.1458 | % | 7.23 | -0.11 | 8.05 | 13 | 7.59 | 0 |
| 2 | CBN | 005.20 | HPLC-MS | 139 | 0.0034 | 0.0024 | 0.0026 | % | 18.9 | 1.43 | 8.05 | 13 | 7.59 | 0 |
| 2 | CBN | 005.30 | GC-FID | 138 | 0.0197 | 0.0128 | 0.0173 | % | 21.1 | 1.72 | 8.05 | 13 | 7.59 | 0 |
| 2 | CBN | 005.99 | Other | 132 | .0241 | .0548 | .0242 | % | 51.5 | 5.72 | 8.05 | 13 | 7.59 | 0 |

Population of Lab CVs

Total ∆9-THC

# HM18NOV-1

| 1 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 110 | <0.1600 | <0.1600 | 0.1964 | % |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 101 | 0.2888 | 0.2777 | 0.2874 | % | 2.12 | -1.04 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 129 | 0.2004 | 0.2089 | 0.2028 | % | 2.15 | -1.03 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 117 | 0.5200 | 0.5400 | 0.5400 | % | 2.17 | -1.03 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.99 | Other | 131 | 0.23 | 0.24 | 0.24 | % | *2.44 | -0.98 | 7.88 | 37 | 5.56 | 0 |

| Sam |  |  |  | Lab |  |  |  |  | Lab |  | Population of Lab CVs |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Num | Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Unit | CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
| 1 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 140 | 0.2157 | 0.2119 | 0.2231 | % | 2.63 | -0.95 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.40 | GC-FID | 138 | 0.1805 | 0.1805 | 0.1721 | % | 2.73 | -0.93 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.40 | GC-FID | 139 | 0.2367 | 0.2485 | 0.2367 | % | 2.83 | -0.91 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.40 | GC-FID | 116 | 0.2498 | 0.2415 | 0.2555 | % | 2.83 | -0.91 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.40 | GC-FID | 143 | 0.1817 | 0.1765 | 0.1711 | % | 3 | -0.88 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.99 | Other | 137 | 0.1989 | 0.1970 | 0.2090 | % | 3.2 | -0.84 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.99 | Other | 132 | 0.2391 | 0.2483 | 0.2550 | % | 3.23 | -0.84 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.30 | HPLC-MS | 122 | 0.2626 | 0.2621 | 0.2838 | % | 4.6 | -0.59 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 142 | 0.1868 | 0.1713 | 0.1859 | % | 4.8 | -0.55 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.30 | HPLC-MS | 105 | 0.223 | 0.23 | 0.209 | % | *4.85 | -0.55 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 109 | 0.2718 | 0.2697 | 0.2946 | % | *4.96 | -0.53 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.30 | HPLC-MS | 102 | 0.2152 | 0.2188 | 0.1967 | % | 5.64 | -0.4 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 123 | 0.2408 | 0.2351 | 0.2136 | % | 6.24 | -0.29 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 108 | 0.1917 | 0.1683 | 0.1779 | % | 6.56 | -0.24 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.30 | HPLC-MS | 119 | 0.2163 | 0.237 | 0.2465 | % | 6.62 | -0.23 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.30 | HPLC-MS | 139 | 0.2766 | 0.2491 | 0.2444 | % | 6.78 | -0.2 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 116 | 0.2238 | 0.1963 | 0.2196 | % | 6.95 | -0.17 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 141 | 0.1677 | 0.1948 | 0.1751 | % | 7.82 | -0.01 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.40 | GC-FID | 112 | 0.2487 | 0.2243 | 0.2130 | % | 7.98 | 0.02 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.40 | GC-FID | 114 | 0.2005 | 0.1798 | 0.1702 | % | 8.44 | 0.1 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.50 | GC-MS | 128 | 0.2542 | 0.2207 | 0.2136 | % | 9.45 | 0.28 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.50 | GC-MS | 127 | 0.2672 | 0.2218 | 0.2185 | % | *11.5 | 0.66 | 7.88 | 37 | 5.56 | 0 |

| Sam |  |  |  | Lab |  |  |  |  | Lab |  | Population of Lab CVs |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Num | Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Unit | CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
| 1 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 125 | 0.2026 | 0.2449 | 0.2604 | % | 12.7 | 0.86 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 124 | 0.2607 | 0.2704 | 0.3297 | % | 13 | 0.92 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 136 | .1942 | .1954 | .2442 | % | 13.5 | 1.01 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 135 | 0.2589 | 0.2689 | 0.3305 | % | 13.6 | 1.02 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.40 | GC-FID | 107 | 0.2095 | 0.2396 | 0.2755 | % | 13.7 | 1.04 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 104 | 0.19 | 0.25 | 0.23 | % | *13.7 | 1.04 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.40 | GC-FID | 121 | 0.1990 | 0.2054 | 0.2554 | % | 14 | 1.11 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 118 | 0.199 | 0.195 | 0.142 | % | 17.8 | 1.79 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 106 | 0.267 | 0.169 | 0.193 | % | 24.4 | 2.97 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 111 | 0.1126 | 0.2019 | 0.1233 | % | 33.4 | 4.6 | 7.88 | 37 | 5.56 | 0 |
| 1 | Total ∆9-THC | 006.40 | GC-FID | 106 | 0.097 | 0.162 | 0.082 | % | 37.4 | 5.31 | 7.88 | 37 | 5.56 | 0 |
| HM18NOV-2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 2 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 117 | 0.6400 | 0.6400 | 0.6400 | % | 0 | -1.33 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.40 | GC-FID | 121 | 0.3593 | 0.3638 | 0.3644 | % | 0.77 | -1.14 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 124 | 0.4701 | 0.4731 | 0.4798 | % | 1.05 | -1.07 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.99 | Other | 132 | 0.3366 | .3455 | .3444 | % | 1.42 | -0.98 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.99 | Other | 137 | 0.3322 | 0.3230 | 0.3264 | % | 1.42 | -0.98 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.40 | GC-FID | 116 | 0.3881 | 0.4012 | 0.3938 | % | 1.67 | -0.92 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.40 | GC-FID | 143 | 0.2452 | 0.2398 | 0.2479 | % | 1.69 | -0.91 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.30 | HPLC-MS | 122 | 0.3649 | 0.3781 | 0.3806 | % | 2.25 | -0.77 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 118 | 0.281 | 0.286 | 0.294 | % | 2.28 | -0.76 | 5.35 | 38 | 4.03 | 0 |

| Sam |  |  |  | Lab |  |  |  |  | Lab |  | Population of Lab CVs |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Num | Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Unit | CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
| 2 | Total ∆9-THC | 006.50 | GC-MS | 128 | 0.3278 | 0.3130 | 0.3187 | % | 2.33 | -0.75 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.30 | HPLC-MS | 105 | 0.318 | 0.309 | 0.324 | % | *2.38 | -0.74 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 101 | 0.4074 | 0.4267 | 0.4264 | % | 2.63 | -0.68 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.30 | HPLC-MS | 139 | 0.3431 | 0.3594 | 0.3617 | % | 2.86 | -0.62 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 129 | 0.3198 | 0.3347 | 0.3167 | % | 2.97 | -0.59 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.40 | GC-FID | 138 | 0.2414 | 0.2399 | 0.2285 | % | 2.98 | -0.59 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 116 | 0.3253 | 0.3497 | 0.3422 | % | 3.69 | -0.41 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 109 | 0.4127 | 0.3846 | 0.4105 | % | *3.88 | -0.37 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 140 | 0.3458 | 0.3377 | 0.3194 | % | 4.05 | -0.32 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 108 | 0.2967 | 0.2810 | 0.2699 | % | 4.77 | -0.15 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 123 | 0.3751 | 0.3398 | 0.3635 | % | 5.01 | -0.09 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 141 | 0.2903 | 0.3225 | 0.3113 | % | 5.31 | -0.01 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.99 | Other | 131 | 0.4 | 0.36 | 0.37 | % | *5.53 | 0.04 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.40 | GC-FID | 107 | 0.3429 | 0.3494 | 0.3836 | % | 6.1 | 0.18 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.40 | GC-FID | 112 | 0.3330 | 0.3039 | 0.2973 | % | 6.1 | 0.19 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.40 | GC-FID | 139 | 0.3324 | 0.3736 | 0.3658 | % | 6.13 | 0.19 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.40 | GC-FID | 114 | 0.2836 | 0.2448 | 0.2633 | % | 7.35 | 0.5 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 142 | 0.3042 | 0.2829 | 0.2619 | % | 7.47 | 0.53 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 104 | 0.31 | 0.33 | 0.36 | % | *7.55 | 0.54 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.30 | HPLC-MS | 102 | 0.3412 | 0.337 | 0.2965 | % | 7.6 | 0.56 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 136 | .2993 | .3480 | .3093 | % | 8.07 | 0.67 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 111 | 0.2468 | 0.2622 | 0.2910 | % | 8.41 | 0.76 | 5.35 | 38 | 4.03 | 0 |

| Sam Num | Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Unit | Lab CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 2 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 135 | 0.4507 | 0.3806 | 0.3899 | % | 9.35 | 0.99 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.30 | HPLC-MS | 119 | 0.312 | 0.3921 | 0.3792 | % | 11.9 | 1.63 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.50 | GC-MS | 127 | 0.3791 | 0.2992 | 0.2955 | % | *14.6 | 2.28 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 106 | 0.244 | 0.213 | 0.166 | % | 18.9 | 3.36 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.40 | GC-FID | 106 | 0.119 | 0.199 | 0.216 | % | 29.1 | 5.89 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 125 | 0.3350 | 0.5940 | 0.3741 | % | 32.1 | 6.65 | 5.35 | 38 | 4.03 | 0 |
| 2 | Total ∆9-THC | 006.10 | HPLC-UVVIS | 110 | 0.3500 | 0.3368 | 0.1821 | % | 32.2 | 6.67 | 5.35 | 38 | 4.03 | 0 |

Population of Lab CVs

Total CBD

# HM18NOV-1

| 1 | Total CBD | 007.40 | GC-FID | 114 | >2.0 | >2.0 | >2.0 | % |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Total CBD | 007.10 | HPLC-UVVIS | 117 | 5.3700 | 5.4100 | 5.4100 | % | 0.43 | -1.36 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.10 | HPLC-UVVIS | 141 | 5.4553 | 5.4396 | 5.5458 | % | 1.05 | -1.18 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.10 | HPLC-UVVIS | 104 | 7.19 | 7.35 | 7.33 | % | *1.2 | -1.13 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.10 | HPLC-UVVIS | 108 | 0.5013 | 0.5111 | 0.5156 | % | 1.44 | -1.06 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.99 | Other | 131 | 5.53 | 5.4 | 5.37 | % | *1.57 | -1.02 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.10 | HPLC-UVVIS | 116 | 4.6797 | 4.6251 | 4.7730 | % | 1.59 | -1.01 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.40 | GC-FID | 143 | 4.9405 | 4.8872 | 4.7879 | % | 1.59 | -1.01 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.10 | HPLC-UVVIS | 135 | 5.692 | 5.8384 | 5.9134 | % | 1.94 | -0.91 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.30 | HPLC-MS | 122 | 5.5438 | 5.6328 | 5.7719 | % | 2.03 | -0.88 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.40 | GC-FID | 139 | 5.9683 | 6.2819 | 6.0264 | % | 2.74 | -0.67 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.10 | HPLC-UVVIS | 110 | 4.3154 | 4.0745 | 4.2650 | % | 3.01 | -0.59 | 4.99 | 34 | 3.36 | 0 |

| Sam |  |  |  | Lab |  |  |  |  | Lab |  | Population of Lab CVs |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Num | Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Unit | CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
| 1 | Total CBD | 007.10 | HPLC-UVVIS | 140 | 5.1169 | 5.0451 | 5.3596 | % | 3.19 | -0.54 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.99 | Other | 137 | 4.6377 | 4.5858 | 4.8762 | % | 3.3 | -0.51 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.99 | Other | 132 | 4.2411 | 4.3339 | 4.5611 | % | 3.76 | -0.37 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.10 | HPLC-UVVIS | 111 | 4.2323 | 4.6003 | 4.3849 | % | 4.2 | -0.24 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.10 | HPLC-UVVIS | 129 | 4.7864 | 4.8018 | 5.1547 | % | 4.24 | -0.22 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.40 | GC-FID | 138 | 4.9379 | 5.2608 | 5.3843 | % | 4.44 | -0.17 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.10 | HPLC-UVVIS | 142 | 4.7258 | 4.2692 | 4.5964 | % | 5.19 | 0.06 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.10 | HPLC-UVVIS | 109 | 6.2551 | 6.2611 | 6.8599 | % | *5.38 | 0.11 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.10 | HPLC-UVVIS | 136 | 4.1916 | 4.5635 | 4.6887 | % | 5.77 | 0.23 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.10 | HPLC-UVVIS | 123 | 5.691 | 5.26 | 5.088 | % | 5.81 | 0.24 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.50 | GC-MS | 128 | 3.0377 | 3.2106 | 3.4402 | % | 6.25 | 0.37 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.10 | HPLC-UVVIS | 101 | 5.9157 | 6.4024 | 6.7400 | % | 6.52 | 0.46 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.30 | HPLC-MS | 105 | 5.09 | 5.57 | 4.85 | % | *7.09 | 0.62 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.40 | GC-FID | 112 | 6.2169 | 5.6611 | 5.3899 | % | 7.32 | 0.69 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.10 | HPLC-UVVIS | 125 | 4.9499 | 5.7067 | 5.5349 | % | 7.35 | 0.7 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.30 | HPLC-MS | 139 | 4.8595 | 4.5678 | 4.1754 | % | 7.57 | 0.77 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.40 | GC-FID | 116 | 4.3120 | 3.7734 | 3.8085 | % | 7.6 | 0.78 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.30 | HPLC-MS | 119 | 4.2785 | 4.9603 | 4.8892 | % | 7.96 | 0.88 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.10 | HPLC-UVVIS | 106 | 5.168 | 4.856 | 5.703 | % | 8.17 | 0.95 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.50 | GC-MS | 127 | 5.9837 | 4.8977 | 4.5720 | % | *14.3 | 2.79 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.10 | HPLC-UVVIS | 118 | 4.815 | 4.681 | 3.59 | % | 15.4 | 3.1 | 4.99 | 34 | 3.36 | 0 |
| 1 | Total CBD | 007.40 | GC-FID | 107 | 3.7958 | 5.1665 | 4.9484 | % | 15.9 | 3.24 | 4.99 | 34 | 3.36 | 0 |

| Sam |  |  |  | Lab |  |  |  |  | Lab |  | Population of Lab CVs |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Num | Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Unit | CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
| 1 | Total CBD | 007.40 | GC-FID | 106 | 1.045 | 2.936 | 0.661 | % | 78.7 | 21.95 | 4.99 | 34 | 3.36 | 0 |
| HM18NOV-2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 2 | Total CBD | 007.40 | GC-FID | 114 | >2.0 | >2.0 | >2.0 | % |  |  |  |  |  | 1 |
| 2 | Total CBD | 007.10 | HPLC-UVVIS | 117 | 7.9900 | 7.9500 | 8.0300 | % | 0.5 | -1.35 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.10 | HPLC-UVVIS | 106 | 7.266 | 7.254 | 7.355 | % | 0.76 | -1.24 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.30 | HPLC-MS | 122 | 7.5075 | 7.5803 | 7.6335 | % | 0.84 | -1.2 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.99 | Other | 132 | 6.2518 | 6.2392 | 6.3461 | % | 0.93 | -1.16 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.10 | HPLC-UVVIS | 141 | 8.0417 | 7.9467 | 8.0966 | % | 0.94 | -1.15 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.99 | Other | 137 | 7.4700 | 7.2444 | 7.3332 | % | 1.55 | -0.88 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.99 | Other | 131 | 8.12 | 8.07 | 7.86 | % | *1.72 | -0.8 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.10 | HPLC-UVVIS | 125 | 8.4107 | 8.6127 | 8.7005 | % | 1.73 | -0.8 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.40 | GC-FID | 138 | 7.1812 | 7.4170 | 7.4325 | % | 1.92 | -0.71 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.10 | HPLC-UVVIS | 118 | 6.909 | 6.793 | 7.061 | % | 1.94 | -0.7 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.10 | HPLC-UVVIS | 101 | 8.5122 | 8.516 | 8.813 | % | 2 | -0.67 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.10 | HPLC-UVVIS | 109 | 9.2557 | 9.063 | 9.4854 | % | *2.28 | -0.55 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.40 | GC-FID | 143 | 7.0505 | 6.8599 | 7.2217 | % | 2.57 | -0.42 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.10 | HPLC-UVVIS | 110 | 6.1856 | 6.3601 | 6.5189 | % | 2.62 | -0.39 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.10 | HPLC-UVVIS | 142 | 7.1488 | 7.1606 | 6.8088 | % | 2.84 | -0.3 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.10 | HPLC-UVVIS | 129 | 7.4423 | 7.8245 | 7.4583 | % | 2.85 | -0.29 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.10 | HPLC-UVVIS | 111 | 6.1975 | 6.5470 | 6.4820 | % | 2.9 | -0.27 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.10 | HPLC-UVVIS | 135 | 8.6249 | 9.1816 | 8.7909 | % | 3.22 | -0.12 | 3.49 | 34 | 2.21 | 0 |

| Sam |  |  |  | Lab |  |  |  |  | Lab |  | Population of Lab CVs |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Num | Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Unit | CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
| 2 | Total CBD | 007.30 | HPLC-MS | 105 | 7.81 | 8.45 | 8.15 | % | *3.94 | 0.2 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.10 | HPLC-UVVIS | 104 | 10.47 | 9.97 | 10.83 | % | *4.14 | 0.3 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.10 | HPLC-UVVIS | 123 | 8.138 | 7.491 | 7.873 | % | 4.15 | 0.3 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.40 | GC-FID | 139 | 8.4382 | 9.1890 | 8.7798 | % | 4.27 | 0.35 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.30 | HPLC-MS | 139 | 6.1382 | 6.5505 | 6.0226 | % | 4.45 | 0.43 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.10 | HPLC-UVVIS | 116 | 6.7289 | 7.3715 | 7.1742 | % | 4.64 | 0.52 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.10 | HPLC-UVVIS | 108 | 0.5067 | 0.5092 | 0.5510 | % | 4.76 | 0.58 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.10 | HPLC-UVVIS | 140 | 8.2578 | 8.0693 | 7.4940 | % | 5.01 | 0.69 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.40 | GC-FID | 112 | 8.2627 | 7.5912 | 7.5323 | % | 5.21 | 0.78 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.40 | GC-FID | 116 | 6.2504 | 6.2529 | 5.6867 | % | 5.38 | 0.85 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.10 | HPLC-UVVIS | 136 | 7.1280 | 7.888 | 7.203 | % | 5.65 | 0.98 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.50 | GC-MS | 128 | 3.9170 | 3.5947 | 3.4778 | % | 6.21 | 1.23 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.40 | GC-FID | 107 | 6.9785 | 7.1266 | 7.8431 | % | 6.32 | 1.28 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.50 | GC-MS | 127 | 8.9365 | 7.4967 | 8.0810 | % | *8.86 | 2.43 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.30 | HPLC-MS | 119 | 6.2769 | 7.8989 | 7.6398 | % | 12 | 3.84 | 3.49 | 34 | 2.21 | 0 |
| 2 | Total CBD | 007.40 | GC-FID | 106 | 1.745 | 4 | 4.337 | % | 41.9 | 17.4 | 3.49 | 34 | 2.21 | 0 |
|  |  |  |  |  |  | Moisture |  |  |  |  |  |  |  |  |
| HM18NOV-1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 1 | Moisture | 500.10 | Oven | 104 | 12.67 | 12.67 | 12.67 | % | 0 | -1.08 | 3.81 | 7 | 3.51 | 0 |
| 1 | Moisture | 500.10 | Oven | 117 | 7.7600 | 7.86 | 7.7700 | % | 0.71 | -0.88 | 3.81 | 7 | 3.51 | 0 |
| 1 | Moisture | 500.99 | Other | 131 | 8.81 | 9.36 | 9.17 | % | 3.07 | -0.21 | 3.81 | 7 | 3.51 | 0 |

| Sam Num | Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Unit | Lab CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Moisture | 500.10 | Oven | 110 | 8.2700 | 8.8000 | 8.4300 | % | 3.2 | -0.17 | 3.81 | 7 | 3.51 | 0 |
| 1 | Moisture | 500.10 | Oven | 129 | 9.35 | 9.92 | 10.33 | % | 4.99 | 0.34 | 3.81 | 7 | 3.51 | 0 |
| 1 | Moisture | 500.10 | Oven | 109 | 8.71 | 7.78 | 8.28 | % | 5.64 | 0.52 | 3.81 | 7 | 3.51 | 0 |
| 1 | Moisture | 500.10 | Oven | 127 | 6.5467 | 11.7092 | 8.3984 | % | 29.4 | 7.3 | 3.81 | 7 | 3.51 | 0 |
| 1 | Moisture | 500.10 | Oven | 105 | 11.8 |  |  | % |  |  |  |  |  | 1 |
| 1 | Moisture | 500.99 | Other | 132 |  |  | 4.73 | % |  |  |  |  |  | 1 |
| 1 | Moisture | 500.10 | Oven | 120 | 8.809 | 8.508 |  | % |  |  |  |  |  | 1 |
| 1 | Moisture | 500.10 | Oven | 118 | 8.73 | N/A | N/A | % |  |  |  |  |  | 1 |
| 1 | Moisture | 500.10 | Oven | 139 | 9.3757 | 9.6007 |  | % |  |  |  |  |  | 1 |
| HM18NOV-2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 2 | Moisture | 500.99 | Other | 141 |  |  |  | % |  |  |  |  |  | 1 |
| 2 | Moisture | 500.99 | Other | 132 |  |  | 1.57 | % |  |  |  |  |  | 1 |
| 2 | Moisture | 500.10 | Oven | 104 | 12.58 | 12.58 | 12.58 | % | 0 | -1.25 | 2.86 | 7 | 2.29 | 0 |
| 2 | Moisture | 500.10 | Oven | 117 | 7.4200 | 7.3200 | 7.3800 | % | 0.68 | -0.95 | 2.86 | 7 | 2.29 | 0 |
| 2 | Moisture | 500.10 | Oven | 110 | 6.9200 | 7.2100 | 7.2200 | % | 2.39 | -0.2 | 2.86 | 7 | 2.29 | 0 |
| 2 | Moisture | 500.99 | Other | 131 | 8.48 | 8.94 | 8.73 | % | 2.64 | -0.09 | 2.86 | 7 | 2.29 | 0 |
| 2 | Moisture | 500.10 | Oven | 129 | 9.72 | 10.61 | 10.23 | % | 4.38 | 0.67 | 2.86 | 7 | 2.29 | 0 |
| 2 | Moisture | 500.10 | Oven | 109 | 7.21 | 7.89 | 7.48 | % | 4.55 | 0.74 | 2.86 | 7 | 2.29 | 0 |
| 2 | Moisture | 500.10 | Oven | 127 | 7.1497 | 7.9017 | 7.7985 | % | 5.35 | 1.09 | 2.86 | 7 | 2.29 | 0 |
| 2 | Moisture | 500.10 | Oven | 105 | 11.8 |  |  | % |  |  |  |  |  | 1 |
| 2 | Moisture | 500.10 | Oven | 120 | 9.256 | 9.18 |  | % |  |  |  |  |  | 1 |

Population of Lab CVs

| Sam Num | Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Unit | Lab CV,% | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 2 | Moisture | 500.10 | Oven | 139 | 9.6390 | 9.4636 |  | % |  |  |  |  |  | 1 |
| 2 | Moisture | 500.10 | Oven | 118 | N/A | N/A | N/A | % |  |  |  |  |  | 1 |

Population of Lab CVs

Flag definitions: 1 (Lab CV not calculated because there were less than 3 numeric lab results), 2 (Lab CV not calculated because there were less than 3 numeric lab results greater than 0).

Asterick on Lab CV denotes results were reported on an oven-dried basis. No asterick on Lab CV denotes results were reported on as-received basis.