Hemp Proficiency Testing

All Labs Trueness HM20NOV-2

Sample: HM20NOV-2

All Labs Trueness

Issue Date: December 17, 2020

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 |  | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Δ9-THC (%AR) |  |  |  |  |  |  |  |  |  |  |  |  |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 192 | 0.0409 | 0.0474 | 0.0493 | 0.0459 |  | -2.84 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 207 | 0.058 | 0.049 | 0.054 | 0.0537 |  | -2.22 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 188 | 0.0623 | 0.067 | 0.059 | 0.0628 |  | -1.5 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 141 | 0.0625 | 0.0664 | 0.0612 | 0.0634 |  | -1.45 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 191 | 0.0663 | 0.0639 | 0.0688 | 0.0663 |  | -1.21 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 161 | 0.0697 | 0.0611 | 0.0668 | 0.0659 |  | -1.25 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 103 | 0.0668 | 0.0688 | 0.0643 | 0.0666 |  | -1.19 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 193 | 0.07 | 0.07 | 0.06 | 0.0667 |  | -1.19 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 125 | 0.0693 | 0.0686 | 0.0651 | 0.0677 |  | -1.11 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 102 | 0.0721 | 0.0703 | 0.0654 | 0.0693 |  | -0.98 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 124 | 0.0767 | 0.0715 | 0.0661 | 0.0714 |  | -0.81 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 203 | 0.073 | 0.072 |  | 0.0725 |  | -0.72 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 148 | 0.073 | 0.076 | 0.074 | 0.0743 |  | -0.58 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 149 | 0.074 | 0.078 | 0.071 | 0.0743 |  | -0.58 | 0.0816 | 55 | 0.0126 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 113 | 0.0758 | 0.0745 | 0.0745 | 0.0749 | -0.53 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 132 | 0.081 | 0.077 | 0.068 | 0.0753 | -0.5 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 108 | 0.0829 | 0.0657 | 0.0774 | 0.0753 | -0.5 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 122 | 0.0773 | 0.0737 | 0.0739 | 0.075 | -0.53 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 197 | <0.0761 | 0.0753 | <0.0746 | 0.0753 |  |  |  |  | 1 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 186 | 0.0785 | 0.0712 | 0.0769 | 0.0755 | -0.48 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 182 | 0.0723 | 0.0865 | 0.0726 | 0.0771 | -0.36 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 175 | 0.0757 | 0.0782 | 0.0772 | 0.077 | -0.36 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 120 | 0.0794 | 0.0752 | 0.0796 | 0.0781 | -0.28 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 168 | 0.0797 | 0.0776 | 0.0773 | 0.0782 | -0.27 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 198 | 0.0801 | 0.0791 | 0.0766 | 0.0786 | -0.24 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 200 | 0.0797 | 0.0802 | 0.0787 | 0.0795 | -0.16 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 199 | 0.079 | 0.086 | 0.0755 | 0.0802 | -0.11 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 109 | 0.075 | 0.073 | 0.095 | 0.081 | -0.05 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 202 | 0.0858 | 0.0797 | 0.08 | 0.0818 | 0.02 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 144 | 0.0781 | 0.0787 | 0.0895 | 0.0821 | 0.04 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.99 | Other | 137 | 0.0799 | 0.0848 | 0.0816 | 0.0821 | 0.04 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 106 | 0.087 | 0.081 | 0.081 | 0.083 | 0.11 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 164 | 0.0835 | 0.0832 | 0.0818 | 0.0828 | 0.1 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 149 | 0.082 | 0.091 | 0.079 | 0.084 | 0.19 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.03 | AOAC 2018.11, M | 139 | 0.0896 | 0.0832 | 0.0804 | 0.0844 | 0.22 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 142 | 0.0867 | 0.0815 | 0.0874 | 0.0852 | 0.29 | 0.0816 | 55 | 0.0126 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 178 | 0.086 |  |  | 0.086 |  |  |  |  | 1 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 155 | 0.093 | 0.087 | 0.077 | 0.0857 | 0.32 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 101 | 0.097 | 0.073 | 0.092 | 0.0873 | 0.46 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.99 | Other | 132 | 0.09 | 0.08 | 0.09 | 0.0867 | 0.4 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 150 | 0.088 | 0.085 | 0.09 | 0.0877 | 0.48 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 165 | 0.0864 | 0.0814 | 0.0964 | 0.0881 | 0.51 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 184 | 0.0903 | 0.0805 | 0.0922 | 0.0877 | 0.48 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 172 | 0.094 | 0.09 | 0.089 | 0.091 | 0.75 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 189 | 0.091 | 0.091 | 0.091 | 0.091 | 0.75 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 123 | 0.088 | 0.095 | <0.0890 | 0.0907 | 0.72 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 119 | 0.0951 | 0.0893 | 0.0937 | 0.0927 | 0.88 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 204 | 0.0962 | 0.0975 | 0.0944 | 0.096 | 1.15 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 152 | 0.0994 | 0.0943 | 0.0942 | 0.096 | 1.14 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 115 | 0.095 | 0.093 | 0.099 | 0.0957 | 1.12 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 116 | 0.0978 | 0.095 | 0.0969 | 0.0966 | 1.19 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.99 | Other | 172 | 0.1045 | 0.09 | 0.0989 | 0.0978 | 1.29 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 162 | 0.0983 | 0.0988 | 0.0966 | 0.0979 | 1.3 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 196 | 0.1049 | 0.1063 | 0.098 | 0.103 | 1.71 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 163 | 0.106 | 0.109 | 0.11 | 0.108 | 2.13 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.60 | NIR | 187 | 0.1406 | 0.1437 | 0.15 | 0.145 | 5.02 | 0.0816 | 55 | 0.0126 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 145 | 0.5 | 0.55 | 0.48 | 0.51 | 34.07 | 0.0816 | 55 | 0.0126 | 0 |

Population of Lab Values

| Analyte | Code |  | Method | Lab Num | Result1 | Result2 | Result3 |  | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Δ9-THCA (%AR) |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Δ9-THCA (%AR) |  | 002.01 | AOAC 2018.10 | 193 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV | 197 | <0.0761 | <0.0753 | <0.0746 | <0.0761 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV | 194 | <0.0800 | <0.0800 | <0.0800 | <0.0800 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV | 142 | 0.028 | 0.0256 | 0.0252 | 0.0263 |  | -2.09 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) |  | 002.99 | Other | 132 | 0.02 | 0.03 | 0.03 | 0.0267 |  | -2.05 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV | 132 | 0.025 | 0.031 | 0.028 | 0.028 |  | -1.94 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV | 141 | 0.0286 | 0.0281 | 0.0278 | 0.0282 |  | -1.93 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV | 124 | 0.0302 | 0.0304 | 0.0262 | 0.0289 |  | -1.87 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) |  | 002.60 | NIR | 187 | 0.0371 | 0.038 | 0.0294 | 0.0348 |  | -1.37 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) |  | 002.02 | AOAC 2018.11, U | 103 | 0.037 | 0.0304 | 0.039 | 0.0355 |  | -1.32 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV | 207 | 0.037 | 0.037 | 0.039 | 0.0377 |  | -1.13 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) |  | 002.02 | AOAC 2018.11, U | 199 | 0.0395 | 0.0405 | 0.0375 | 0.0392 |  | -1.01 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV | 186 | 0.0422 | 0.0417 | 0.0419 | 0.0419 |  | -0.78 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV | 184 | 0.0436 | 0.0422 | 0.0444 | 0.0434 |  | -0.65 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) |  | 002.01 | AOAC 2018.10 | 101 | 0.046 | 0.038 | 0.048 | 0.044 |  | -0.6 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV | 178 | 0.044 |  |  | 0.044 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) |  | 002.01 | AOAC 2018.10 | 115 | 0.045 | 0.045 | 0.046 | 0.0453 |  | -0.49 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV | 198 | 0.0501 | 0.0434 | 0.0432 | 0.0456 |  | -0.47 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV | 202 | 0.0437 | 0.0488 | 0.0451 | 0.0459 |  | -0.45 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV | 200 | 0.0487 | 0.0475 | 0.0457 | 0.0473 |  | -0.33 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV | 162 | 0.0512 | 0.047 | 0.0461 | 0.0481 |  | -0.26 | 0.0512 | 51 | 0.012 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 175 | 0.0508 | 0.0466 | 0.0483 | 0.0486 | -0.22 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 204 | 0.0523 | 0.0473 | 0.047 | 0.0489 | -0.2 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.99 | Other | 137 | 0.0495 | 0.051 | 0.0474 | 0.0493 | -0.16 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 161 | 0.0501 | 0.0483 | 0.0488 | 0.0491 | -0.18 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 203 | 0.049 | 0.048 |  | 0.0485 | -0.23 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 149 | 0.048 | 0.051 | 0.047 | 0.0487 | -0.21 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 120 | 0.0509 | 0.0514 | 0.0514 | 0.0512 | 0 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 155 | 0.055 | 0.059 | 0.043 | 0.0523 | 0.09 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 144 | 0.0497 | 0.0523 | 0.0553 | 0.0524 | 0.1 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 113 | 0.0518 | 0.0516 | 0.0523 | 0.0519 | 0.06 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 106 | 0.053 | 0.052 | 0.051 | 0.052 | 0.06 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.03 | AOAC 2018.11, M | 139 | 0.0568 | 0.0574 | 0.0436 | 0.0526 | 0.11 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 149 | 0.052 | 0.055 | 0.053 | 0.0533 | 0.18 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 152 | 0.0557 | 0.0509 | 0.0554 | 0.054 | 0.23 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 122 | 0.0552 | 0.0538 | 0.0535 | 0.0542 | 0.25 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 196 | 0.0562 | 0.0558 | 0.0506 | 0.0542 | 0.25 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 188 | 0.0534 | 0.0556 | 0.0526 | 0.0539 | 0.22 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 191 | 0.0561 | 0.0523 | 0.0552 | 0.0545 | 0.28 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 172 | 0.058 | 0.054 | 0.052 | 0.0547 | 0.29 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 150 | 0.065 | 0.05 | 0.053 | 0.056 | 0.4 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 119 | 0.0588 | 0.056 | 0.0563 | 0.057 | 0.49 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 116 | 0.0587 | 0.052 | 0.0593 | 0.0567 | 0.45 | 0.0512 | 51 | 0.012 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 182 | 0.056 | 0.0625 | 0.0552 | 0.0579 | 0.56 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.99 | Other | 172 | 0.0576 | 0.0609 | 0.0565 | 0.0583 | 0.59 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 165 | 0.0666 | 0.0555 | 0.0607 | 0.0609 | 0.81 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 108 | 0.0638 | 0.063 | 0.0604 | 0.0624 | 0.93 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 109 | 0.062 | 0.067 | 0.061 | 0.0633 | 1.01 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 125 | 0.0647 | 0.063 | 0.0745 | 0.0674 | 1.35 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 168 | 0.0698 | 0.0726 | 0.0604 | 0.0676 | 1.37 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 164 | 0.0651 | 0.0733 | 0.0689 | 0.0691 | 1.49 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 163 | 0.069 | 0.07 | 0.071 | 0.07 | 1.57 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 192 | 0.0813 | 0.0714 | 0.0741 | 0.0756 | 2.04 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 189 | 0.083 | 0.084 | 0.084 | 0.0837 | 2.71 | 0.0512 | 51 | 0.012 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 123 | <0.0880 | <0.0910 | 0.106 | 0.095 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 145 | 0.54 | 0.55 | 0.57 | 0.553 | 42 | 0.0512 | 51 | 0.012 | 0 |
|  |  |  |  | CBD (%AR) |  |  |  |  |  |  |  |  |
| CBD (%AR) | 003.01 | AOAC 2018.10 | 193 | 1 | 0.99 | 0.98 | 0.99 | -3.1 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.01 | AOAC 2018.10 | 155 | 1.161 | 1.185 | 1.137 | 1.16 | -1.97 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 122 | 1.2105 | 1.2023 | 1.1765 | 1.2 | -1.74 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 132 | 1.278 | 1.169 | 1.256 | 1.23 | -1.49 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 141 | 1.2351 | 1.2315 | 1.3157 | 1.26 | -1.31 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 202 | 1.2587 | 1.2275 | 1.3163 | 1.27 | -1.27 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 207 | 1.29 | 1.282 | 1.315 | 1.3 | -1.08 | 1.4603 | 46 | 0.1519 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%AR) | 003.10 | Other LC-UV | 150 | 1.333 | 1.3 | 1.333 | 1.32 | -0.91 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 199 | 1.2615 | 1.391 | 1.2935 | 1.32 | -0.95 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.03 | AOAC 2018.11, M | 139 | 1.4727 | 1.3065 | 1.2758 | 1.35 | -0.72 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 161 | 1.3583 | 1.3104 | 1.387 | 1.35 | -0.71 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.60 | NIR | 187 | 1.3781 | 1.3415 | 1.3983 | 1.37 | -0.58 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.99 | Other | 132 | 1.36 | 1.32 | 1.42 | 1.37 | -0.62 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 200 | 1.3294 | 1.3601 | 1.4438 | 1.38 | -0.54 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 203 | 1.391 | 1.377 |  | 1.38 | -0.5 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 184 | 1.4857 | 1.4046 | 1.3148 | 1.4 | -0.39 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 178 | 1.413 |  |  | 1.41 |  |  |  |  | 1 |
| CBD (%AR) | 003.10 | Other LC-UV | 186 | 1.4165 | 1.3982 | 1.4206 | 1.41 | -0.32 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 108 | 1.43 | 1.437 | 1.427 | 1.43 | -0.19 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 149 | 1.353 | 1.558 | 1.377 | 1.43 | -0.2 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 149 | 1.395 | 1.452 | 1.497 | 1.45 | -0.08 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 125 | 1.4781 | 1.4658 | 1.3941 | 1.45 | -0.09 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 191 | 1.4331 | 1.4483 | 1.4591 | 1.45 | -0.09 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 189 | 1.454 | 1.467 | 1.449 | 1.46 | -0.02 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.99 | Other | 137 | 1.44 | 1.51 | 1.44 | 1.46 | 0.02 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 113 | 1.4675 | 1.4556 | 1.4565 | 1.46 | 0 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 196 | 1.4836 | 1.4903 | 1.4515 | 1.48 | 0.1 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.01 | AOAC 2018.10 | 152 | 1.5493 | 1.472 | 1.4564 | 1.49 | 0.21 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 204 | 1.4918 | 1.4935 | 1.4937 | 1.49 | 0.22 | 1.4603 | 46 | 0.1519 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%AR) | 003.10 | Other LC-UV | 197 | 1.537 | 1.446 | 1.547 | 1.51 | 0.33 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 142 | 1.5242 | 1.465 | 1.5697 | 1.52 | 0.39 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.01 | AOAC 2018.10 | 168 | 1.5444 | 1.472 | 1.5603 | 1.53 | 0.43 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 164 | 1.5241 | 1.5673 | 1.5231 | 1.54 | 0.51 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 198 | 1.5899 | 1.5193 | 1.4981 | 1.54 | 0.5 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 172 | 1.611 | 1.567 | 1.503 | 1.56 | 0.66 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 144 | 1.5596 | 1.557 | 1.549 | 1.56 | 0.62 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 162 | 1.5754 | 1.5791 | 1.5869 | 1.58 | 0.79 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 192 | 1.7007 | 1.5378 | 1.5281 | 1.59 | 0.85 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 123 | 1.59 | 1.615 | 1.564 | 1.59 | 0.85 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 116 | 1.5826 | 1.597 | 1.6333 | 1.6 | 0.95 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 109 | 1.635 | 1.646 | 1.523 | 1.6 | 0.93 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 188 | 1.6876 | 1.6555 | 1.5396 | 1.63 | 1.1 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 106 | 1.648 | 1.658 | 1.659 | 1.65 | 1.28 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 163 | 1.666 | 1.673 | 1.658 | 1.67 | 1.35 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 119 | 1.7378 | 1.6243 | 1.7101 | 1.69 | 1.52 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.99 | Other | 172 | 1.7902 | 1.7413 | 1.6702 | 1.73 | 1.8 | 1.4603 | 46 | 0.1519 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 145 | 15.61 | 15.3 | 15.14 | 15.4 | 91.45 | 1.4603 | 46 | 0.1519 | 0 |
|  |  |  |  | CBDA (%AR) |  |  |  |  |  |  |  |  |
| CBDA (%AR) | 004.10 | Other LC-UV | 132 | 2.096 | 1.946 | 2.075 | 2.04 | -2.93 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.99 | Other | 132 | 2.23 | 2.06 | 2.16 | 2.15 | -2.65 | 3.2232 | 46 | 0.4047 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDA (%AR) | 004.10 | Other LC-UV | 142 | 2.3597 | 2.3571 | 2.4413 | 2.39 | -2.07 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 150 | 2.667 | 2.633 | 2.667 | 2.66 | -1.4 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 122 | 2.7282 | 2.6628 | 2.6506 | 2.68 | -1.34 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.01 | AOAC 2018.10 | 155 | 2.783 | 2.822 | 2.736 | 2.78 | -1.09 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 149 | 2.46 | 2.915 | 3.06 | 2.81 | -1.02 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 202 | 2.806 | 2.7212 | 2.9067 | 2.81 | -1.02 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 207 | 2.738 | 2.756 | 2.922 | 2.81 | -1.03 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 161 | 2.7637 | 2.7509 | 2.9115 | 2.81 | -1.02 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 199 | 2.6445 | 3.032 | 2.8225 | 2.83 | -0.96 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.03 | AOAC 2018.11, M | 139 | 3.1473 | 2.7214 | 2.7615 | 2.88 | -0.86 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 108 | 3.042 | 2.904 | 2.845 | 2.93 | -0.72 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 123 | 3.003 | 3.008 | 2.971 | 2.99 | -0.57 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 198 | 3.3162 | 2.9941 | 2.9696 | 3.09 | -0.32 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 189 | 3.139 | 3.145 | 3.128 | 3.14 | -0.21 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 192 | 3.2288 | 3.1409 | 3.0685 | 3.15 | -0.19 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 152 | 3.1935 | 3.1362 | 3.1553 | 3.16 | -0.15 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 141 | 3.2458 | 3.2937 | 3.0362 | 3.19 | -0.08 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 119 | 3.35 | 3.1169 | 3.1252 | 3.2 | -0.06 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 125 | 3.387 | 3.1706 | 3.0653 | 3.21 | -0.04 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.01 | AOAC 2018.10 | 193 | 3.2 | 3.27 | 3.2 | 3.22 | 0 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 178 | 3.226 |  |  | 3.23 |  |  |  |  | 1 |
| CBDA (%AR) | 004.10 | Other LC-UV | 116 | 3.207 | 3.2175 | 3.3296 | 3.25 | 0.07 | 3.2232 | 46 | 0.4047 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDA (%AR) | 004.10 | Other LC-UV | 149 | 3.09 | 3.576 | 3.142 | 3.27 | 0.11 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 186 | 3.3024 | 3.2907 | 3.3313 | 3.31 | 0.21 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 144 | 3.1197 | 3.2427 | 3.5965 | 3.32 | 0.24 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 196 | 3.359 | 3.378 | 3.296 | 3.34 | 0.3 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 204 | 3.3638 | 3.3478 | 3.3581 | 3.36 | 0.33 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 200 | 3.3313 | 3.3714 | 3.3971 | 3.37 | 0.35 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 184 | 3.3473 | 3.3333 | 3.473 | 3.38 | 0.4 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 106 | 3.389 | 3.412 | 3.409 | 3.4 | 0.45 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.01 | AOAC 2018.10 | 168 | 3.4733 | 3.3339 | 3.3986 | 3.4 | 0.44 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 188 | 3.3956 | 3.475 | 3.4025 | 3.42 | 0.5 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 203 | 3.476 | 3.388 |  | 3.43 | 0.52 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.99 | Other | 137 | 3.43 | 3.6 | 3.36 | 3.46 | 0.59 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 162 | 3.4483 | 3.4095 | 3.5939 | 3.48 | 0.64 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.60 | NIR | 187 | 3.5625 | 3.5741 | 3.3894 | 3.51 | 0.71 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 197 | 3.392 | 3.64 | 3.553 | 3.53 | 0.75 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 172 | 3.717 | 3.577 | 3.393 | 3.56 | 0.84 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 191 | 3.4969 | 3.5627 | 3.6348 | 3.56 | 0.84 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 163 | 3.707 | 3.699 | 3.702 | 3.7 | 1.18 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 164 | 3.7497 | 3.7327 | 3.6256 | 3.7 | 1.18 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 194 | 3.693 | 3.746 | 3.971 | 3.8 | 1.43 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.99 | Other | 172 | 4.1305 | 3.9749 | 3.7704 | 3.96 | 1.82 | 3.2232 | 46 | 0.4047 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 109 | 4.237 | 4.299 | 3.927 | 4.15 | 2.3 | 3.2232 | 46 | 0.4047 | 0 |

Population of Lab Values

Lab

Lab

Population of Lab Values

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDA (%AR) | 004.10 | Other LC-UV | 145 | 36.12 | 36.63 | 37.04 | 36.6 | 82.47 | 3.2232 | 46 | 0.4047 | 0 |

|  |  |  |  | CBN (%AR) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 108 | <0.0030 | <0.0030 | <0.0030 | <0.0030 |  |  |  |  | 1 |
| CBN (%AR) | 005.01 | AOAC 2018.10 | 168 | <0.0250 | <0.0250 | <0.0250 | <0.0250 |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 194 | <0.0270 | <0.0270 | <0.0270 | <0.0270 |  |  |  |  | 1 |
| CBN (%AR) | 005.99 | Other | 137 | <0.0298 | <0.0297 | <0.0298 | <0.0298 |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 132 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| CBN (%AR) | 005.20 | Other LC-MS | 188 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 109 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| CBN (%AR) | 005.01 | AOAC 2018.10 | 193 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 197 | <0.0761 | <0.0753 | <0.0746 | <0.0761 |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 123 | <0.0880 | <0.0910 | <0.0890 | <0.0880 |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 145 | <0.2500 | <0.2500 | <0.2500 | <0.2500 |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 141 | 0.0143 | 0.0164 | 0.0149 | 0.0152 | -1.98 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.99 | Other | 132 | 0.02 | 0.02 | 0.02 | 0.02 | -0.91 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 106 | 0.018 | 0.018 | 0.024 | 0.02 | -0.91 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 144 | 0.0193 | 0.0198 | 0.0211 | 0.0201 | -0.9 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 162 | 0.0192 | 0.0201 | 0.02 | 0.0198 | -0.96 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 199 | 0.018 | 0.021 | 0.02 | 0.0197 | -0.99 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 192 | 0.0206 | 0.0199 | 0.0184 | 0.0196 | -0.99 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.20 | Other LC-MS | 149 | 0.023 | 0.022 | 0.019 | 0.0213 | -0.61 | 0.0241 | 28 | 0.0045 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%AR) | 005.03 | AOAC 2018.11, M | 139 | 0.0227 | 0.0191 | 0.0212 | 0.021 | -0.69 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 207 | 0.021 | 0.021 | 0.023 | 0.0217 | -0.54 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 202 | 0.0207 | 0.0221 | 0.0234 | 0.0221 | -0.45 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 186 | 0.0224 | 0.0229 | 0.024 | 0.0231 | -0.22 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 142 | 0.0233 | 0.0237 | 0.0232 | 0.0234 | -0.15 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 203 | 0.022 | 0.023 |  | 0.0225 | -0.35 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 196 | 0.0244 | 0.0252 | 0.0238 | 0.0245 | 0.09 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 204 | 0.024 | 0.0241 | 0.0239 | 0.024 | -0.02 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 149 | 0.024 | 0.026 | 0.023 | 0.0243 | 0.06 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 178 | 0.024 |  |  | 0.024 |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 200 | 0.0244 | 0.0257 | 0.0252 | 0.0251 | 0.23 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.20 | Other LC-MS | 122 | 0.0258 | 0.025 | 0.0249 | 0.0252 | 0.26 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 189 | 0.026 | 0.025 | 0.027 | 0.026 | 0.43 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 185 | 0.0173 | 0.0238 | 0.0412 | 0.0274 | 0.75 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.01 | AOAC 2018.10 | 155 | 0.026 | 0.03 | 0.024 | 0.0267 | 0.58 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 125 | 0.0271 | 0.0287 | 0.0287 | 0.0282 | 0.91 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.20 | Other LC-MS | 119 | 0.0284 | 0.0276 | 0.0283 | 0.0281 | 0.9 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 107 | 0.0335 | 0.0403 | 0.0343 | 0.036 | 2.67 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 116 | 0.0414 | <0.0400 | <0.0400 | 0.0405 |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 116 | <0.0400 | <0.0400 | 0.0416 | 0.0405 |  |  |  |  | 1 |
| CBN (%AR) | 005.30 | GC-FID | 138 | 0.0425 | 0.047 | 0.0413 | 0.0436 | 4.36 | 0.0241 | 28 | 0.0045 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 198 | 0.0502 | 0.0469 | 0.0464 | 0.0478 | 5.31 | 0.0241 | 28 | 0.0045 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%AR) | 005.30 | GC-FID | 163 | 0.046 | 0.048 | 0.051 | 0.0483 | 5.42 | 0.0241 | 28 | 0.0045 | 0 |
|  |  |  |  | Total Δ9-THC (%AR) |  |  |  |  |  |  |  |  |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 197 | <0.0761 | 0.0753 | <0.0746 | 0.0753 |  |  |  |  | 1 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 207 | 0.09 | 0.081 | 0.088 | 0.0863 | -1.86 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 141 | 0.0875 | 0.0911 | 0.0856 | 0.0881 | -1.79 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 123 | 0.088 | 0.095 | 0.093 | 0.092 | -1.62 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 114 | 0.0927 | 0.0946 | 0.1028 | 0.0967 | -1.42 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 124 | 0.1032 | 0.0982 | 0.0891 | 0.0968 | -1.42 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 103 | 0.0992 | 0.0955 | 0.0985 | 0.0977 | -1.38 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 132 | 0.103 | 0.105 | 0.092 | 0.1 | -1.28 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 142 | 0.0867 | 0.104 | 0.1095 | 0.1 | -1.28 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 161 | 0.1137 | 0.1035 | 0.1096 | 0.109 | -0.91 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 188 | 0.1091 | 0.1158 | 0.1051 | 0.11 | -0.86 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 192 | 0.1122 | 0.1101 | 0.1143 | 0.112 | -0.77 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 186 | 0.1155 | 0.1077 | 0.1136 | 0.112 | -0.77 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.99 | Other | 132 | 0.11 | 0.11 | 0.12 | 0.113 | -0.72 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 107 | 0.1048 | 0.1261 | 0.1116 | 0.114 | -0.69 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 191 | 0.1155 | 0.1098 | 0.1172 | 0.114 | -0.69 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 203 | 0.12 | 0.11 |  | 0.115 | -0.65 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 199 | 0.1136 | 0.1215 | 0.1084 | 0.115 | -0.67 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 198 | 0.1241 | 0.1171 | 0.1145 | 0.119 | -0.5 | 0.1304 | 61 | 0.0237 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 175 | 0.1203 | 0.1191 | 0.1196 | 0.12 | -0.45 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 149 | 0.12 | 0.126 | 0.117 | 0.121 | -0.4 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 200 | 0.1225 | 0.1218 | 0.1188 | 0.121 | -0.4 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 202 | 0.1241 | 0.1225 | 0.1196 | 0.122 | -0.35 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 122 | 0.1257 | 0.1209 | 0.1209 | 0.123 | -0.33 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 120 | 0.124 | 0.1202 | 0.1247 | 0.123 | -0.31 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 121 | 0.138 | 0.115 | 0.12 | 0.124 | -0.26 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 178 | 0.125 |  |  | 0.125 |  |  |  |  | 1 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 101 | 0.138 | 0.106 | 0.134 | 0.126 | -0.19 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 149 | 0.124 | 0.136 | 0.121 | 0.127 | -0.14 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 125 | 0.126 | 0.1239 | 0.1304 | 0.127 | -0.15 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 116 | 0.1316 | 0.1301 | 0.1221 | 0.128 | -0.1 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 144 | 0.1217 | 0.1246 | 0.138 | 0.128 | -0.1 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 182 | 0.1214 | 0.1413 | 0.121 | 0.128 | -0.11 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 108 | 0.1389 | 0.121 | 0.1303 | 0.13 | -0.01 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.03 | AOAC 2018.11, M | 139 | 0.1394 | 0.1335 | 0.1186 | 0.131 | 0 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 155 | 0.142 | 0.139 | 0.115 | 0.132 | 0.07 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 115 | 0.13 | 0.13 | 0.14 | 0.133 | 0.12 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 138 | 0.1378 | 0.1403 | 0.1264 | 0.135 | 0.19 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 109 | 0.129 | 0.131 | 0.148 | 0.136 | 0.24 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 150 | 0.144 | 0.129 | 0.137 | 0.137 | 0.26 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 168 | 0.1409 | 0.1413 | 0.1303 | 0.138 | 0.3 | 0.1304 | 61 | 0.0237 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 204 | 0.1421 | 0.139 | 0.1362 | 0.139 | 0.37 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 172 | 0.1449 | 0.1374 | 0.1346 | 0.139 | 0.36 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 162 | 0.1431 | 0.14 | 0.1369 | 0.14 | 0.41 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 165 | 0.1448 | 0.1301 | 0.1496 | 0.141 | 0.47 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 194 | 0.1374 | 0.1364 | 0.1492 | 0.141 | 0.45 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 112 | 0.14 | 0.15 | 0.14 | 0.143 | 0.55 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 119 | 0.1467 | 0.1384 | 0.1431 | 0.143 | 0.52 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 164 | 0.1406 | 0.1475 | 0.1422 | 0.143 | 0.55 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 152 | 0.1482 | 0.1389 | 0.1427 | 0.143 | 0.54 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 116 | 0.1492 | 0.1406 | 0.1489 | 0.146 | 0.67 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.99 | Other | 172 | 0.155 | 0.1434 | 0.1485 | 0.149 | 0.78 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 196 | 0.1542 | 0.1552 | 0.1424 | 0.151 | 0.85 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 153 | 0.1451 | 0.1695 | 0.1521 | 0.156 | 1.06 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 185 | 0.1445 | 0.1873 | 0.1402 | 0.157 | 1.14 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 189 | 0.165 | 0.164 | 0.164 | 0.164 | 1.43 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.60 | NIR | 187 | 0.1719 | 0.1774 | 0.1756 | 0.175 | 1.88 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 163 | 0.184 | 0.183 | 0.185 | 0.184 | 2.26 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 169 | 0.2151 | 0.1754 | 0.1612 | 0.184 | 2.26 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 170 | 0.228 | 0.212 | 0.2156 | 0.219 | 3.72 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 128 | 0.3175 | 0.3054 | 0.3259 | 0.316 | 7.85 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.60 | NIR | 138 | 0.49 | 0.415 | 0.46 | 0.455 | 13.71 | 0.1304 | 61 | 0.0237 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 145 | 0.97 | 1.03 | 0.98 | 0.993 | 36.46 | 0.1304 | 61 | 0.0237 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Total CBD (%AR) |  |  |  |  |  |  |  |  |
| Total CBD (%AR) | 007.10 | Other LC-UV | 132 | 3.118 | 2.88 | 3.078 | 3.03 | -2.41 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.99 | Other | 132 | 3.32 | 3.13 | 3.32 | 3.26 | -2 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 122 | 3.6031 | 3.5376 | 3.5011 | 3.55 | -1.47 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.01 | AOAC 2018.10 | 155 | 3.601 | 3.659 | 3.537 | 3.6 | -1.38 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 142 | 3.5937 | 3.5322 | 3.7107 | 3.61 | -1.36 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 150 | 3.672 | 3.609 | 3.672 | 3.65 | -1.29 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 202 | 3.7196 | 3.614 | 3.8655 | 3.73 | -1.14 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 207 | 3.691 | 3.699 | 3.878 | 3.76 | -1.1 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 199 | 3.5807 | 4.0501 | 3.7688 | 3.8 | -1.02 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 161 | 3.782 | 3.7229 | 3.9404 | 3.82 | -0.99 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.03 | AOAC 2018.11, M | 139 | 4.2329 | 3.6932 | 3.6976 | 3.87 | -0.88 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 149 | 3.552 | 4.009 | 4.181 | 3.91 | -0.81 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 108 | 4.097 | 3.984 | 3.922 | 4 | -0.66 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 141 | 4.0817 | 4.12 | 3.9784 | 4.06 | -0.55 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 189 | 4.207 | 4.225 | 4.192 | 4.21 | -0.28 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 123 | 4.224 | 4.253 | 4.17 | 4.22 | -0.27 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 178 | 4.249 |  |  | 4.25 |  |  |  |  | 1 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 125 | 4.4485 | 4.2464 | 4.0823 | 4.26 | -0.19 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 152 | 4.3499 | 4.2225 | 4.2236 | 4.27 | -0.18 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 149 | 4.062 | 4.695 | 4.133 | 4.3 | -0.12 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 186 | 4.3127 | 4.2842 | 4.3422 | 4.31 | -0.09 | 4.3651 | 49 | 0.5552 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%AR) | 007.40 | GC-FID | 169 | 4.9239 | 4.1514 | 3.9288 | 4.33 | -0.05 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 200 | 4.2509 | 4.3168 | 4.4231 | 4.33 | -0.06 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 198 | 4.4982 | 4.4152 | 4.1025 | 4.34 | -0.05 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 192 | 4.5323 | 4.2923 | 4.2191 | 4.35 | -0.03 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 203 | 4.44 | 4.35 |  | 4.39 | 0.05 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 196 | 4.4294 | 4.4528 | 4.3421 | 4.41 | 0.08 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 204 | 4.4419 | 4.4295 | 4.4388 | 4.44 | 0.13 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.60 | NIR | 187 | 4.5025 | 4.4756 | 4.3722 | 4.45 | 0.15 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 116 | 4.3952 | 4.4187 | 4.5534 | 4.46 | 0.16 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 144 | 4.2962 | 4.4015 | 4.7038 | 4.47 | 0.18 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 119 | 4.6758 | 4.3579 | 4.4509 | 4.49 | 0.23 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.01 | AOAC 2018.10 | 168 | 4.5905 | 4.3958 | 4.5409 | 4.51 | 0.26 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 191 | 4.4999 | 4.5727 | 4.6453 | 4.57 | 0.37 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 163 | 4.549 | 4.614 | 4.597 | 4.59 | 0.4 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 197 | 4.511 | 4.639 | 4.663 | 4.6 | 0.43 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 188 | 4.6664 | 4.7032 | 4.5241 | 4.63 | 0.48 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 162 | 4.5995 | 4.5692 | 4.7388 | 4.64 | 0.49 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 116 | 4.775 | 4.6037 | 4.5611 | 4.65 | 0.51 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 172 | 4.8708 | 4.704 | 4.4787 | 4.68 | 0.58 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 164 | 4.8126 | 4.8409 | 4.7028 | 4.79 | 0.76 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 170 | 4.9267 | 4.7445 | 4.7446 | 4.81 | 0.79 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 107 | 5.4187 | 4.816 | 4.3956 | 4.88 | 0.92 | 4.3651 | 49 | 0.5552 | 0 |

Population of Lab Values

Lab

Lab

Population of Lab Values

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%AR) | 007.10 | Other LC-UV | 194 | 4.8116 | 4.8553 | 5.0906 | 4.92 | 1 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.99 | Other | 172 | 5.4126 | 5.2273 | 4.9768 | 5.21 | 1.51 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 109 | 5.351 | 5.416 | 4.967 | 5.24 | 1.58 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 138 | 5.5518 | 6.5945 | 5.4367 | 5.86 | 2.69 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.60 | NIR | 138 | 5.975 | 6.1 | 6.01 | 6.03 | 3 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 185 | 7.2736 | 7.4277 | 7.7077 | 7.47 | 5.59 | 4.3651 | 49 | 0.5552 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 145 | 47.29 | 47.43 | 47.62 | 47.4 | 77.59 | 4.3651 | 49 | 0.5552 | 0 |

|  |  |  |  | CBDV (%AR) |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDV (%AR) | 008.99 | Other | 202 | <0.0003 | <0.0003 | <0.0003 | <0.0003 | 1 |
| CBDV (%AR) | 008.99 | Other | 203 | <0.0040 | <0.0040 |  | <0.0040 | 1 |
| CBDV (%AR) | 008.02 | AOAC 2018.11, U | 108 | <0.0050 | <0.0050 | <0.0050 | <0.0050 | 1 |
| CBDV (%AR) | 008.99 | Other | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| CBDV (%AR) | 008.99 | Other | 109 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| CBDV (%AR) | 008.99 | Other | 207 | <0.0130 | <0.0130 | <0.0130 | <0.0130 | 1 |
| CBDV (%AR) | 008.99 | Other | 186 | <0.0200 | <0.0200 | <0.0200 | <0.0200 | 1 |
| CBDV (%AR) | 008.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 | <0.0205 | 1 |
| CBDV (%AR) | 008.99 | Other | 194 | <0.0270 | <0.0270 | <0.0270 | <0.0270 | 1 |
| CBDV (%AR) | 008.99 | Other | 193 | <0.0500 | <0.0500 | <0.0500 | <0.0500 | 1 |
| CBDV (%AR) | 008.99 | Other | 132 | <0.0500 | <0.0500 | <0.0500 | <0.0500 | 1 |
| CBDV (%AR) | 008.99 | Other | 188 | <0.0500 | <0.0500 | <0.0500 | <0.0500 | 1 |
| CBDV (%AR) | 008.99 | Other | 197 | <0.0761 | <0.0753 | <0.0746 | <0.0761 | 1 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDV (%AR) | 008.99 | Other | 200 | 0.0029 | 0.003 | 0.0036 | 0.0032 | -1.03 | 0.0369 | 7 | 0.0327 | 0 |
| CBDV (%AR) | 008.99 | Other | 149 | 0.013 | 0.012 | 0.011 | 0.012 | -0.76 | 0.0369 | 7 | 0.0327 | 0 |
| CBDV (%AR) | 008.99 | Other | 142 | 0.021 | 0.025 | 0.026 | 0.024 | -0.39 | 0.0369 | 7 | 0.0327 | 0 |
| CBDV (%AR) | 008.99 | Other | 155 | 0.031 | 0.031 | 0.026 | 0.0293 | -0.23 | 0.0369 | 7 | 0.0327 | 0 |
| CBDV (%AR) | 008.99 | Other | 106 | 0.04 | 0.048 | 0.041 | 0.043 | 0.19 | 0.0369 | 7 | 0.0327 | 0 |
| CBDV (%AR) | 008.02 | AOAC 2018.11, U | 132 | 0.074 | 0.058 | <0.0500 | 0.0607 | 0.73 | 0.0369 | 7 | 0.0327 | 0 |
| CBDV (%AR) | 008.99 | Other | 189 | 0.163 | 0.162 | 0.16 | 0.162 | 3.82 | 0.0369 | 7 | 0.0327 | 0 |
|  |  |  |  | CBDVA (%AR) |  |  |  |  |  |  |  |  |
| CBDVA (%AR) | 009.99 | Other | 203 | <0.0100 | <0.0110 |  | <0.0100 |  |  |  |  | 1 |
| CBDVA (%AR) | 009.99 | Other | 186 | 0.0189 | 0.0201 | 0.0266 | 0.0219 | -0.93 | 0.0444 | 6 | 0.0241 | 0 |
| CBDVA (%AR) | 009.99 | Other | 149 | 0.035 | 0.033 | 0.029 | 0.0323 | -0.5 | 0.0444 | 6 | 0.0241 | 0 |
| CBDVA (%AR) | 009.01 | AOAC 2018.10 | 155 | 0.036 | 0.036 | 0.034 | 0.0353 | -0.38 | 0.0444 | 6 | 0.0241 | 0 |
| CBDVA (%AR) | 009.99 | Other | 196 | 0.0385 | 0.0392 | 0.0375 | 0.0384 | -0.25 | 0.0444 | 6 | 0.0241 | 0 |
| CBDVA (%AR) | 009.02 | AOAC 2018.11, U | 108 | 0.0569 | 0.0584 | 0.0585 | 0.0579 | 0.56 | 0.0444 | 6 | 0.0241 | 0 |
| CBDVA (%AR) | 009.01 | AOAC 2018.10 | 168 | 0.0881 | 0.069 | 0.0853 | 0.0808 | 1.51 | 0.0444 | 6 | 0.0241 | 0 |
|  |  |  |  | CBG (%AR) |  |  |  |  |  |  |  |  |
| CBG (%AR) | 010.99 | Other | 106 | <0.0000 | <0.0000 | <0.0000 | <0.0000 |  |  |  |  | 1 |
| CBG (%AR) | 010.99 | Other | 186 | <0.0200 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  | 1 |
| CBG (%AR) | 010.99 | Other | 194 | <0.0270 | <0.0270 | <0.0270 | <0.0270 |  |  |  |  | 1 |
| CBG (%AR) | 010.99 | Other | 109 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |

Population of Lab Values

Lab

Lab

Population of Lab Values

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%AR) | 010.01 | AOAC 2018.10 | 193 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| CBG (%AR) | 010.99 | Other | 188 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| CBG (%AR) | 010.99 | Other | 197 | <0.0761 | <0.0753 | <0.0746 | <0.0761 |  |  |  |  | 1 |
| CBG (%AR) | 010.99 | Other | 123 | <0.0880 | <0.0910 | <0.0890 | <0.0880 |  |  |  |  | 1 |
| CBG (%AR) | 010.99 | Other | 178 | 0.013 |  |  | 0.013 |  |  |  |  | 1 |
| CBG (%AR) | 010.99 | Other | 162 | <0.0100 | 0.0331 | 0.0207 | 0.0213 | -0.93 | 0.0342 | 19 | 0.0139 | 0 |
| CBG (%AR) | 010.99 | Other | 142 | 0.0325 | 0.0155 | 0.0168 | 0.0216 | -0.9 | 0.0342 | 19 | 0.0139 | 0 |
| CBG (%AR) | 010.99 | Other | 196 | 0.0223 | 0.0215 | 0.0233 | 0.0224 | -0.85 | 0.0342 | 19 | 0.0139 | 0 |
| CBG (%AR) | 010.99 | Other | 149 | 0.024 | 0.026 | 0.022 | 0.024 | -0.73 | 0.0342 | 19 | 0.0139 | 0 |
| CBG (%AR) | 010.99 | Other | 202 | 0.023 | 0.0233 | 0.0244 | 0.0236 | -0.76 | 0.0342 | 19 | 0.0139 | 0 |
| CBG (%AR) | 010.01 | AOAC 2018.10 | 155 | <0.0227 | <0.0227 | 0.026 | 0.0238 |  |  |  |  | 1 |
| CBG (%AR) | 010.99 | Other | 189 | 0.019 | 0.026 | 0.026 | 0.0237 | -0.75 | 0.0342 | 19 | 0.0139 | 0 |
| CBG (%AR) | 010.99 | Other | 203 | 0.026 | 0.024 |  | 0.025 | -0.66 | 0.0342 | 19 | 0.0139 | 0 |
| CBG (%AR) | 010.99 | Other | 122 | 0.0265 | 0.0256 | 0.0256 | 0.0259 | -0.59 | 0.0342 | 19 | 0.0139 | 0 |
| CBG (%AR) | 010.99 | Other | 200 | 0.0253 | 0.0292 | 0.0311 | 0.0285 | -0.4 | 0.0342 | 19 | 0.0139 | 0 |
| CBG (%AR) | 010.99 | Other | 119 | 0.0311 | 0.0301 | 0.0308 | 0.0307 | -0.25 | 0.0342 | 19 | 0.0139 | 0 |
| CBG (%AR) | 010.99 | Other | 141 | 0.0301 | 0.0375 | 0.0272 | 0.0316 | -0.18 | 0.0342 | 19 | 0.0139 | 0 |
| CBG (%AR) | 010.99 | Other | 207 | 0.033 | 0.026 | 0.043 | 0.034 | -0.01 | 0.0342 | 19 | 0.0139 | 0 |
| CBG (%AR) | 010.99 | Other | 204 | 0.0349 | 0.0333 | 0.0361 | 0.0348 | 0.04 | 0.0342 | 19 | 0.0139 | 0 |
| CBG (%AR) | 010.01 | AOAC 2018.10 | 168 | 0.051 | 0.0281 | 0.0466 | 0.0419 | 0.56 | 0.0342 | 19 | 0.0139 | 0 |
| CBG (%AR) | 010.99 | Other | 132 | <0.0500 | 0.04 | 0.05 | 0.0467 | 0.9 | 0.0342 | 19 | 0.0139 | 0 |
| CBG (%AR) | 010.02 | AOAC 2018.11, U | 108 | 0.049 | 0.049 | 0.0468 | 0.0483 | 1.01 | 0.0342 | 19 | 0.0139 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%AR) | 010.02 | AOAC 2018.11, U | 152 | 0.1104 | 0.1092 | 0.1164 | 0.112 | 5.59 | 0.0342 | 19 | 0.0139 | 0 |
| CBG (%AR) | 010.99 | Other | 125 | 0.1402 | 0.134 | 0.1362 | 0.137 | 7.37 | 0.0342 | 19 | 0.0139 | 0 |
| CBG (%AR) | 010.02 | AOAC 2018.11, U | 132 | 0.229 | 0.217 | 0.224 | 0.223 | 13.58 | 0.0342 | 19 | 0.0139 | 0 |

Population of Lab Values

|  |  |  |  | CBGA (%AR) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%AR) | 011.02 | AOAC 2018.11, U | 108 | <0.0008 | <0.0008 | <0.0008 | <0.0008 |  |  |  |  | 1 |
| CBGA (%AR) | 011.99 | Other | 203 | <0.0100 | <0.0110 |  | <0.0100 |  |  |  |  | 1 |
| CBGA (%AR) | 011.99 | Other | 109 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| CBGA (%AR) | 011.01 | AOAC 2018.10 | 193 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| CBGA (%AR) | 011.99 | Other | 197 | <0.0761 | <0.0753 | <0.0746 | <0.0761 |  |  |  |  | 1 |
| CBGA (%AR) | 011.99 | Other | 194 | <0.0800 | <0.0800 | <0.0800 | <0.0800 |  |  |  |  | 1 |
| CBGA (%AR) | 011.99 | Other | 123 | <0.0880 | <0.0910 | <0.0890 | <0.0880 |  |  |  |  | 1 |
| CBGA (%AR) | 011.99 | Other | 149 | 0.026 | 0.027 | 0.026 | 0.0263 | -1.56 | 0.0434 | 19 | 0.011 | 0 |
| CBGA (%AR) | 011.99 | Other | 106 | 0.032 | 0.033 | 0.032 | 0.0323 | -1.01 | 0.0434 | 19 | 0.011 | 0 |
| CBGA (%AR) | 011.99 | Other | 207 | 0.027 | 0.034 | 0.036 | 0.0323 | -1.01 | 0.0434 | 19 | 0.011 | 0 |
| CBGA (%AR) | 011.99 | Other | 141 | 0.0307 | 0.0361 | 0.0329 | 0.0332 | -0.93 | 0.0434 | 19 | 0.011 | 0 |
| CBGA (%AR) | 011.02 | AOAC 2018.11, U | 132 | 0.035 | 0.031 | 0.034 | 0.0333 | -0.92 | 0.0434 | 19 | 0.011 | 0 |
| CBGA (%AR) | 011.02 | AOAC 2018.11, U | 152 | 0.0349 | 0.0347 | 0.0324 | 0.034 | -0.86 | 0.0434 | 19 | 0.011 | 0 |
| CBGA (%AR) | 011.99 | Other | 122 | 0.0379 | 0.038 | 0.0376 | 0.0378 | -0.51 | 0.0434 | 19 | 0.011 | 0 |
| CBGA (%AR) | 011.99 | Other | 178 | 0.038 |  |  | 0.038 |  |  |  |  | 1 |
| CBGA (%AR) | 011.01 | AOAC 2018.10 | 155 | 0.041 | 0.043 | 0.035 | 0.0397 | -0.34 | 0.0434 | 19 | 0.011 | 0 |
| CBGA (%AR) | 011.99 | Other | 196 | 0.0423 | 0.042 | 0.0412 | 0.0418 | -0.14 | 0.0434 | 19 | 0.011 | 0 |

| Analyte Code |  | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%AR) | 011.99 | Other | 162 | 0.0471 | 0.0424 | 0.0403 | 0.0433 | -0.01 | 0.0434 | 19 | 0.011 | 0 |
| CBGA (%AR) | 011.99 | Other | 202 | 0.0387 | 0.0454 | 0.0454 | 0.0432 | -0.02 | 0.0434 | 19 | 0.011 | 0 |
| CBGA (%AR) | 011.99 | Other | 142 | 0.052 | 0.0473 | 0.0503 | 0.0499 | 0.59 | 0.0434 | 19 | 0.011 | 0 |
| CBGA (%AR) | 011.99 | Other | 186 | 0.054 | 0.0521 | 0.048 | 0.0514 | 0.73 | 0.0434 | 19 | 0.011 | 0 |
| CBGA (%AR) | 011.01 | AOAC 2018.10 | 168 | 0.0536 | 0.0497 | 0.0533 | 0.0522 | 0.8 | 0.0434 | 19 | 0.011 | 0 |
| CBGA (%AR) | 011.99 | Other | 125 | 0.0528 | 0.0536 | 0.0536 | 0.0533 | 0.91 | 0.0434 | 19 | 0.011 | 0 |
| CBGA (%AR) | 011.99 | Other | 132 | 0.04 | <0.0500 | 0.07 | 0.0533 | 0.91 | 0.0434 | 19 | 0.011 | 0 |
| CBGA (%AR) | 011.99 | Other | 188 | 0.0575 | 0.056 | 0.0506 | 0.0547 | 1.03 | 0.0434 | 19 | 0.011 | 0 |
| CBGA (%AR) | 011.99 | Other | 204 | 0.056 | 0.0545 | 0.0544 | 0.055 | 1.05 | 0.0434 | 19 | 0.011 | 0 |
| CBGA (%AR) | 011.99 | Other | 189 | 0.055 | 0.059 | 0.057 | 0.057 | 1.24 | 0.0434 | 19 | 0.011 | 0 |

Population of Lab Values

|  |  |  |  | THCV (%AR) |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| THCV (%AR) | 012.99 | Other | 106 | <0.0000 | <0.0000 | <0.0000 | <0.0000 | 1 |
| THCV (%AR) | 012.99 | Other | 203 | <0.0020 | <0.0020 |  | <0.0020 | 1 |
| THCV (%AR) | 012.02 | AOAC 2018.11, U | 108 | <0.0051 | <0.0051 | <0.0051 | <0.0051 | 1 |
| THCV (%AR) | 012.99 | Other | 109 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| THCV (%AR) | 012.99 | Other | 142 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| THCV (%AR) | 012.99 | Other | 186 | <0.0200 | <0.0200 | <0.0200 | <0.0200 | 1 |
| THCV (%AR) | 012.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 | <0.0205 | 1 |
| THCV (%AR) | 012.99 | Other | 168 | <0.0250 | <0.0250 | <0.0250 | <0.0250 | 1 |
| THCV (%AR) | 012.99 | Other | 194 | <0.0270 | <0.0270 | <0.0270 | <0.0270 | 1 |
| THCV (%AR) | 012.99 | Other | 188 | <0.0500 | <0.0500 | <0.0500 | <0.0500 | 1 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| THCV (%AR) | 012.99 | Other | 132 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 197 | <0.0761 | <0.0753 | <0.0746 | <0.0761 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 123 | <0.0880 | <0.0910 | <0.0890 | <0.0880 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 202 | <0.0003 | <0.0003 | 0.0026 | 0.0011 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 149 | 0.007 | 0.007 | 0.005 | 0.0063 |  |  |  |  | 0 |
| THCV (%AR) | 012.99 | Other | 200 | 0.0105 | 0.0049 | 0.0043 | 0.0066 |  |  |  |  | 0 |
| THCV (%AR) | 012.99 | Other | 119 | 0.0104 | 0.0097 | 0.0104 | 0.0102 |  |  |  |  | 0 |
| THCV (%AR) | 012.02 | AOAC 2018.11, U | 132 | 0.04 | 0.04 | <0.0500 | 0.0433 |  |  |  |  | 0 |
| THCV (%AR) | 012.99 | Other | 189 | 0.075 | 0.083 | 0.081 | 0.0797 |  |  |  |  | 0 |
|  |  |  |  | Δ8-THC (%AR) |  |  |  |  |  |  |  |  |
| Δ8-THC (%AR) | 013.99 | Other | 106 | <0.0000 | <0.0000 | <0.0000 | <0.0000 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 202 | <0.0001 | <0.0001 | <0.0001 | <0.0001 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 203 | <0.0020 | <0.0020 |  | <0.0020 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 200 | <0.0025 | <0.0025 | <0.0025 | <0.0025 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.02 | AOAC 2018.11, U | 108 | <0.0045 | <0.0045 | <0.0045 | <0.0045 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 109 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 142 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 141 | <0.0120 | <0.0120 | <0.0120 | <0.0120 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 207 | <0.0130 | <0.0130 | <0.0130 | <0.0130 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 186 | <0.0200 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  | 1 |

Population of Lab Values

Lab

Lab

Population of Lab Values

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%AR) | 013.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 | <0.0205 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 149 | <0.0240 | <0.0240 | <0.0220 | <0.0240 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 194 | <0.0270 | <0.0270 | <0.0270 | <0.0270 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 188 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.01 | AOAC 2018.10 | 193 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 197 | <0.0761 | <0.0753 | <0.0746 | <0.0761 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 123 | <0.0880 | <0.0910 | <0.0890 | <0.0880 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 189 | 0.013 | 0.012 | 0.012 | 0.0123 |  |  |  |  | 0 |

|  |  |  |  | CBC (%AR) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%AR) | 014.99 | Other | 202 | 0.0669 | 0.1103 | 0.055 | 0.0774 | -3.72 | 0.1175 | 23 | 0.0108 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, U | 132 | 0.119 | 0.101 | 0.098 | 0.106 | -1.07 | 0.1175 | 23 | 0.0108 | 0 |
| CBC (%AR) | 014.99 | Other | 186 | 0.1068 | 0.1072 | 0.1072 | 0.107 | -0.97 | 0.1175 | 23 | 0.0108 | 0 |
| CBC (%AR) | 014.99 | Other | 197 | 0.104 | 0.111 | 0.105 | 0.107 | -1.01 | 0.1175 | 23 | 0.0108 | 0 |
| CBC (%AR) | 014.99 | Other | 207 | 0.109 | 0.105 | 0.112 | 0.109 | -0.82 | 0.1175 | 23 | 0.0108 | 0 |
| CBC (%AR) | 014.99 | Other | 203 | 0.111 | 0.108 |  | 0.109 | -0.74 | 0.1175 | 23 | 0.0108 | 0 |
| CBC (%AR) | 014.99 | Other | 200 | 0.1124 | 0.1123 | 0.1081 | 0.111 | -0.61 | 0.1175 | 23 | 0.0108 | 0 |
| CBC (%AR) | 014.99 | Other | 149 | 0.11 | 0.118 | 0.108 | 0.112 | -0.51 | 0.1175 | 23 | 0.0108 | 0 |
| CBC (%AR) | 014.99 | Other | 132 | 0.1 | 0.11 | 0.13 | 0.113 | -0.39 | 0.1175 | 23 | 0.0108 | 0 |
| CBC (%AR) | 014.99 | Other | 125 | 0.1124 | 0.1139 | 0.1115 | 0.113 | -0.46 | 0.1175 | 23 | 0.0108 | 0 |
| CBC (%AR) | 014.99 | Other | 142 | 0.1188 | 0.1088 | 0.114 | 0.114 | -0.34 | 0.1175 | 23 | 0.0108 | 0 |
| CBC (%AR) | 014.01 | AOAC 2018.10 | 155 | 0.119 | 0.129 | 0.097 | 0.115 | -0.23 | 0.1175 | 23 | 0.0108 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%AR) | 014.01 | AOAC 2018.10 | 168 | 0.116 | 0.1159 | 0.1152 | 0.116 | -0.17 | 0.1175 | 23 | 0.0108 | 0 |
| CBC (%AR) | 014.99 | Other | 204 | 0.1191 | 0.1198 | 0.1193 | 0.119 | 0.18 | 0.1175 | 23 | 0.0108 | 0 |
| CBC (%AR) | 014.99 | Other | 106 | 0.123 | 0.122 | 0.123 | 0.123 | 0.48 | 0.1175 | 23 | 0.0108 | 0 |
| CBC (%AR) | 014.99 | Other | 123 | 0.13 | 0.122 | 0.122 | 0.125 | 0.66 | 0.1175 | 23 | 0.0108 | 0 |
| CBC (%AR) | 014.99 | Other | 196 | 0.1279 | 0.125 | 0.1233 | 0.125 | 0.73 | 0.1175 | 23 | 0.0108 | 0 |
| CBC (%AR) | 014.99 | Other | 162 | 0.1259 | 0.1316 | 0.1191 | 0.126 | 0.75 | 0.1175 | 23 | 0.0108 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, U | 108 | 0.1196 | 0.1341 | 0.1261 | 0.127 | 0.84 | 0.1175 | 23 | 0.0108 | 0 |
| CBC (%AR) | 014.99 | Other | 188 | 0.1273 | 0.1337 | 0.125 | 0.129 | 1.04 | 0.1175 | 23 | 0.0108 | 0 |
| CBC (%AR) | 014.99 | Other | 189 | 0.13 | 0.13 | 0.129 | 0.13 | 1.13 | 0.1175 | 23 | 0.0108 | 0 |
| CBC (%AR) | 014.99 | Other | 109 | 0.138 | 0.14 | 0.13 | 0.136 | 1.72 | 0.1175 | 23 | 0.0108 | 0 |
| CBC (%AR) | 014.99 | Other | 119 | 0.1555 | 0.1493 | 0.1533 | 0.153 | 3.27 | 0.1175 | 23 | 0.0108 | 0 |
|  |  |  |  | Cd (ug/g AR) |  |  |  |  |  |  |  |  |
| Cd (ug/g AR) | 101.04 | ICP-MS | 201 | 0.002 | 0.002 | 0.001 | 0.0017 | -6.18 | 0.1737 | 13 | 0.0278 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 183 | 0.1225 | 0.128 | 0.1575 | 0.136 | -1.35 | 0.1737 | 13 | 0.0278 | 0 |
| Cd (ug/g AR) | 101.03 | ICP-OES | 132 | 0.152 | 0.14 | 0.156 | 0.149 | -0.87 | 0.1737 | 13 | 0.0278 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 162 | 0.1605 | 0.161 | 0.1699 | 0.164 | -0.35 | 0.1737 | 13 | 0.0278 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 145 | 0.181 | 0.168 | 0.161 | 0.17 | -0.13 | 0.1737 | 13 | 0.0278 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 196 | 0.167 | 0.169 | 0.174 | 0.17 | -0.13 | 0.1737 | 13 | 0.0278 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 142 | 0.164 | 0.181 | 0.174 | 0.173 | -0.02 | 0.1737 | 13 | 0.0278 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 137 | 0.17 | 0.187 | 0.179 | 0.179 | 0.18 | 0.1737 | 13 | 0.0278 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 122 | 0.1811 | 0.1818 | 0.1811 | 0.181 | 0.28 | 0.1737 | 13 | 0.0278 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cd (ug/g AR) | 101.04 | ICP-MS | 194 | 0.1876 | 0.1869 | 0.1834 | 0.186 | 0.44 | 0.1737 | 13 | 0.0278 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 197 | 0.2046 | 0.2011 | 0.2049 | 0.204 | 1.07 | 0.1737 | 13 | 0.0278 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 141 | 0.179 | 0.222 | 0.217 | 0.206 | 1.16 | 0.1737 | 13 | 0.0278 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 168 | 0.194 | 0.216 | 0.214 | 0.208 | 1.23 | 0.1737 | 13 | 0.0278 | 0 |
|  |  |  |  | As (ug/g AR) |  |  |  |  |  |  |  |  |
| As (ug/g AR) | 102.04 | ICP-MS | 201 | 0.011 | 0.009 | 0.011 | 0.0103 | -6.61 | 0.6422 | 12 | 0.0956 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 183 | 0.528 | 0.538 | 0.634 | 0.567 | -0.79 | 0.6422 | 12 | 0.0956 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 168 | 0.54 | 0.604 | 0.61 | 0.585 | -0.6 | 0.6422 | 12 | 0.0956 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 145 | 0.629 | 0.584 | 0.552 | 0.588 | -0.56 | 0.6422 | 12 | 0.0956 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 142 | 0.556 | 0.604 | 0.612 | 0.591 | -0.54 | 0.6422 | 12 | 0.0956 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 162 | 0.6063 | 0.5931 | 0.707 | 0.635 | -0.07 | 0.6422 | 12 | 0.0956 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 197 | 0.6545 | 0.5979 | 0.665 | 0.639 | -0.03 | 0.6422 | 12 | 0.0956 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 137 | 0.661 | 0.699 | 0.601 | 0.654 | 0.12 | 0.6422 | 12 | 0.0956 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 122 | 0.6736 | 0.6734 | 0.6739 | 0.674 | 0.33 | 0.6422 | 12 | 0.0956 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 194 | 0.7038 | 0.7531 | 0.7457 | 0.734 | 0.96 | 0.6422 | 12 | 0.0956 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 141 | 0.803 | 0.723 | 0.741 | 0.756 | 1.19 | 0.6422 | 12 | 0.0956 | 0 |
| As (ug/g AR) | 102.03 | ICP-OES | 132 | 0.8 |  |  | 0.8 |  |  |  |  | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 196 | 1.14 | 1.14 | 1.16 | 1.15 | 5.27 | 0.6422 | 12 | 0.0956 | 0 |
|  |  |  |  | Hg (ug/g AR) |  |  |  |  |  |  |  |  |
| Hg (ug/g AR) | 103.04 | ICP-MS | 142 | <0.0005 <0.0005 <0.0005 |  |  | <0.0005 |  |  |  |  | 1 |

Population of Lab Values

Lab

Lab

Population of Lab Values

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Hg (ug/g AR) | 103.04 | ICP-MS | 168 | <0.0130 | <0.0130 | <0.0140 | <0.0130 |  |  |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 141 | <0.0139 | <0.0139 | <0.0139 | <0.0139 |  |  |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 194 | <0.0200 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 197 | <0.0990 | <0.0990 | <0.0958 | <0.0990 |  |  |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 162 | <0.1000 | <0.1000 | <0.1000 | <0.1000 |  |  |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 122 | 0.0054 | 0.005 | 0.0055 | 0.0053 |  |  |  |  | 0 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 137 | 0.01 | 0.0105 | 0.0113 | 0.0106 |  |  |  |  | 0 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 145 | 0.013 | 0.012 | 0.011 | 0.012 |  |  |  |  | 0 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 196 | 0.0133 | 0.0139 | 0.0154 | 0.0142 |  |  |  |  | 0 |

|  |  |  |  | Pb (ug/g AR) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Pb (ug/g AR) | 104.04 | ICP-MS | 201 | 0.016 | 0.017 | 0.017 | 0.0167 | -5.46 | 1.7685 | 13 | 0.3206 | 0 |
| Pb (ug/g AR) | 104.03 | ICP-OES | 132 | 1.6 | 1.2 | 1.2 | 1.33 | -1.36 | 1.7685 | 13 | 0.3206 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 197 | 1.6846 | 1.5391 | 1.6215 | 1.62 | -0.48 | 1.7685 | 13 | 0.3206 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 194 | 1.6034 | 1.6949 | 1.6043 | 1.63 | -0.42 | 1.7685 | 13 | 0.3206 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 142 | 1.488 | 1.813 | 1.613 | 1.64 | -0.41 | 1.7685 | 13 | 0.3206 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 145 | 1.741 | 1.663 | 1.562 | 1.66 | -0.35 | 1.7685 | 13 | 0.3206 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 196 | 1.84 | 1.74 | 1.81 | 1.8 | 0.09 | 1.7685 | 13 | 0.3206 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 168 | 1.83 | 1.78 | 1.93 | 1.85 | 0.24 | 1.7685 | 13 | 0.3206 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 162 | 1.7725 | 1.8222 | 2.0493 | 1.88 | 0.35 | 1.7685 | 13 | 0.3206 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 137 | 1.79 | 2.01 | 1.9 | 1.9 | 0.41 | 1.7685 | 13 | 0.3206 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 183 | 1.89 | 2.205 | 1.91 | 2 | 0.73 | 1.7685 | 13 | 0.3206 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Pb (ug/g AR) | 104.04 | ICP-MS | 141 | 2.22 | 2.11 | 2.15 | 2.16 | 1.22 | 1.7685 | 13 | 0.3206 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 122 | 2.2405 | 2.2404 | 2.2401 | 2.24 | 1.47 | 1.7685 | 13 | 0.3206 | 0 |
|  |  |  |  | Cu (ug/g AR) |  |  |  |  |  |  |  |  |
| Cu (ug/g AR) | 105.03 | ICP-OES | 132 | 10 | 10.4 | 10.4 | 10.3 |  |  |  |  | 0 |
| Cu (ug/g AR) | 105.04 | ICP-MS | 196 | 11 | 10.9 | 10.6 | 10.8 |  |  |  |  | 0 |
| Cu (ug/g AR) | 105.04 | ICP-MS | 142 | 7.653 | 12.899 | 12.607 | 11.1 |  |  |  |  | 0 |
|  |  |  |  | Ni (ug/g AR) |  |  |  |  |  |  |  |  |
| Ni (ug/g AR) | 106.03 | ICP-OES | 132 | 2 | 2 | 2 | 2 |  |  |  |  | 0 |
| Ni (ug/g AR) | 106.04 | ICP-MS | 142 | 7.02 | 7.769 | 6.893 | 7.23 |  |  |  |  | 0 |
| Ni (ug/g AR) | 106.04 | ICP-MS | 196 | 7.66 | 7.34 | 7.54 | 7.51 |  |  |  |  | 0 |
| Ni (ug/g AR) | 106.04 | ICP-MS | 141 | 8.59 | 8.32 | 8.36 | 8.42 |  |  |  |  | 0 |
|  |  |  |  | Cr (ug/g AR) |  |  |  |  |  |  |  |  |
| Cr (ug/g AR) | 107.03 | ICP-OES | 132 | 2 | 2 | 2.4 | 2.13 |  |  |  |  | 0 |
| Cr (ug/g AR) | 107.04 | ICP-MS | 142 | 4.362 | 4.311 | 6.976 | 5.22 |  |  |  |  | 0 |
| Cr (ug/g AR) | 107.04 | ICP-MS | 196 | 15.9 | 15 | 15 | 15.3 |  |  |  |  | 0 |
| Cr (ug/g AR) | 107.04 | ICP-MS | 141 | 17.8 | 17.4 | 17.3 | 17.5 |  |  |  |  | 0 |
| Cr (ug/g AR) | 107.04 | ICP-MS | 197 | 18.9829 | 18.0105 | 19.1447 | 18.7 |  |  |  |  | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Se (ug/g AR) |  |  |  |  |  |  |  |  |  |  |  |
| Se (ug/g AR) | 108.04 | ICP-MS | 142 | 0.191 | 0.215 | 0.213 | 0.206 | 0 |  |  |  |  |
| Se (ug/g AR) | 108.04 | ICP-MS | 197 | 0.2597 | 0.2535 | 0.2462 | 0.253 | 0 |  |  |  |  |
| Se (ug/g AR) | 108.04 | ICP-MS | 141 | 0.509 | 0.397 | 0.426 | 0.444 | 0 |  |  |  |  |
| Se (ug/g AR) | 108.03 | ICP-OES | 132 | 0.8 |  | 0.8 | 0.8 | 0 |  |  |  |  |
| Se (ug/g AR) | 108.04 | ICP-MS | 196 | 20.1 | 22.5 | 22.3 | 21.6 | 0 |  |  |  |  |

Population of Lab Values

Ag (ug/g AR)

Ag (ug/g AR)

109.04	ICP-MS

142

<0.0005  <0.0005	<0.0005	<0.0005	1

Ba (ug/g AR)

Ba (ug/g AR)

110.03	ICP-OES

132

44	46.4	47.6	46	0

|  |  |  |  | Zn (ug/g AR) |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Zn (ug/g AR) | 111.03 | ICP-OES | 132 | 48.8 | 48.8 | 52.8 | 50.1 | 0 |
| Zn (ug/g AR) | 111.04 | ICP-MS | 142 | 62.049 | 53.206 | 53.839 | 56.4 | 0 |
| Zn (ug/g AR) | 111.04 | ICP-MS | 141 | 69.4 | 70.4 | 71.9 | 70.6 | 0 |
| Zn (ug/g AR) | 111.04 | ICP-MS | 196 | 82.4 | 78.6 | 81.6 | 80.9 | 0 |

|  |  |  |  | Sb (ug/g AR) |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Sb (ug/g AR) | 112.04 | ICP-MS | 196 | 0.0382 | 0.0442 | 0.0403 | 0.0409 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Moisture (%AR) |  |  |  |  |  |  |  |  |
| Moisture (%AR) | 500.99 | Other | 103 | 0.9637 | 0.93 | 0.9321 | 0.942 | -3.86 | 7.3771 | 23 | 1.6662 | 0 |
| Moisture (%AR) | 500.99 | Other | 201 | 4 | 3 | 5 | 4 | -2.03 | 7.3771 | 23 | 1.6662 | 0 |
| Moisture (%AR) | 500.99 | Other | 207 | 5.44 | 5.19 | 5.18 | 5.27 | -1.26 | 7.3771 | 23 | 1.6662 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 198 | 6.11 | 6.25 | 5.73 | 6.03 | -0.81 | 7.3771 | 23 | 1.6662 | 0 |
| Moisture (%AR) | 500.72 | NIR | 187 | 6.3938 | 6.2741 | 6.0444 | 6.24 | -0.68 | 7.3771 | 23 | 1.6662 | 0 |
| Moisture (%AR) | 500.99 | Other | 142 | 6.25 | 6.25 | 6.25 | 6.25 | -0.68 | 7.3771 | 23 | 1.6662 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 109 | 7.47 | 6.63 | 5.17 | 6.42 | -0.57 | 7.3771 | 23 | 1.6662 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 186 | 6.42 | 6.92 | 6.54 | 6.63 | -0.45 | 7.3771 | 23 | 1.6662 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 116 | 6.84 | 6.97 | 6.92 | 6.91 | -0.28 | 7.3771 | 23 | 1.6662 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 204 | 7.03 | 7.03 | 7.03 | 7.03 | -0.21 | 7.3771 | 23 | 1.6662 | 0 |
| Moisture (%AR) | 500.99 | Other | 120 | 7.1029 | 7.4943 | 7.2006 | 7.27 | -0.07 | 7.3771 | 23 | 1.6662 | 0 |
| Moisture (%AR) | 500.99 | Other | 132 | 6.9 | 7.69 | 7.63 | 7.41 | 0.02 | 7.3771 | 23 | 1.6662 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 119 | 7.56 | 7.47 | 7.48 | 7.5 | 0.08 | 7.3771 | 23 | 1.6662 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 182 | 7.8 | 7.87 | 7.595 | 7.76 | 0.23 | 7.3771 | 23 | 1.6662 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 196 | 7.68 | 7.53 | 8.33 | 7.85 | 0.28 | 7.3771 | 23 | 1.6662 | 0 |
| Moisture (%AR) | 500.99 | Other | 102 | 7.9 | 7.8 | 8 | 7.9 | 0.31 | 7.3771 | 23 | 1.6662 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 183 | 7.98 | 8.28 | 7.97 | 8.08 | 0.42 | 7.3771 | 23 | 1.6662 | 0 |
| Moisture (%AR) | 500.72 | NIR | 138 | 8 | 8.25 | 8.15 | 8.13 | 0.45 | 7.3771 | 23 | 1.6662 | 0 |
| Moisture (%AR) | 500.99 | Other | 165 | 8.684 | 8.5234 | 8.679 | 8.63 | 0.75 | 7.3771 | 23 | 1.6662 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 199 | 9.3826 | 9.1335 | 9.0886 | 9.2 | 1.09 | 7.3771 | 23 | 1.6662 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 105 | 9.71 | 9.53 | 9.77 | 9.67 | 1.38 | 7.3771 | 23 | 1.6662 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 172 | 10.014 | 10.014 | 10.014 | 10 | 1.58 | 7.3771 | 23 | 1.6662 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 192 | 9.9706 | 9.9555 | 10.3067 | 10.1 | 1.62 | 7.3771 | 23 | 1.6662 | 0 |
|  |  |  |  | Δ9-THC (%DW) |  |  |  |  |  |  |  |  |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 207 | 0.061 | 0.051 | 0.057 | 0.0563 | -2.38 | 0.0916 | 22 | 0.0148 | 0 |
| Δ9-THC (%DW) | 501.30 | Other LC-MS | 102 | 0.0783 | 0.0762 | 0.0711 | 0.0752 | -1.11 | 0.0916 | 22 | 0.0148 | 0 |
| Δ9-THC (%DW) | 501.99 | Other | 154 | 0.079 | 0.074 | 0.0733 | 0.0754 | -1.09 | 0.0916 | 22 | 0.0148 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 118 | 0.083 | 0.08 | 0.08 | 0.081 | -0.71 | 0.0916 | 22 | 0.0148 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 186 | 0.084 | 0.0762 | 0.0824 | 0.0809 | -0.72 | 0.0916 | 22 | 0.0148 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 206 | 0.0804 | 0.0846 | 0.0798 | 0.0816 | -0.67 | 0.0916 | 22 | 0.0148 | 0 |
| Δ9-THC (%DW) | 501.30 | Other LC-MS | 120 | 0.0854 | 0.0813 | 0.0857 | 0.0841 | -0.5 | 0.0916 | 22 | 0.0148 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 198 | 0.0853 | 0.0842 | 0.0815 | 0.0837 | -0.53 | 0.0916 | 22 | 0.0148 | 0 |
| Δ9-THC (%DW) | 501.30 | Other LC-MS | 105 | 0.0866 | 0.0852 | 0.0839 | 0.0852 | -0.43 | 0.0916 | 22 | 0.0148 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 109 | 0.081 | 0.078 | 0.1 | 0.0863 | -0.36 | 0.0916 | 22 | 0.0148 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 194 | 0.088 | 0.081 | 0.094 | 0.0877 | -0.27 | 0.0916 | 22 | 0.0148 | 0 |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, U | 199 | 0.0872 | 0.0946 | 0.083 | 0.0883 | -0.22 | 0.0916 | 22 | 0.0148 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 142 | 0.0925 | 0.0869 | 0.0932 | 0.0909 | -0.05 | 0.0916 | 22 | 0.0148 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 172 | 0.1077 | 0.0708 | 0.1111 | 0.0965 | 0.33 | 0.0916 | 22 | 0.0148 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 165 | 0.0946 | 0.09 | 0.1057 | 0.0968 | 0.35 | 0.0916 | 22 | 0.0148 | 0 |
| Δ9-THC (%DW) | 501.99 | Other | 183 | 0.0999 | 0.0945 | 0.1011 | 0.0985 | 0.46 | 0.0916 | 22 | 0.0148 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 204 | 0.1035 | 0.1049 | 0.1015 | 0.103 | 0.79 | 0.0916 | 22 | 0.0148 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 116 | 0.105 | 0.1021 | 0.1041 | 0.104 | 0.82 | 0.0916 | 22 | 0.0148 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%DW) | 501.99 | Other | 172 | 0.1196 | 0.0787 | 0.1235 | 0.107 | 1.06 | 0.0916 | 22 | 0.0148 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 196 | 0.1136 | 0.115 | 0.1069 | 0.112 | 1.36 | 0.0916 | 22 | 0.0148 | 0 |
| Δ9-THC (%DW) | 501.60 | NIR | 187 | 0.1506 | 0.1537 | 0.16 | 0.155 | 4.26 | 0.0916 | 22 | 0.0148 | 0 |
| Δ9-THC (%DW) | 501.30 | Other LC-MS | 201 | 0.21 | 0.22 | 0.24 | 0.223 | 8.88 | 0.0916 | 22 | 0.0148 | 0 |
|  |  |  |  | Δ9-THCA (%DW) |  |  |  |  |  |  |  |  |
| Δ9-THCA (%DW) | 502.30 | Other LC-MS | 105 | 0.0263 | 0.0195 | 0.0212 | 0.0223 | -1.76 | 0.0529 | 19 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 142 | 0.03 | 0.0273 | 0.0269 | 0.0281 | -1.43 | 0.0529 | 19 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.60 | NIR | 187 | 0.0371 | 0.038 | 0.0294 | 0.0348 | -1.04 | 0.0529 | 19 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 207 | 0.039 | 0.039 | 0.041 | 0.0397 | -0.76 | 0.0529 | 19 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 118 | 0.042 | 0.042 | 0.039 | 0.041 | -0.69 | 0.0529 | 19 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, U | 199 | 0.0436 | 0.0446 | 0.0412 | 0.0431 | -0.56 | 0.0529 | 19 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 186 | 0.0452 | 0.0447 | 0.0448 | 0.0449 | -0.46 | 0.0529 | 19 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 198 | 0.0533 | 0.0462 | 0.0459 | 0.0485 | -0.25 | 0.0529 | 19 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 204 | 0.0563 | 0.0509 | 0.0513 | 0.0528 | 0 | 0.0529 | 19 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 206 | 0.0502 | 0.0575 | 0.0539 | 0.0539 | 0.06 | 0.0529 | 19 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.30 | Other LC-MS | 120 | 0.0548 | 0.0555 | 0.0554 | 0.0552 | 0.14 | 0.0529 | 19 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.99 | Other | 154 | 0.0592 | 0.0593 | 0.0508 | 0.0564 | 0.21 | 0.0529 | 19 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 196 | 0.0609 | 0.0603 | 0.0552 | 0.0588 | 0.34 | 0.0529 | 19 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 116 | 0.063 | 0.0558 | 0.0637 | 0.0608 | 0.46 | 0.0529 | 19 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 165 | 0.0729 | 0.0607 | 0.0665 | 0.0667 | 0.8 | 0.0529 | 19 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.99 | Other | 183 | 0.0678 | 0.067 | 0.0676 | 0.0675 | 0.84 | 0.0529 | 19 | 0.0173 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 109 | 0.067 | 0.072 | 0.064 | 0.0677 | 0.85 | 0.0529 | 19 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 172 | 0.0886 | 0.098 | 0.0915 | 0.0927 | 2.3 | 0.0529 | 19 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.99 | Other | 172 | 0.0985 | 0.1089 | 0.1017 | 0.103 | 2.9 | 0.0529 | 19 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.30 | Other LC-MS | 201 | 0.61 | <0.1000 | <0.1000 | 0.27 |  |  |  |  | 1 |
|  |  |  |  | CBD (%DW) |  |  |  |  |  |  |  |  |
| CBD (%DW) | 503.10 | Other LC-UV | 118 | 1.31 | 1.274 | 1.281 | 1.29 | -1.53 | 1.5769 | 19 | 0.1884 | 0 |
| CBD (%DW) | 503.99 | Other | 154 | 1.2973 | 1.4445 | 1.3034 | 1.35 | -1.21 | 1.5769 | 19 | 0.1884 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 207 | 1.364 | 1.352 | 1.387 | 1.37 | -1.11 | 1.5769 | 19 | 0.1884 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 206 | 1.4107 | 1.4223 | 1.4506 | 1.43 | -0.79 | 1.5769 | 19 | 0.1884 | 0 |
| CBD (%DW) | 503.30 | Other LC-MS | 105 | 1.44 | 1.5 | 1.41 | 1.45 | -0.67 | 1.5769 | 19 | 0.1884 | 0 |
| CBD (%DW) | 503.02 | AOAC 2018.11, U | 199 | 1.3921 | 1.5308 | 1.4228 | 1.45 | -0.68 | 1.5769 | 19 | 0.1884 | 0 |
| CBD (%DW) | 503.60 | NIR | 187 | 1.4738 | 1.4322 | 1.4894 | 1.47 | -0.59 | 1.5769 | 19 | 0.1884 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 186 | 1.5171 | 1.4975 | 1.5214 | 1.51 | -0.34 | 1.5769 | 19 | 0.1884 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 194 | 1.573 | 1.57 | 1.608 | 1.58 | 0.04 | 1.5769 | 19 | 0.1884 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 196 | 1.607 | 1.612 | 1.583 | 1.6 | 0.13 | 1.5769 | 19 | 0.1884 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 204 | 1.6046 | 1.6064 | 1.6066 | 1.61 | 0.15 | 1.5769 | 19 | 0.1884 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 142 | 1.6258 | 1.5627 | 1.6743 | 1.62 | 0.23 | 1.5769 | 19 | 0.1884 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 198 | 1.6919 | 1.6168 | 1.5943 | 1.63 | 0.3 | 1.5769 | 19 | 0.1884 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 109 | 1.767 | 1.763 | 1.606 | 1.71 | 0.72 | 1.5769 | 19 | 0.1884 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 172 | 1.6447 | 1.8101 | 1.685 | 1.71 | 0.72 | 1.5769 | 19 | 0.1884 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 116 | 1.7001 | 1.7155 | 1.7546 | 1.72 | 0.78 | 1.5769 | 19 | 0.1884 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%DW) | 503.99 | Other | 183 | 1.7563 | 1.7007 | 1.7449 | 1.73 | 0.83 | 1.5769 | 19 | 0.1884 | 0 |
| CBD (%DW) | 503.30 | Other LC-MS | 201 | 1.63 | 1.94 | 2.07 | 1.88 | 1.61 | 1.5769 | 19 | 0.1884 | 0 |
| CBD (%DW) | 503.99 | Other | 172 | 1.8277 | 2.0115 | 1.8725 | 1.9 | 1.74 | 1.5769 | 19 | 0.1884 | 0 |
|  |  |  |  | CBDA (%DW) |  |  |  |  |  |  |  |  |
| CBDA (%DW) | 504.99 | Other | 154 | 2.196 | 1.9251 | 1.9749 | 2.03 | -2 | 3.3808 | 18 | 0.673 | 0 |
| CBDA (%DW) | 504.30 | Other LC-MS | 105 | 2.18 | 2.27 | 2.22 | 2.22 | -1.72 | 3.3808 | 18 | 0.673 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 142 | 2.517 | 2.5142 | 2.6041 | 2.55 | -1.24 | 3.3808 | 18 | 0.673 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 207 | 2.895 | 2.907 | 3.081 | 2.96 | -0.62 | 3.3808 | 18 | 0.673 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 118 | 3.146 | 3.049 | 2.954 | 3.05 | -0.49 | 3.3808 | 18 | 0.673 | 0 |
| CBDA (%DW) | 504.02 | AOAC 2018.11, U | 199 | 2.9183 | 3.3368 | 3.1047 | 3.12 | -0.39 | 3.3808 | 18 | 0.673 | 0 |
| CBDA (%DW) | 504.30 | Other LC-MS | 201 | 2.62 | 3.34 | 3.41 | 3.12 | -0.38 | 3.3808 | 18 | 0.673 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 198 | 3.529 | 3.1862 | 3.1602 | 3.29 | -0.13 | 3.3808 | 18 | 0.673 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 116 | 3.4451 | 3.4564 | 3.5768 | 3.49 | 0.17 | 3.3808 | 18 | 0.673 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 186 | 3.5367 | 3.5243 | 3.5677 | 3.54 | 0.24 | 3.3808 | 18 | 0.673 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 206 | 3.5426 | 3.5767 | 3.6391 | 3.59 | 0.31 | 3.3808 | 18 | 0.673 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 172 | 3.5793 | 3.7637 | 3.4616 | 3.6 | 0.33 | 3.3808 | 18 | 0.673 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 204 | 3.6182 | 3.6009 | 3.612 | 3.61 | 0.34 | 3.3808 | 18 | 0.673 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 196 | 3.638 | 3.653 | 3.596 | 3.63 | 0.37 | 3.3808 | 18 | 0.673 | 0 |
| CBDA (%DW) | 504.60 | NIR | 187 | 3.905 | 3.8126 | 3.6067 | 3.77 | 0.59 | 3.3808 | 18 | 0.673 | 0 |
| CBDA (%DW) | 504.99 | Other | 172 | 3.9776 | 4.1825 | 3.8468 | 4 | 0.92 | 3.3808 | 18 | 0.673 | 0 |
| CBDA (%DW) | 504.99 | Other | 183 | 4.4402 | 4.2799 | 4.461 | 4.39 | 1.51 | 3.3808 | 18 | 0.673 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDA (%DW) | 504.10 | Other LC-UV | 109 | 4.58 | 4.604 | 4.141 | 4.44 | 1.58 | 3.3808 | 18 | 0.673 | 0 |
|  |  |  |  | CBN (%DW) |  |  |  |  |  |  |  |  |
| CBN (%DW) | 505.10 | Other LC-UV | 109 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| CBN (%DW) | 505.99 | Other | 154 | 0.0209 | 0.0193 | 0.0192 | 0.0198 | -1.12 | 0.026 | 14 | 0.0055 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 192 | 0.0229 | 0.0221 | 0.0205 | 0.0218 | -0.75 | 0.026 | 14 | 0.0055 | 0 |
| CBN (%DW) | 505.02 | AOAC 2018.11, U | 199 | 0.0199 | 0.0231 | 0.022 | 0.0217 | -0.78 | 0.026 | 14 | 0.0055 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 207 | 0.022 | 0.023 | 0.023 | 0.0227 | -0.6 | 0.026 | 14 | 0.0055 | 0 |
| CBN (%DW) | 505.20 | Other LC-MS | 105 | 0.0226 | 0.0202 | 0.0253 | 0.0227 | -0.59 | 0.026 | 14 | 0.0055 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 206 | 0.024 | 0.0246 | 0.0221 | 0.0236 | -0.44 | 0.026 | 14 | 0.0055 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 186 | 0.024 | 0.0245 | 0.0257 | 0.0247 | -0.23 | 0.026 | 14 | 0.0055 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 142 | 0.0249 | 0.0253 | 0.0247 | 0.025 | -0.18 | 0.026 | 14 | 0.0055 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 196 | 0.0264 | 0.0273 | 0.0259 | 0.0265 | 0.1 | 0.026 | 14 | 0.0055 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 204 | 0.0258 | 0.0259 | 0.0257 | 0.0258 | -0.03 | 0.026 | 14 | 0.0055 | 0 |
| CBN (%DW) | 505.99 | Other | 183 | 0.0262 | 0.0248 | 0.0293 | 0.0268 | 0.14 | 0.026 | 14 | 0.0055 | 0 |
| CBN (%DW) | 505.30 | GC-FID | 116 | 0.0445 | 0.0405 | <0.0400 | 0.0417 | 2.83 | 0.026 | 14 | 0.0055 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 116 | 0.0424 | <0.0400 | 0.0447 | 0.0424 | 2.95 | 0.026 | 14 | 0.0055 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 198 | 0.0534 | 0.0499 | 0.0494 | 0.0509 | 4.49 | 0.026 | 14 | 0.0055 | 0 |
| CBN (%DW) | 505.20 | Other LC-MS | 201 | <0.1000 | <0.1000 | 0.11 | 0.103 |  |  |  |  | 1 |
|  |  |  |  | Total Δ9-THC (%DW) |  |  |  |  |  |  |  |  |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 207 | 0.095 | 0.086 | 0.093 | 0.0913 | -1.52 | 0.1426 | 25 | 0.0337 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%DW) | 506.30 | Other LC-MS | 105 | 0.11 | 0.102 | 0.102 | 0.105 | -1.12 | 0.1426 | 25 | 0.0337 | 0 |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, U | 103 | 0.106 | 0.1026 | 0.1057 | 0.105 | -1.12 | 0.1426 | 25 | 0.0337 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 142 | 0.0925 | 0.1109 | 0.1168 | 0.107 | -1.06 | 0.1426 | 25 | 0.0337 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 118 | 0.12 | 0.117 | 0.114 | 0.117 | -0.76 | 0.1426 | 25 | 0.0337 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 186 | 0.1237 | 0.1154 | 0.1217 | 0.12 | -0.66 | 0.1426 | 25 | 0.0337 | 0 |
| Total Δ9-THC (%DW) | 506.99 | Other | 154 | 0.1309 | 0.126 | 0.1178 | 0.125 | -0.52 | 0.1426 | 25 | 0.0337 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 192 | 0.1246 | 0.1223 | 0.1274 | 0.125 | -0.53 | 0.1426 | 25 | 0.0337 | 0 |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, U | 199 | 0.1254 | 0.1337 | 0.1191 | 0.126 | -0.49 | 0.1426 | 25 | 0.0337 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 198 | 0.132 | 0.1246 | 0.1218 | 0.126 | -0.49 | 0.1426 | 25 | 0.0337 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 206 | 0.1244 | 0.135 | 0.1271 | 0.129 | -0.41 | 0.1426 | 25 | 0.0337 | 0 |
| Total Δ9-THC (%DW) | 506.30 | Other LC-MS | 120 | 0.1335 | 0.13 | 0.1343 | 0.133 | -0.3 | 0.1426 | 25 | 0.0337 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 116 | 0.1414 | 0.1397 | 0.1312 | 0.137 | -0.15 | 0.1426 | 25 | 0.0337 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 109 | 0.14 | 0.141 | 0.156 | 0.146 | 0.09 | 0.1426 | 25 | 0.0337 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 138 | 0.15 | 0.1527 | 0.1376 | 0.147 | 0.12 | 0.1426 | 25 | 0.0337 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 204 | 0.1528 | 0.1495 | 0.1465 | 0.15 | 0.21 | 0.1426 | 25 | 0.0337 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 165 | 0.1586 | 0.1422 | 0.164 | 0.155 | 0.37 | 0.1426 | 25 | 0.0337 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 116 | 0.1603 | 0.151 | 0.16 | 0.157 | 0.43 | 0.1426 | 25 | 0.0337 | 0 |
| Total Δ9-THC (%DW) | 506.99 | Other | 183 | 0.1594 | 0.1533 | 0.1604 | 0.158 | 0.45 | 0.1426 | 25 | 0.0337 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 196 | 0.167 | 0.1679 | 0.1553 | 0.163 | 0.62 | 0.1426 | 25 | 0.0337 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 172 | 0.1854 | 0.1567 | 0.1913 | 0.178 | 1.04 | 0.1426 | 25 | 0.0337 | 0 |
| Total Δ9-THC (%DW) | 506.60 | NIR | 187 | 0.1825 | 0.1889 | 0.1861 | 0.186 | 1.28 | 0.1426 | 25 | 0.0337 | 0 |
| Total Δ9-THC (%DW) | 506.99 | Other | 172 | 0.206 | 0.1742 | 0.2127 | 0.198 | 1.63 | 0.1426 | 25 | 0.0337 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%DW) | 506.99 | Other | 181 | 0.275 | 0.287 | 0.243 | 0.268 | 3.73 | 0.1426 | 25 | 0.0337 | 0 |
| Total Δ9-THC (%DW) | 506.30 | Other LC-MS | 201 | 0.74 | 0.22 | 0.24 | 0.4 | 7.63 | 0.1426 | 25 | 0.0337 | 0 |
|  |  |  |  | Total CBD (%DW) |  |  |  |  |  |  |  |  |
| Total CBD (%DW) | 507.99 | Other | 154 | 3.3705 | 2.9856 | 3.0353 | 3.13 | -2.13 | 4.6456 | 21 | 0.7123 | 0 |
| Total CBD (%DW) | 507.30 | Other LC-MS | 105 | 3.35 | 3.49 | 3.36 | 3.4 | -1.75 | 4.6456 | 21 | 0.7123 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 142 | 3.8332 | 3.7677 | 3.9581 | 3.85 | -1.11 | 4.6456 | 21 | 0.7123 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 207 | 3.904 | 3.901 | 4.089 | 3.96 | -0.96 | 4.6456 | 21 | 0.7123 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 118 | 4.069 | 3.948 | 3.872 | 3.96 | -0.96 | 4.6456 | 21 | 0.7123 | 0 |
| Total CBD (%DW) | 507.02 | AOAC 2018.11, U | 199 | 3.9514 | 4.4527 | 4.1456 | 4.18 | -0.65 | 4.6456 | 21 | 0.7123 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 198 | 4.7868 | 4.4112 | 4.3657 | 4.52 | -0.17 | 4.6456 | 21 | 0.7123 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 206 | 4.5175 | 4.5591 | 4.6421 | 4.57 | -0.1 | 4.6456 | 21 | 0.7123 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 186 | 4.6188 | 4.5882 | 4.6503 | 4.62 | -0.04 | 4.6456 | 21 | 0.7123 | 0 |
| Total CBD (%DW) | 507.30 | Other LC-MS | 201 | 3.93 | 4.86 | 5.06 | 4.62 | -0.04 | 4.6456 | 21 | 0.7123 | 0 |
| Total CBD (%DW) | 507.60 | NIR | 187 | 4.8106 | 4.7752 | 4.6517 | 4.75 | 0.14 | 4.6456 | 21 | 0.7123 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 204 | 4.7777 | 4.7645 | 4.7744 | 4.77 | 0.18 | 4.6456 | 21 | 0.7123 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 196 | 4.798 | 4.816 | 4.737 | 4.78 | 0.19 | 4.6456 | 21 | 0.7123 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 116 | 4.7215 | 4.7467 | 4.8914 | 4.79 | 0.2 | 4.6456 | 21 | 0.7123 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 192 | 5.0343 | 4.7669 | 4.704 | 4.84 | 0.27 | 4.6456 | 21 | 0.7123 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 172 | 4.7837 | 5.1109 | 4.7208 | 4.87 | 0.32 | 4.6456 | 21 | 0.7123 | 0 |
| Total CBD (%DW) | 507.40 | GC-FID | 116 | 5.1295 | 4.9455 | 4.8997 | 4.99 | 0.49 | 4.6456 | 21 | 0.7123 | 0 |
| Total CBD (%DW) | 507.99 | Other | 172 | 5.3161 | 5.6796 | 5.2461 | 5.41 | 1.08 | 4.6456 | 21 | 0.7123 | 0 |

Population of Lab Values

Lab

Lab

Population of Lab Values

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%DW) | 507.99 | Other | 183 | 5.6504 | 5.4541 | 5.6571 | 5.59 | 1.32 | 4.6456 | 21 | 0.7123 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 109 | 5.783 | 5.8 | 5.238 | 5.61 | 1.35 | 4.6456 | 21 | 0.7123 | 0 |
| Total CBD (%DW) | 507.40 | GC-FID | 138 | 6.0433 | 7.1784 | 5.9181 | 6.38 | 2.43 | 4.6456 | 21 | 0.7123 | 0 |

|  |  |  |  | CBDV (%DW) |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDV (%DW) | 508.99 | Other | 109 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| CBDV (%DW) | 508.99 | Other | 207 | <0.0130 | <0.0130 | <0.0130 | <0.0130 | 1 |
| CBDV (%DW) | 508.99 | Other | 186 | <0.0200 | <0.0200 | <0.0200 | <0.0200 | 1 |
| CBDV (%DW) | 508.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 | <0.0205 | 1 |
| CBDV (%DW) | 508.99 | Other | 201 | <0.1000 | <0.1000 | <0.1000 | <0.1000 | 1 |
| CBDV (%DW) | 508.99 | Other | 206 | 0 | 0 | 0 | 0 | 2 |
| CBDV (%DW) | 508.99 | Other | 183 | 0 | 0 | 0 | 0 | 2 |
| CBDV (%DW) | 508.99 | Other | 142 | 0.0224 | 0.0267 | 0.0277 | 0.0256 | 0 |

|  |  |  |  | CBDVA (%DW) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDVA (%DW) | 509.99 | Other | 186 | 0.0202 | 0.0215 | 0.0285 | 0.0234 | -1.08 | 0.0467 | 6 | 0.0215 | 0 |
| CBDVA (%DW) | 509.99 | Other | 206 | 0.0359 | 0.036 | 0.0363 | 0.0361 | -0.49 | 0.0467 | 6 | 0.0215 | 0 |
| CBDVA (%DW) | 509.99 | Other | 196 | 0.0417 | 0.0424 | 0.0409 | 0.0417 | -0.23 | 0.0467 | 6 | 0.0215 | 0 |
| CBDVA (%DW) | 509.99 | Other | 154 | 0.0448 | 0.0444 | 0.0443 | 0.0445 | -0.1 | 0.0467 | 6 | 0.0215 | 0 |
| CBDVA (%DW) | 509.99 | Other | 183 | 0.0546 | 0.055 | 0.0567 | 0.0554 | 0.41 | 0.0467 | 6 | 0.0215 | 0 |
| CBDVA (%DW) | 509.99 | Other | 201 | <0.1000 | 0.45 | 0.45 | 0.333 | 13.31 | 0.0467 | 6 | 0.0215 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 |  | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | CBG (%DW) |  |  |  |  |  |  |  |  |  |  |  |  |
| CBG (%DW) | 510.99 | Other | 186 | <0.0200 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  |  | 1 |
| CBG (%DW) | 510.99 | Other | 109 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| CBG (%DW) | 510.99 | Other | 142 | 0.0347 | 0.0165 | 0.0179 | 0.023 |  | -1.08 | 0.0343 | 8 | 0.0104 | 0 |
| CBG (%DW) | 510.99 | Other | 196 | 0.0241 | 0.0233 | 0.0254 | 0.0243 |  | -0.96 | 0.0343 | 8 | 0.0104 | 0 |
| CBG (%DW) | 510.99 | Other | 206 | 0.0263 | 0.0272 | 0.0272 | 0.0269 |  | -0.71 | 0.0343 | 8 | 0.0104 | 0 |
| CBG (%DW) | 510.99 | Other | 183 | 0.0305 | 0.0342 | 0.0411 | 0.0353 |  | 0.1 | 0.0343 | 8 | 0.0104 | 0 |
| CBG (%DW) | 510.99 | Other | 207 | 0.035 | 0.028 | 0.045 | 0.036 |  | 0.17 | 0.0343 | 8 | 0.0104 | 0 |
| CBG (%DW) | 510.99 | Other | 204 | 0.0375 | 0.0358 | 0.0388 | 0.0374 |  | 0.3 | 0.0343 | 8 | 0.0104 | 0 |
| CBG (%DW) | 510.99 | Other | 105 | 0.0193 | 0.0526 | 0.0522 | 0.0414 |  | 0.69 | 0.0343 | 8 | 0.0104 | 0 |
| CBG (%DW) | 510.99 | Other | 201 | 0.16 | 0.18 | 0.18 | 0.173 |  | 13.4 | 0.0343 | 8 | 0.0104 | 0 |

Population of Lab Values

|  |  |  |  | CBGA (%DW) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%DW) | 511.99 | Other | 109 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| CBGA (%DW) | 511.99 | Other | 206 | 0 | 0 | 0 | 0 |  |  |  |  | 2 |
| CBGA (%DW) | 511.99 | Other | 207 | 0.028 | 0.036 | 0.038 | 0.034 | -1.08 | 0.0626 | 8 | 0.0265 | 0 |
| CBGA (%DW) | 511.99 | Other | 196 | 0.0459 | 0.0454 | 0.045 | 0.0454 | -0.65 | 0.0626 | 8 | 0.0265 | 0 |
| CBGA (%DW) | 511.99 | Other | 142 | 0.0555 | 0.0505 | 0.0537 | 0.0532 | -0.35 | 0.0626 | 8 | 0.0265 | 0 |
| CBGA (%DW) | 511.99 | Other | 186 | 0.0579 | 0.0558 | 0.0514 | 0.055 | -0.28 | 0.0626 | 8 | 0.0265 | 0 |
| CBGA (%DW) | 511.99 | Other | 183 | 0.0531 | 0.0526 | 0.0679 | 0.0579 | -0.18 | 0.0626 | 8 | 0.0265 | 0 |
| CBGA (%DW) | 511.99 | Other | 204 | 0.0602 | 0.0586 | 0.0585 | 0.0591 | -0.13 | 0.0626 | 8 | 0.0265 | 0 |
| CBGA (%DW) | 511.99 | Other | 154 | 0.0938 | 0.0932 | 0.0934 | 0.0935 | 1.17 | 0.0626 | 8 | 0.0265 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%DW) | 511.99 | Other | 201 | <0.1000 | 0.44 | 0.45 | 0.33 | 10.1 | 0.0626 | 8 | 0.0265 | 0 |
|  |  |  |  | THCV (%DW) |  |  |  |  |  |  |  |  |
| THCV (%DW) | 512.99 | Other | 142 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| THCV (%DW) | 512.99 | Other | 109 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| THCV (%DW) | 512.99 | Other | 186 | <0.0200 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  | 1 |
| THCV (%DW) | 512.99 | Other | 196 | <0.2050 | <0.0205 | <0.0205 | <0.2050 |  |  |  |  | 1 |
| THCV (%DW) | 512.99 | Other | 183 | 0 | 0 | 0 | 0 |  |  |  |  | 2 |
| THCV (%DW) | 512.99 | Other | 206 | 0.0016 | 0.0026 | 0.0021 | 0.0021 |  |  |  |  | 0 |
| THCV (%DW) | 512.99 | Other | 201 | <0.1000 | <0.1000 | 0.1 | 0.1 |  |  |  |  | 1 |
|  |  |  |  | Δ8-THC (%DW) |  |  |  |  |  |  |  |  |
| Δ8-THC (%DW) | 513.99 | Other | 109 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Δ8-THC (%DW) | 513.99 | Other | 142 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Δ8-THC (%DW) | 513.99 | Other | 207 | <0.0130 | <0.0130 | <0.0130 | <0.0130 |  |  |  |  | 1 |
| Δ8-THC (%DW) | 513.99 | Other | 186 | <0.0200 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  | 1 |
| Δ8-THC (%DW) | 513.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 | <0.0205 |  |  |  |  | 1 |
| Δ8-THC (%DW) | 513.99 | Other | 183 | 0 | 0 | 0 | 0 |  |  |  |  | 2 |
| Δ8-THC (%DW) | 513.99 | Other | 206 | 0 | 0 | 0 | 0 |  |  |  |  | 2 |
|  |  |  |  | CBC (%DW) |  |  |  |  |  |  |  |  |
| CBC (%DW) | 514.99 | Other | 206 | 0.1113 | 0.1158 | 0.114 | 0.114 | -0.9 | 0.1263 | 12 | 0.0139 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%DW) | 514.99 | Other | 207 | 0.115 | 0.111 | 0.118 | 0.115 | -0.83 | 0.1263 | 12 | 0.0139 | 0 |
| CBC (%DW) | 514.99 | Other | 186 | 0.1144 | 0.1148 | 0.1148 | 0.115 | -0.83 | 0.1263 | 12 | 0.0139 | 0 |
| CBC (%DW) | 514.99 | Other | 154 | 0.1113 | 0.1129 | 0.1224 | 0.116 | -0.77 | 0.1263 | 12 | 0.0139 | 0 |
| CBC (%DW) | 514.99 | Other | 194 | 0.117 | 0.117 | 0.122 | 0.119 | -0.55 | 0.1263 | 12 | 0.0139 | 0 |
| CBC (%DW) | 514.99 | Other | 142 | 0.1267 | 0.1161 | 0.1216 | 0.121 | -0.34 | 0.1263 | 12 | 0.0139 | 0 |
| CBC (%DW) | 514.99 | Other | 105 | 0.125 | 0.117 | 0.125 | 0.122 | -0.28 | 0.1263 | 12 | 0.0139 | 0 |
| CBC (%DW) | 514.99 | Other | 204 | 0.1281 | 0.1289 | 0.1283 | 0.128 | 0.16 | 0.1263 | 12 | 0.0139 | 0 |
| CBC (%DW) | 514.99 | Other | 196 | 0.1385 | 0.1352 | 0.1345 | 0.136 | 0.7 | 0.1263 | 12 | 0.0139 | 0 |
| CBC (%DW) | 514.99 | Other | 183 | 0.136 | 0.1311 | 0.1443 | 0.137 | 0.78 | 0.1263 | 12 | 0.0139 | 0 |
| CBC (%DW) | 514.99 | Other | 109 | 0.149 | 0.15 | 0.137 | 0.145 | 1.37 | 0.1263 | 12 | 0.0139 | 0 |
| CBC (%DW) | 514.99 | Other | 201 | 0.24 | 0.37 | 0.44 | 0.35 | 16.08 | 0.1263 | 12 | 0.0139 | 0 |

Population of Lab Values

Cd (ug/g DW)

Cd (ug/g DW)

601.04	ICP-MS

105

0.206	0.195	0.205	0.202	0

As (ug/g DW)

As (ug/g DW)

602.04	ICP-MS

105

0.683	0.669	0.729	0.694	0

Hg (ug/g DW)

Hg (ug/g DW)

603.04	ICP-MS

105

0.013	0.0128	0.0119	0.0126	0

Pb (ug/g DW)

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

Value	Z score

Population of Lab Values	 Rob Mean	n	Rob StDev

Flag

Pb (ug/g DW)

604.04	ICP-MS

105

2.2	2.93	2.2	2.44	0

Cu (ug/g DW)

Cu (ug/g DW)

605.04	ICP-MS

105

15	14.1	13.4	14.2	0

Ni (ug/g DW)

Ni (ug/g DW)

606.04	ICP-MS

105

9.02	9.59	10.9	9.84	0

Cr (ug/g DW)

Cr (ug/g DW)

607.04	ICP-MS

105

17.8	18.8	18.5	18.4	0

Se (ug/g DW)

Se (ug/g DW)

608.04	ICP-MS

105

0.0915	0.102	0.126	0.106	0

Zn (ug/g DW)

Zn (ug/g DW)

611.04	ICP-MS

105

62.1	69	70.2	67.1	0

- Flag definitions: 1 (Lab Value not used because there were less than 2 numeric lab results), 2 (Lab Value not used because there were less than 2 numeric lab results greater than 0).

- Robust statistics performed to determine robust mean and Z scores only if there were 6 or more numeric lab values.

- AR refers to concentration on as-received basis. DW refers to concentration on dry weight basis.