Hemp Proficiency Testing

All Labs Precision HM20SEP-2

Sample: HM20SEP-2

All Labs Precision

Issue Date:

November 2, 2020

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

Δ9-THC (%AR)

RSDr	HorRat(r)

Population of Lab RSDr	 Rob Mean	min	max

Flag

| Δ9-THC (%AR) | 001.10 | Other LC-UV | 178 | 0.2300 |  |  |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.60 | NIR | 187 | 0.1751 | 0.1747 | 0.1751 | 0.13 | 0.03 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 123 | 0.2560 | 0.2550 | 0.2540 | 0.39 | 0.08 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 156 | 0.3320 | 0.3289 | 0.3297 | 0.49 | 0.1 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.99 | Other | 137 | 0.2080 | 0.2070 | 0.2090 | 0.48 | 0.1 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 162 | 0.2436 | 0.2407 | 0.2438 | 0.72 | 0.14 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 202 | 0.1886 | 0.1915 | 0.1898 | 0.77 | 0.15 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 124 | 0.2069 | 0.2036 | 0.2047 | 0.82 | 0.16 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.03 | AOAC 2018.11, M | 139 | 0.2296 | 0.2274 | 0.2313 | 0.85 | 0.17 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 149 | 0.2470 | 0.2440 | 0.2430 | 0.85 | 0.17 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 189 | 0.2676 | 0.2627 | 0.2651 | 0.92 | 0.19 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 159 | 0.1920 | 0.1960 | 0.1950 | 1.07 | 0.21 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 204 | 0.2420 | 0.2377 | 0.2429 | 1.15 | 0.23 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 199 | 0.2515 | 0.2495 | 0.2455 | 1.23 | 0.25 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 122 | 0.2775 | 0.2732 | 0.2703 | 1.32 | 0.27 | 3.31 | 0.13 | 16.2 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 186 | 0.2355 | 0.2300 | 0.2304 | 1.32 | 0.27 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 148 | 0.2040 | 0.2070 | 0.2100 | 1.45 | 0.29 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 150 | 0.2333 | 0.2267 | 0.2267 | 1.66 | 0.33 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 164 | 0.2192 | 0.2247 | 0.2271 | 1.81 | 0.36 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 168 | 0.2125 | 0.2114 | 0.2053 | 1.85 | 0.37 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 161 | 0.1904 | 0.1977 | 0.1922 | 1.97 | 0.39 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 152 | 0.2429 | 0.2495 | 0.2524 | 1.96 | 0.4 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 200 | 0.2281 | 0.2195 | 0.2211 | 2.05 | 0.41 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 196 | 0.1943 | 0.2026 | 0.1974 | 2.12 | 0.42 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 144 | 0.2202 | 0.2249 | 0.2299 | 2.16 | 0.43 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 165 | 0.2381 | 0.2477 | 0.2391 | 2.18 | 0.44 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 103 | 0.2109 | 0.2023 | 0.2094 | 2.21 | 0.44 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 113 | 0.2543 | 0.2461 | 0.2433 | 2.31 | 0.47 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 155 | 0.1960 | 0.2040 | 0.2050 | 2.45 | 0.48 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.99 | Other | 154 | 0.2206 | 0.2108 | 0.2195 | 2.47 | 0.49 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 116 | 0.2593 | 0.2632 | 0.2508 | 2.46 | 0.5 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 172 | 0.2610 | 0.2720 | 0.2742 | 2.63 | 0.54 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 141 | 0.1921 | 0.1825 | 0.1834 | 2.85 | 0.56 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 163 | 0.2450 | 0.2590 | 0.2520 | 2.78 | 0.57 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 182 | 0.2382 | 0.2346 | 0.2253 | 2.86 | 0.58 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 120 | 0.2385 | 0.2354 | 0.2487 | 2.89 | 0.58 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 195 | 0.2083 | 0.2002 | 0.2125 | 3.02 | 0.6 | 3.31 | 0.13 | 16.2 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 184 | 0.2180 | 0.2316 | 0.2306 | 3.34 | 0.67 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 108 | 0.2408 | 0.2261 | 0.2404 | 3.55 | 0.72 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 125 | 0.2303 | 0.2357 | 0.2474 | 3.68 | 0.74 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 142 | 0.2376 | 0.2206 | 0.2262 | 3.8 | 0.76 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 149 | 0.2400 | 0.2500 | 0.2300 | 4.17 | 0.84 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 198 | 0.2458 | 0.2263 | 0.2420 | 4.34 | 0.88 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 191 | 0.2048 | 0.2256 | 0.2159 | 4.83 | 0.96 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 101 | 0.2500 | 0.2260 | 0.2400 | 5.05 | 1.02 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 119 | 0.2564 | 0.2415 | 0.2671 | 5.04 | 1.03 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 109 | 0.2270 | 0.2510 | 0.2300 | 5.54 | 1.12 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 175 | 0.2376 | 0.2389 | 0.2158 | 5.62 | 1.13 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 197 | 0.2190 | 0.2250 | 0.2000 | 6.08 | 1.21 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 192 | 0.1906 | 0.2055 | 0.1819 | 6.19 | 1.21 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 106 | 0.2600 | 0.2700 | 0.2400 | 5.95 | 1.22 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 102 | 0.2595 | 0.2321 | 0.2362 | 6.09 | 1.23 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 203 | 0.2120 | 0.1880 | 0.2080 | 6.34 | 1.25 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 131 | 0.2300 | 0.2200 | 0.2500 | 6.55 | 1.32 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 201 | 0.8400 | 0.8000 | 0.7500 | 5.66 | 1.37 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 145 | 0.2200 | 0.2300 | 0.2000 | 7.05 | 1.4 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 132 | 0.2809 | 0.2515 | 0.2451 | 7.37 | 1.51 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 188 | 0.2159 | 0.2107 | 0.2430 | 7.77 | 1.55 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.99 | Other | 132 | 0.3400 | 0.3000 | 0.3000 | 7.37 | 1.55 | 3.31 | 0.13 | 16.2 | 0 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 115 | 0.2500 | 0.2500 | 0.2900 | 8.77 | 1.8 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.99 | Other | 205 | 0.2140 | 0.2440 | 0.2060 | 9.05 | 1.81 | 3.31 | 0.13 | 16.2 | 0 |
| Δ9-THC (%AR) | 001.99 | Other | 172 | 0.2031 | 0.2175 | 0.2741 | 16.2 | 3.26 | 3.31 | 0.13 | 16.2 | 0 |

Lab

Population of Lab RSDr

Δ9-THCA (%AR)

| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 194 | <0.0724 | <0.0724 | <0.0724 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 178 | 0.1100 |  |  |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 123 | <0.0880 | <0.0930 | <0.0880 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 188 | 0.0603 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 197 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 131 | 0.0600 | 0.0600 | 0.0600 | 0 | 0 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 106 | 0.0600 | 0.0600 | 0.0600 | 0 | 0 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 163 | 0.0990 | 0.0980 | 0.0990 | 0.58 | 0.1 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 204 | 0.0683 | 0.0690 | 0.0693 | 0.74 | 0.13 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 122 | 0.0620 | 0.0643 | 0.0630 | 1.83 | 0.3 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 184 | 0.0493 | 0.0511 | 0.0511 | 2.06 | 0.33 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 159 | 0.0667 | 0.0692 | 0.0690 | 2.03 | 0.34 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 196 | 0.0466 | 0.0473 | 0.0486 | 2.14 | 0.34 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 162 | 0.0494 | 0.0518 | 0.0503 | 2.4 | 0.38 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 116 | 0.0609 | 0.0633 | 0.0636 | 2.36 | 0.39 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 150 | 0.0733 | 0.0700 | 0.0733 | 2.64 | 0.45 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.99 | Other | 154 | 0.0854 | 0.0900 | 0.0863 | 2.79 | 0.49 | 6.92 | 0 | 173 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 186 | 0.0585 | 0.0555 | 0.0555 | 3.07 | 0.5 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.99 | Other | 137 | 0.0513 | 0.0541 | 0.0513 | 3.09 | 0.5 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 142 | 0.0514 | 0.0485 | 0.0489 | 3.17 | 0.51 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 200 | 0.0466 | 0.0446 | 0.0434 | 3.6 | 0.57 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 101 | 0.0620 | 0.0590 | 0.0580 | 3.49 | 0.57 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 156 | 0.0686 | 0.0639 | 0.0661 | 3.55 | 0.59 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 199 | 0.0530 | 0.0570 | 0.0565 | 3.93 | 0.64 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 109 | 0.0690 | 0.0650 | 0.0640 | 4.01 | 0.67 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 195 | 0.0579 | 0.0537 | 0.0539 | 4.29 | 0.7 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 113 | 0.0476 | 0.0445 | 0.0435 | 4.73 | 0.74 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 165 | 0.0706 | 0.0652 | 0.0655 | 4.52 | 0.76 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 149 | 0.0640 | 0.0640 | 0.0590 | 4.63 | 0.77 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 144 | 0.0522 | 0.0573 | 0.0571 | 5.2 | 0.84 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 161 | 0.0514 | 0.0545 | 0.0572 | 5.34 | 0.86 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 141 | 0.0452 | 0.0411 | 0.0454 | 5.53 | 0.87 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 119 | 0.0640 | 0.0589 | 0.0654 | 5.45 | 0.9 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 168 | 0.0523 | 0.0570 | 0.0589 | 6.06 | 0.99 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 164 | 0.0641 | 0.0628 | 0.0571 | 6.07 | 1 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 103 | 0.0627 | 0.0712 | 0.0671 | 6.34 | 1.06 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 189 | 0.0881 | 0.0778 | 0.0849 | 6.31 | 1.09 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 191 | 0.0659 | 0.0763 | 0.0680 | 7.85 | 1.32 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 149 | 0.0660 | 0.0630 | 0.0560 | 8.32 | 1.37 | 6.92 | 0 | 173 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.03 | AOAC 2018.11, M | 139 | 0.0778 | 0.0662 | 0.0746 | 8.22 | 1.39 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 124 | 0.0415 | 0.0458 | 0.0498 | 9.08 | 1.43 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 145 | 0.0700 | 0.0700 | 0.0600 | 8.66 | 1.45 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.60 | NIR | 187 | 0.0319 | 0.0296 | 0.0262 | 9.81 | 1.45 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 115 | 0.0530 | 0.0450 | 0.0440 | 10.4 | 1.65 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 108 | 0.0832 | 0.0685 | 0.0779 | 9.73 | 1.66 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 152 | 0.0438 | 0.0457 | 0.0533 | 10.6 | 1.68 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 120 | 0.0658 | 0.0655 | 0.0778 | 10.1 | 1.69 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 125 | 0.0981 | 0.0967 | 0.1158 | 10.3 | 1.83 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.99 | Other | 205 | 0.0320 | 0.0380 | 0.0410 | 12.4 | 1.89 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 155 | 0.0810 | 0.0700 | 0.0640 | 12 | 2.03 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 198 | 0.0447 | 0.0392 | 0.0525 | 14.7 | 2.32 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 192 | 0.0497 | 0.0405 | 0.0558 | 15.8 | 2.52 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.99 | Other | 132 | 0.0400 | 0.0300 | 0.0300 | 17.3 | 2.61 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 172 | 0.0590 | 0.0520 | 0.0723 | 16.9 | 2.78 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.99 | Other | 172 | 0.0786 | 0.0579 | 0.0805 | 17.3 | 2.93 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 182 | 0.1029 | 0.0573 | 0.0574 | 36.3 | 6.13 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 175 | 0.0665 | 0.0695 | 0.0306 | 39 | 6.33 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 201 | 0.0000 | 0.0000 | 0.0400 | 173 | 22.71 | 6.92 | 0 | 173 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 202 | 0.5110 | 0.0531 | 0.0521 | 129 | 25.47 | 6.92 | 0 | 173 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | CBD (%AR) |  |  |  |  |  |  |  |  |
| CBD (%AR) | 003.10 | Other LC-UV | 178 | 3.2000 |  |  |  |  |  |  |  | 1 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 199 | 3.6635 | 3.6695 | 3.6725 | 0.12 | 0.04 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 125 | 4.2051 | 4.2302 | 4.2021 | 0.37 | 0.11 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 161 | 3.1836 | 3.1961 | 3.1737 | 0.35 | 0.11 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.60 | NIR | 187 | 2.7287 | 2.7371 | 2.7494 | 0.38 | 0.11 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 150 | 3.8333 | 3.8000 | 3.8000 | 0.5 | 0.15 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 123 | 3.8200 | 3.8230 | 3.7830 | 0.58 | 0.18 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 204 | 3.6490 | 3.6084 | 3.6195 | 0.58 | 0.18 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 163 | 4.1980 | 4.2610 | 4.2300 | 0.74 | 0.23 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.03 | AOAC 2018.11, M | 139 | 3.2825 | 3.2977 | 3.2426 | 0.87 | 0.26 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.99 | Other | 137 | 3.6100 | 3.6200 | 3.6700 | 0.88 | 0.27 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 142 | 3.7949 | 3.7221 | 3.7620 | 0.97 | 0.3 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 189 | 3.5790 | 3.5026 | 3.5333 | 1.09 | 0.33 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.01 | AOAC 2018.10 | 131 | 4.0600 | 4.1100 | 4.0200 | 1.11 | 0.34 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 109 | 3.8630 | 3.8420 | 3.7810 | 1.11 | 0.34 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 159 | 3.0900 | 3.0600 | 3.0200 | 1.15 | 0.34 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 164 | 3.4781 | 3.4049 | 3.3888 | 1.39 | 0.42 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 145 | 3.7700 | 3.6800 | 3.6800 | 1.4 | 0.43 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.01 | AOAC 2018.10 | 155 | 3.1240 | 3.1400 | 3.0490 | 1.57 | 0.46 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 156 | 4.1901 | 4.3227 | 4.2764 | 1.58 | 0.49 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 195 | 3.5005 | 3.3930 | 3.4122 | 1.67 | 0.5 | 2.64 | 0.12 | 26.1 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%AR) | 003.99 | Other | 205 | 3.7230 | 3.6210 | 3.6020 | 1.78 | 0.54 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 113 | 4.0605 | 3.9744 | 3.9210 | 1.77 | 0.54 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 197 | 3.8220 | 3.7530 | 3.6900 | 1.76 | 0.54 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 108 | 3.4020 | 3.3370 | 3.5040 | 2.47 | 0.74 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.01 | AOAC 2018.10 | 168 | 3.5964 | 3.4240 | 3.5118 | 2.46 | 0.74 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 144 | 4.0247 | 3.9095 | 3.8345 | 2.44 | 0.75 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 196 | 3.2790 | 3.4470 | 3.4020 | 2.58 | 0.77 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 200 | 3.4192 | 3.3300 | 3.2463 | 2.6 | 0.78 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 116 | 3.6426 | 3.8493 | 3.7603 | 2.76 | 0.84 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 152 | 4.1251 | 4.0107 | 3.8948 | 2.87 | 0.89 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.99 | Other | 154 | 3.2624 | 3.0651 | 3.1664 | 3.12 | 0.93 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 172 | 4.0130 | 4.1930 | 4.2653 | 3.13 | 0.97 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 141 | 3.0403 | 2.8579 | 2.8901 | 3.32 | 0.98 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 203 | 3.4340 | 3.6040 | 3.6550 | 3.25 | 0.98 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 122 | 3.9699 | 3.7482 | 3.7661 | 3.22 | 0.99 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 184 | 3.7340 | 3.9111 | 4.0008 | 3.5 | 1.07 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 192 | 3.1801 | 3.3440 | 3.1225 | 3.57 | 1.07 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 198 | 3.4076 | 3.1881 | 3.3950 | 3.7 | 1.11 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 188 | 3.8670 | 3.8480 | 4.1240 | 3.91 | 1.2 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 107 | 0.2501 | 0.2412 | 0.2705 | 5.92 | 1.21 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.99 | Other | 172 | 3.6861 | 3.9618 | 3.9496 | 4.03 | 1.24 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 106 | 4.3900 | 4.4900 | 4.1300 | 4.29 | 1.34 | 2.64 | 0.12 | 26.1 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%AR) | 003.10 | Other LC-UV | 132 | 4.5902 | 4.4434 | 4.2065 | 4.39 | 1.37 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 149 | 3.7000 | 4.0000 | 3.7000 | 4.56 | 1.4 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 202 | 3.0409 | 3.0377 | 3.3018 | 4.85 | 1.44 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 162 | 3.8327 | 3.8275 | 4.1654 | 4.91 | 1.51 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 119 | 3.8013 | 3.6042 | 3.9904 | 5.08 | 1.56 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 191 | 3.3508 | 3.6992 | 3.6417 | 5.24 | 1.59 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.99 | Other | 132 | 4.1000 | 3.7500 | 3.7300 | 5.39 | 1.65 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 149 | 4.3040 | 4.1040 | 3.7230 | 7.3 | 2.26 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 201 | 11.0900 | 11.9400 | 14.0600 | 12.4 | 4.52 | 2.64 | 0.12 | 26.1 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 186 | 3.6671 | 3.6386 | 5.5988 | 26.1 | 8.15 | 2.64 | 0.12 | 26.1 | 0 |
|  |  |  |  | CBDA (%AR) |  |  |  |  |  |  |  |  |
| CBDA (%AR) | 004.10 | Other LC-UV | 178 | 4.0500 |  |  |  |  |  |  |  | 1 |
| CBDA (%AR) | 004.60 | NIR | 187 | 3.1376 | 3.1407 | 3.1404 | 0.05 | 0.02 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.99 | Other | 137 | 3.7500 | 3.7400 | 3.7600 | 0.27 | 0.08 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 163 | 3.2270 | 3.2440 | 3.2360 | 0.26 | 0.08 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 204 | 3.8889 | 3.8639 | 3.8872 | 0.36 | 0.11 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 123 | 4.0500 | 4.0130 | 4.0060 | 0.59 | 0.18 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 186 | 3.9306 | 3.9559 | 3.8988 | 0.73 | 0.22 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 156 | 3.7549 | 3.8098 | 3.7928 | 0.74 | 0.23 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 161 | 3.0788 | 3.0307 | 3.0672 | 0.82 | 0.24 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.03 | AOAC 2018.11, M | 139 | 3.5492 | 3.5014 | 3.5559 | 0.84 | 0.25 | 2.65 | 0.05 | 10.6 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDA (%AR) | 004.30 | Other LC-MS | 159 | 3.9400 | 4.0000 | 3.9300 | 0.96 | 0.29 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 199 | 3.9905 | 3.9935 | 3.9250 | 0.98 | 0.3 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.99 | Other | 205 | 3.5730 | 3.5480 | 3.5020 | 1.02 | 0.31 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 150 | 3.5667 | 3.6000 | 3.5000 | 1.43 | 0.43 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 109 | 4.6700 | 4.6560 | 4.5530 | 1.38 | 0.44 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.01 | AOAC 2018.10 | 131 | 4.0500 | 4.1700 | 4.1100 | 1.46 | 0.45 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 203 | 3.8770 | 3.9290 | 3.8080 | 1.57 | 0.48 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 184 | 3.9764 | 3.9740 | 4.0858 | 1.59 | 0.49 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.01 | AOAC 2018.10 | 168 | 3.6575 | 3.6318 | 3.7499 | 1.69 | 0.51 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 189 | 3.7134 | 3.6021 | 3.6082 | 1.72 | 0.52 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 108 | 3.7010 | 3.6110 | 3.7340 | 1.73 | 0.53 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.99 | Other | 172 | 3.9720 | 4.0980 | 4.0887 | 1.73 | 0.54 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 116 | 3.6183 | 3.7541 | 3.7223 | 1.92 | 0.59 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 196 | 3.2860 | 3.4380 | 3.3200 | 2.38 | 0.72 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 164 | 4.2535 | 4.0841 | 4.2576 | 2.36 | 0.73 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 197 | 4.0810 | 4.0530 | 3.9040 | 2.37 | 0.73 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 152 | 3.7841 | 3.6090 | 3.7321 | 2.42 | 0.74 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 149 | 3.9000 | 4.1000 | 4.0000 | 2.5 | 0.77 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 195 | 3.8041 | 3.6827 | 3.6141 | 2.6 | 0.79 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 122 | 4.1084 | 4.2731 | 4.3201 | 2.63 | 0.82 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 200 | 3.4098 | 3.3209 | 3.2179 | 2.9 | 0.87 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 188 | 4.1550 | 4.0080 | 3.9260 | 2.88 | 0.89 | 2.65 | 0.05 | 10.6 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDA (%AR) | 004.10 | Other LC-UV | 145 | 4.2900 | 4.0900 | 4.0400 | 3.2 | 0.99 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 144 | 4.0323 | 3.8998 | 3.7750 | 3.3 | 1.01 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.99 | Other | 154 | 3.2884 | 3.0631 | 3.1303 | 3.66 | 1.09 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 172 | 4.1790 | 4.4170 | 4.4771 | 3.62 | 1.13 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 162 | 3.7664 | 3.7637 | 4.0195 | 3.82 | 1.17 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 194 | 3.9402 | 4.0925 | 3.7848 | 3.91 | 1.2 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 141 | 3.4566 | 3.1745 | 3.2789 | 4.32 | 1.29 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.01 | AOAC 2018.10 | 155 | 3.1390 | 3.4090 | 3.3560 | 4.33 | 1.3 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 198 | 3.9916 | 3.6846 | 3.9602 | 4.35 | 1.34 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 202 | 3.3537 | 3.4299 | 3.6506 | 4.43 | 1.34 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 125 | 4.1920 | 4.5759 | 4.4002 | 4.38 | 1.37 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 142 | 3.8482 | 3.5363 | 3.5354 | 4.95 | 1.51 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 201 | 17.6100 | 16.6600 | 18.1400 | 4.29 | 1.65 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 191 | 3.8391 | 4.2637 | 4.2073 | 5.62 | 1.74 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 192 | 2.7492 | 2.4381 | 2.5521 | 6.1 | 1.76 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 119 | 3.8389 | 3.3704 | 3.7164 | 6.67 | 2.03 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 149 | 4.3400 | 4.2940 | 3.8420 | 6.62 | 2.05 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.99 | Other | 132 | 3.9300 | 3.4600 | 3.3800 | 8.28 | 2.51 | 2.65 | 0.05 | 10.6 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 106 | 4.4900 | 4.6100 | 3.7700 | 10.6 | 3.3 | 2.65 | 0.05 | 10.6 | 0 |
|  |  |  |  | CBN (%AR) |  |  |  |  |  |  |  |  |
| CBN (%AR) | 005.10 | Other LC-UV | 123 | <0.0880 <0.0930 |  | <0.0880 |  |  |  |  |  | 1 |

Population of Lab RSDr

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%AR) | 005.10 | Other LC-UV | 203 | <0.0040 | 0.0200 | 0.0180 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 108 | <0.0030 | <0.0030 | <0.0030 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.20 | Other LC-MS | 188 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 194 | <0.0239 | <0.0724 | <0.0239 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 178 | 0.0300 |  |  |  |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 116 | 0.0491 | <0.0400 | 0.0410 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.99 | Other | 137 | <0.0290 | <0.0295 | <0.0293 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 109 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.30 | GC-FID | 185 | <0.0100 | <0.1000 | <0.1000 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 199 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 197 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.01 | AOAC 2018.10 | 131 | 0.0300 | 0.0300 | 0.0300 | 0 | 0 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 145 | 0.0200 | 0.0200 | 0.0200 | 0 | 0 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 204 | 0.0280 | 0.0278 | 0.0281 | 0.55 | 0.08 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 189 | 0.0271 | 0.0267 | 0.0272 | 0.98 | 0.14 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.20 | Other LC-MS | 122 | 0.0257 | 0.0261 | 0.0255 | 1.19 | 0.17 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 163 | 0.0480 | 0.0490 | 0.0490 | 1.19 | 0.19 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.20 | Other LC-MS | 159 | 0.0241 | 0.0249 | 0.0240 | 2.03 | 0.29 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 196 | 0.0205 | 0.0210 | 0.0214 | 2.15 | 0.3 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 202 | 0.0252 | 0.0263 | 0.0257 | 2.14 | 0.31 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.99 | Other | 154 | 0.0306 | 0.0298 | 0.0311 | 2.15 | 0.32 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 186 | 0.0295 | 0.0285 | 0.0280 | 2.66 | 0.39 | 5.29 | 0 | 43.3 | 0 |

Lab

Lab

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%AR) | 005.10 | Other LC-UV | 142 | 0.0256 | 0.0251 | 0.0265 | 2.76 | 0.4 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.01 | AOAC 2018.10 | 155 | 0.0340 | 0.0350 | 0.0330 | 2.94 | 0.44 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 116 | 0.0475 | 0.0463 | 0.0493 | 3.17 | 0.5 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 200 | 0.0280 | 0.0260 | 0.0275 | 3.83 | 0.56 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.01 | AOAC 2018.10 | 168 | 0.0283 | 0.0279 | 0.0301 | 4.07 | 0.6 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.20 | Other LC-MS | 119 | 0.0303 | 0.0298 | 0.0322 | 4.12 | 0.61 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 144 | 0.0231 | 0.0241 | 0.0252 | 4.35 | 0.62 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 195 | 0.0267 | 0.0246 | 0.0250 | 4.38 | 0.63 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.20 | Other LC-MS | 156 | 0.0327 | 0.0359 | 0.0352 | 4.86 | 0.74 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 132 | 0.0246 | 0.0226 | 0.0223 | 5.4 | 0.77 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 192 | 0.0299 | 0.0332 | 0.0306 | 5.57 | 0.83 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 198 | 0.0837 | 0.0760 | 0.0824 | 5.11 | 0.88 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 162 | 0.0314 | 0.0276 | 0.0289 | 6.59 | 0.97 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 125 | 0.0338 | 0.0344 | 0.0297 | 7.84 | 1.18 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.03 | AOAC 2018.11, M | 139 | 0.0212 | 0.0240 | 0.0250 | 8.42 | 1.2 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.20 | Other LC-MS | 149 | 0.0300 | 0.0290 | 0.0250 | 9.45 | 1.38 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 141 | 0.0191 | 0.0179 | 0.0226 | 12.3 | 1.71 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 138 | 0.0304 | 0.0378 | 0.0384 | 12.5 | 1.9 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 149 | 0.0350 | 0.0340 | 0.0270 | 13.6 | 2.04 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.99 | Other | 205 | 0.0260 | 0.0270 | 0.0200 | 15.6 | 2.23 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 107 | 6.4749 | 7.0620 | 7.7840 | 9.23 | 3.1 | 5.29 | 0 | 43.3 | 0 |
| CBN (%AR) | 005.99 | Other | 132 | 0.0300 | 0.0200 | 0.0200 | 24.7 | 3.53 | 5.29 | 0 | 43.3 | 0 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%AR) | 005.10 | Other LC-UV | 106 | 0.0400 | 0.0200 | 0.0200 | 43.3 | 6.3 | 5.29 | 0 | 43.3 | 0 |

Lab

Population of Lab RSDr

Total Δ9-THC (%AR)

| Total Δ9-THC (%AR) | 006.40 | GC-FID | 185 | <0.0800 | <0.1000 | <0.1000 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 178 | 0.3400 |  |  |  |  |  |  |  | 1 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 162 | 0.2870 | 0.2860 | 0.2879 | 0.33 | 0.07 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 123 | 0.2560 | 0.2550 | 0.2540 | 0.39 | 0.08 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 103 | 0.2659 | 0.2647 | 0.2682 | 0.67 | 0.14 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 128 | 0.3285 | 0.3238 | 0.3270 | 0.74 | 0.16 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 199 | 0.2980 | 0.2995 | 0.2951 | 0.75 | 0.16 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 164 | 0.2754 | 0.2797 | 0.2772 | 0.78 | 0.16 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 155 | 0.2610 | 0.2630 | 0.2650 | 0.76 | 0.16 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 168 | 0.2580 | 0.2614 | 0.2570 | 0.89 | 0.18 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 204 | 0.3019 | 0.2982 | 0.3037 | 0.93 | 0.19 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 202 | 0.2334 | 0.2381 | 0.2355 | 1 | 0.2 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 122 | 0.3319 | 0.3296 | 0.3256 | 0.97 | 0.21 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.99 | Other | 154 | 0.2955 | 0.2898 | 0.2951 | 1.08 | 0.23 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 124 | 0.2433 | 0.2438 | 0.2484 | 1.15 | 0.23 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 156 | 0.3916 | 0.3833 | 0.3869 | 1.07 | 0.23 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 159 | 0.2510 | 0.2570 | 0.2540 | 1.18 | 0.24 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.60 | NIR | 187 | 0.2031 | 0.2007 | 0.1981 | 1.25 | 0.25 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 153 | 0.2623 | 0.2639 | 0.2686 | 1.24 | 0.25 | 3.48 | 0.33 | 14 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 116 | 0.2974 | 0.2903 | 0.2964 | 1.3 | 0.27 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 149 | 0.3030 | 0.3000 | 0.2950 | 1.35 | 0.28 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 165 | 0.3000 | 0.3049 | 0.2965 | 1.4 | 0.29 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 163 | 0.3360 | 0.3460 | 0.3410 | 1.47 | 0.31 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 186 | 0.2868 | 0.2787 | 0.2791 | 1.62 | 0.34 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 150 | 0.2976 | 0.2881 | 0.2910 | 1.67 | 0.35 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 196 | 0.2351 | 0.2441 | 0.2400 | 1.88 | 0.38 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 112 | 0.3000 | 0.2900 | 0.3000 | 1.95 | 0.41 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 116 | 0.3126 | 0.3187 | 0.3065 | 1.95 | 0.41 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 161 | 0.2355 | 0.2455 | 0.2423 | 2.12 | 0.43 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 189 | 0.3449 | 0.3310 | 0.3395 | 2.07 | 0.44 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 192 | 0.2342 | 0.2410 | 0.2308 | 2.21 | 0.45 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 200 | 0.2690 | 0.2586 | 0.2592 | 2.23 | 0.46 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.03 | AOAC 2018.11, M | 139 | 0.2978 | 0.2855 | 0.2967 | 2.32 | 0.48 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 170 | 0.3439 | 0.3486 | 0.3596 | 2.3 | 0.49 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 144 | 0.2660 | 0.2752 | 0.2800 | 2.6 | 0.54 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 194 | 0.3008 | 0.3004 | 0.2870 | 2.65 | 0.55 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 172 | 0.3127 | 0.3273 | 0.3290 | 2.77 | 0.59 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 141 | 0.2318 | 0.2185 | 0.2233 | 3 | 0.6 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 152 | 0.2813 | 0.2896 | 0.2991 | 3.07 | 0.64 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 190 | 0.4014 | 0.3800 | 0.3830 | 2.99 | 0.65 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 142 | 0.2827 | 0.2631 | 0.2691 | 3.7 | 0.76 | 3.48 | 0.33 | 14 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 120 | 0.2962 | 0.2928 | 0.3169 | 4.32 | 0.9 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 109 | 0.2880 | 0.3090 | 0.2860 | 4.33 | 0.9 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 101 | 0.3050 | 0.2770 | 0.2910 | 4.81 | 1 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 108 | 0.3138 | 0.2862 | 0.3088 | 4.85 | 1.02 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 119 | 0.3125 | 0.2932 | 0.3245 | 5.09 | 1.07 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 149 | 0.3000 | 0.3100 | 0.2800 | 5.15 | 1.08 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 131 | 0.2800 | 0.2700 | 0.3000 | 5.39 | 1.12 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 198 | 0.2851 | 0.2607 | 0.2880 | 5.4 | 1.12 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 191 | 0.2627 | 0.2925 | 0.2755 | 5.4 | 1.12 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 125 | 0.3164 | 0.3205 | 0.3490 | 5.4 | 1.15 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 107 | 0.0286 | 0.0286 | 0.0328 | 8.08 | 1.2 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 197 | 0.2190 | 0.2250 | 0.2000 | 6.08 | 1.21 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 203 | 0.2600 | 0.2300 | 0.2500 | 6.19 | 1.26 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 145 | 0.2800 | 0.2800 | 0.2500 | 6.42 | 1.32 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 138 | 0.2869 | 0.3206 | 0.3248 | 6.69 | 1.41 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.99 | Other | 132 | 0.3700 | 0.3300 | 0.3300 | 6.73 | 1.44 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 115 | 0.2900 | 0.2800 | 0.3200 | 7.02 | 1.47 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 205 | 0.2420 | 0.2770 | 0.2410 | 8.09 | 1.65 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 114 | 0.2376 | 0.2593 | 0.2794 | 8.08 | 1.65 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.60 | NIR | 138 | 0.6300 | 0.5500 | 0.5650 | 7.31 | 1.69 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 182 | 0.3283 | 0.2848 | 0.2755 | 9.52 | 1.99 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 121 | 0.2850 | 0.3210 | 0.3530 | 10.6 | 2.25 | 3.48 | 0.33 | 14 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 175 | 0.2959 | 0.2999 | 0.2427 | 11.4 | 2.36 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 188 | 0.2688 | 0.2107 | 0.2430 | 12.1 | 2.45 | 3.48 | 0.33 | 14 | 0 |
| Total Δ9-THC (%AR) | 006.99 | Other | 172 | 0.2720 | 0.2631 | 0.3375 | 14 | 2.91 | 3.48 | 0.33 | 14 | 0 |
|  |  |  |  | Total CBD (%AR) |  |  |  |  |  |  |  |  |
| Total CBD (%AR) | 007.10 | Other LC-UV | 178 | 7.2500 |  |  |  |  |  |  |  | 1 |
| Total CBD (%AR) | 007.60 | NIR | 187 | 5.4804 | 5.4914 | 5.5035 | 0.21 | 0.07 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 161 | 5.8837 | 5.8540 | 5.8636 | 0.26 | 0.08 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.03 | AOAC 2018.11, M | 139 | 6.3951 | 6.3684 | 6.3611 | 0.28 | 0.09 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 199 | 7.1632 | 7.1718 | 7.1147 | 0.43 | 0.14 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 204 | 7.0596 | 6.9970 | 7.0286 | 0.44 | 0.15 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 123 | 7.3710 | 7.3430 | 7.2970 | 0.51 | 0.17 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 122 | 7.5729 | 7.4958 | 7.5548 | 0.54 | 0.18 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 150 | 6.9613 | 6.9572 | 6.8695 | 0.75 | 0.25 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 186 | 7.1142 | 7.1079 | 7.0180 | 0.76 | 0.26 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 159 | 6.5400 | 6.5700 | 6.4700 | 0.79 | 0.26 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.01 | AOAC 2018.10 | 131 | 7.6200 | 7.7700 | 7.6300 | 1.09 | 0.37 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 109 | 7.9590 | 7.9260 | 7.7730 | 1.26 | 0.43 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 156 | 7.4744 | 7.6605 | 7.5998 | 1.25 | 0.43 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 189 | 6.8357 | 6.6617 | 6.6977 | 1.36 | 0.46 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 205 | 6.8560 | 6.7330 | 6.6740 | 1.37 | 0.46 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 188 | 7.5110 | 7.3630 | 7.5670 | 1.41 | 0.48 | 2.37 | 0.21 | 17 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%AR) | 007.30 | Other LC-MS | 164 | 7.2084 | 6.9867 | 7.1227 | 1.57 | 0.53 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 163 | 6.9690 | 7.1930 | 7.0810 | 1.58 | 0.53 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.01 | AOAC 2018.10 | 168 | 6.8040 | 6.6091 | 6.8004 | 1.65 | 0.55 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 203 | 6.8300 | 7.0500 | 6.9900 | 1.63 | 0.55 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 170 | 8.1712 | 8.2846 | 8.4434 | 1.65 | 0.57 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 152 | 7.4044 | 7.1755 | 7.1679 | 1.85 | 0.63 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 197 | 7.4010 | 7.3080 | 7.1140 | 2.01 | 0.68 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 192 | 5.5912 | 5.4821 | 5.3608 | 2.1 | 0.68 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 108 | 6.6480 | 6.5030 | 6.7790 | 2.08 | 0.69 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 116 | 7.3953 | 7.1083 | 7.1232 | 2.24 | 0.76 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 145 | 7.5300 | 7.2600 | 7.2300 | 2.25 | 0.76 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 125 | 7.8814 | 8.2432 | 8.0611 | 2.24 | 0.77 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 116 | 6.8158 | 7.1416 | 7.0247 | 2.36 | 0.79 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 196 | 6.1610 | 6.4620 | 6.3140 | 2.38 | 0.79 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 200 | 6.4096 | 6.2424 | 6.0684 | 2.73 | 0.9 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 142 | 7.1698 | 6.8234 | 6.8625 | 2.73 | 0.91 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 144 | 7.5618 | 7.3304 | 7.1459 | 2.84 | 0.96 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.99 | Other | 172 | 7.1704 | 7.5568 | 7.5363 | 2.93 | 0.99 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.01 | AOAC 2018.10 | 155 | 6.0700 | 6.1300 | 5.7700 | 3.22 | 1.06 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.99 | Other | 154 | 6.1464 | 5.7515 | 5.9116 | 3.35 | 1.1 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 172 | 7.6780 | 8.0667 | 8.1918 | 3.36 | 1.15 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 194 | 7.3928 | 7.5034 | 7.0240 | 3.44 | 1.16 | 2.37 | 0.21 | 17 | 0 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%AR) | 007.10 | Other LC-UV | 149 | 7.1200 | 7.5960 | 7.2080 | 3.47 | 1.17 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 141 | 6.0717 | 5.6420 | 5.7657 | 3.8 | 1.24 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 198 | 6.9082 | 6.4195 | 6.8681 | 4.03 | 1.34 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.60 | NIR | 138 | 9.2000 | 8.5050 | 8.7450 | 4 | 1.39 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 162 | 7.1396 | 7.1321 | 7.6945 | 4.41 | 1.49 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 202 | 5.9821 | 6.0457 | 6.5034 | 4.6 | 1.52 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 191 | 3.3669 | 3.7392 | 3.6898 | 5.62 | 1.71 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 138 | 7.6707 | 8.3211 | 8.4641 | 5.19 | 1.78 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 119 | 7.1680 | 6.5600 | 7.2497 | 5.39 | 1.81 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.99 | Other | 132 | 7.5500 | 6.7800 | 6.7000 | 6.7 | 2.25 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 149 | 8.1100 | 7.8700 | 7.0930 | 6.91 | 2.35 | 2.37 | 0.21 | 17 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 185 | 10.0400 | 12.7080 | 14.1690 | 17 | 6.21 | 2.37 | 0.21 | 17 | 0 |

Lab

Population of Lab RSDr

CBDV (%AR)

| CBDV (%AR) | 008.99 | Other | 194 | <0.0239 | <0.0239 | <0.0239 | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| CBDV (%AR) | 008.99 | Other | 188 | <0.0500 | <0.0500 | <0.0500 | 1 |
| CBDV (%AR) | 008.99 | Other | 197 | <0.0500 | <0.0500 | <0.0500 | 1 |
| CBDV (%AR) | 008.99 | Other | 203 | <0.0040 | 0.0190 | 0.0180 | 1 |
| CBDV (%AR) | 008.99 | Other | 162 | <0.0100 | <0.0100 | <0.0100 | 1 |
| CBDV (%AR) | 008.02 | AOAC 2018.11, U | 108 | <0.0050 | <0.0050 | <0.0050 | 1 |
| CBDV (%AR) | 008.99 | Other | 142 | <0.0100 | <0.0100 | <0.0100 | 1 |
| CBDV (%AR) | 008.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 | 1 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDV (%AR) | 008.99 | Other | 132 | 0.0200 | 0.0200 | 0.0200 | 0 | 0 | 6.71 | 0 | 128 | 0 |
| CBDV (%AR) | 008.99 | Other | 202 | 0.0157 | 0.0165 | 0.0161 | 2.48 | 0.34 | 6.71 | 0 | 128 | 0 |
| CBDV (%AR) | 008.99 | Other | 159 | 0.0357 | 0.0365 | 0.0347 | 2.53 | 0.38 | 6.71 | 0 | 128 | 0 |
| CBDV (%AR) | 008.99 | Other | 200 | 0.0294 | 0.0286 | 0.0279 | 2.62 | 0.39 | 6.71 | 0 | 128 | 0 |
| CBDV (%AR) | 008.99 | Other | 109 | 0.1100 | 0.1090 | 0.1050 | 2.45 | 0.44 | 6.71 | 0 | 128 | 0 |
| CBDV (%AR) | 008.99 | Other | 189 | 0.0361 | 0.0345 | 0.0376 | 4.3 | 0.65 | 6.71 | 0 | 128 | 0 |
| CBDV (%AR) | 008.99 | Other | 205 | 0.0450 | 0.0520 | 0.0510 | 7.67 | 1.22 | 6.71 | 0 | 128 | 0 |
| CBDV (%AR) | 008.99 | Other | 149 | 0.0362 | 0.0320 | 0.0300 | 9.67 | 1.45 | 6.71 | 0 | 128 | 0 |
| CBDV (%AR) | 008.99 | Other | 106 | 0.0600 | 0.0500 | 0.0500 | 10.8 | 1.75 | 6.71 | 0 | 128 | 0 |
| CBDV (%AR) | 008.99 | Other | 195 | 0.0431 | 0.0428 | 0.0326 | 15.1 | 2.34 | 6.71 | 0 | 128 | 0 |
| CBDV (%AR) | 008.99 | Other | 186 | 0.0396 | 0.3840 | 0.0422 | 128 | 24.18 | 6.71 | 0 | 128 | 0 |
|  |  |  |  | CBDVA (%AR) |  |  |  |  |  |  |  |  |
| CBDVA (%AR) | 009.01 | AOAC 2018.10 | 131 | 0.0400 | 0.0400 | 0.0400 | 0 | 0 | 4.37 | 0 | 8.79 | 0 |
| CBDVA (%AR) | 009.99 | Other | 159 | 0.0355 | 0.0354 | 0.0353 | 0.28 | 0.04 | 4.37 | 0 | 8.79 | 0 |
| CBDVA (%AR) | 009.99 | Other | 186 | 0.0888 | 0.0874 | 0.0860 | 1.6 | 0.28 | 4.37 | 0 | 8.79 | 0 |
| CBDVA (%AR) | 009.99 | Other | 154 | 0.0561 | 0.0556 | 0.0587 | 2.93 | 0.48 | 4.37 | 0 | 8.79 | 0 |
| CBDVA (%AR) | 009.99 | Other | 196 | 0.0293 | 0.0311 | 0.0315 | 3.83 | 0.57 | 4.37 | 0 | 8.79 | 0 |
| CBDVA (%AR) | 009.01 | AOAC 2018.10 | 155 | 0.0390 | 0.0360 | 0.0380 | 4.06 | 0.62 | 4.37 | 0 | 8.79 | 0 |
| CBDVA (%AR) | 009.99 | Other | 203 | 0.0320 | 0.0330 | 0.0290 | 6.64 | 0.99 | 4.37 | 0 | 8.79 | 0 |
| CBDVA (%AR) | 009.02 | AOAC 2018.11, U | 108 | 0.0389 | 0.0454 | 0.0427 | 7.71 | 1.2 | 4.37 | 0 | 8.79 | 0 |
| CBDVA (%AR) | 009.01 | AOAC 2018.10 | 168 | 0.0449 | 0.0470 | 0.0522 | 7.82 | 1.24 | 4.37 | 0 | 8.79 | 0 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDVA (%AR) | 009.99 | Other | 149 | 0.0440 | 0.0420 | 0.0370 | 8.79 | 1.36 | 4.37 | 0 | 8.79 | 0 |

Lab

Population of Lab RSDr

CBG (%AR)

| CBG (%AR) | 010.99 | Other | 141 | <0.0120 | <0.0120 | <0.0120 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%AR) | 010.99 | Other | 123 | <0.0880 | <0.0930 | <0.0880 |  |  |  |  |  | 1 |
| CBG (%AR) | 010.99 | Other | 109 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| CBG (%AR) | 010.99 | Other | 197 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| CBG (%AR) | 010.99 | Other | 106 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| CBG (%AR) | 010.99 | Other | 154 | <0.0700 | <0.0700 | <0.0700 |  |  |  |  |  | 1 |
| CBG (%AR) | 010.99 | Other | 200 | 0.1898 | 0.1887 | 0.1886 | 0.35 | 0.07 | 7.45 | 0.35 | 51.8 | 0 |
| CBG (%AR) | 010.01 | AOAC 2018.10 | 155 | 0.0890 | 0.0870 | 0.0880 | 1.14 | 0.2 | 7.45 | 0.35 | 51.8 | 0 |
| CBG (%AR) | 010.99 | Other | 159 | 0.0609 | 0.0597 | 0.0614 | 1.44 | 0.24 | 7.45 | 0.35 | 51.8 | 0 |
| CBG (%AR) | 010.99 | Other | 195 | 0.0710 | 0.0700 | 0.0685 | 1.8 | 0.3 | 7.45 | 0.35 | 51.8 | 0 |
| CBG (%AR) | 010.99 | Other | 203 | 0.0690 | 0.0680 | 0.0720 | 2.99 | 0.5 | 7.45 | 0.35 | 51.8 | 0 |
| CBG (%AR) | 010.99 | Other | 196 | 0.0540 | 0.0575 | 0.0560 | 3.14 | 0.51 | 7.45 | 0.35 | 51.8 | 0 |
| CBG (%AR) | 010.99 | Other | 205 | 0.0960 | 0.0890 | 0.0920 | 3.8 | 0.67 | 7.45 | 0.35 | 51.8 | 0 |
| CBG (%AR) | 010.99 | Other | 119 | 0.0804 | 0.0770 | 0.0838 | 4.23 | 0.73 | 7.45 | 0.35 | 51.8 | 0 |
| CBG (%AR) | 010.02 | AOAC 2018.11, U | 108 | 0.0941 | 0.0893 | 0.0972 | 4.26 | 0.75 | 7.45 | 0.35 | 51.8 | 0 |
| CBG (%AR) | 010.99 | Other | 142 | 0.0670 | 0.0610 | 0.0640 | 4.69 | 0.78 | 7.45 | 0.35 | 51.8 | 0 |
| CBG (%AR) | 010.99 | Other | 204 | 0.1080 | 0.1021 | 0.0984 | 4.71 | 0.84 | 7.45 | 0.35 | 51.8 | 0 |
| CBG (%AR) | 010.01 | AOAC 2018.10 | 168 | 0.1028 | 0.1097 | 0.0980 | 5.68 | 1.01 | 7.45 | 0.35 | 51.8 | 0 |
| CBG (%AR) | 010.02 | AOAC 2018.11, U | 125 | 0.1270 | 0.1266 | 0.1143 | 5.89 | 1.08 | 7.45 | 0.35 | 51.8 | 0 |

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%AR) | 010.99 | Other | 194 | 0.0755 | 0.0845 | 0.0864 | 7.09 | 1.22 | 7.45 | 0.35 | 51.8 | 0 |
| CBG (%AR) | 010.99 | Other | 132 | 0.0800 | 0.0700 | 0.0700 | 7.87 | 1.33 | 7.45 | 0.35 | 51.8 | 0 |
| CBG (%AR) | 010.02 | AOAC 2018.11, U | 152 | 0.1139 | 0.1257 | 0.1077 | 7.9 | 1.43 | 7.45 | 0.35 | 51.8 | 0 |
| CBG (%AR) | 010.99 | Other | 149 | 0.0780 | 0.0750 | 0.0660 | 8.55 | 1.45 | 7.45 | 0.35 | 51.8 | 0 |
| CBG (%AR) | 010.99 | Other | 188 | 0.0780 | 0.0725 | 0.0904 | 11.4 | 1.96 | 7.45 | 0.35 | 51.8 | 0 |
| CBG (%AR) | 010.01 | AOAC 2018.10 | 131 | 0.1100 | 0.1000 | 0.0800 | 15.8 | 2.79 | 7.45 | 0.35 | 51.8 | 0 |
| CBG (%AR) | 010.99 | Other | 202 | 0.0660 | 0.0862 | 0.0507 | 26.3 | 4.4 | 7.45 | 0.35 | 51.8 | 0 |
| CBG (%AR) | 010.99 | Other | 162 | 0.0481 | 0.0402 | 0.0714 | 30.5 | 4.92 | 7.45 | 0.35 | 51.8 | 0 |
| CBG (%AR) | 010.99 | Other | 189 | 0.0552 | 0.0527 | 0.0924 | 33.3 | 5.56 | 7.45 | 0.35 | 51.8 | 0 |
| CBG (%AR) | 010.99 | Other | 186 | 0.0163 | 0.0547 | 0.0426 | 51.8 | 7.95 | 7.45 | 0.35 | 51.8 | 0 |

Lab

Population of Lab RSDr

CBGA (%AR)

| CBGA (%AR) | 011.99 | Other | 109 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%AR) | 011.02 | AOAC 2018.11, U | 108 | <0.0008 | <0.0008 | <0.0008 |  |  |  |  |  | 1 |
| CBGA (%AR) | 011.99 | Other | 132 | 0.0900 | <0.0100 | 0.0700 |  |  |  |  |  | 1 |
| CBGA (%AR) | 011.01 | AOAC 2018.10 | 131 | 0.1100 | 0.1100 | 0.1100 | 0 | 0 | 5.28 | 0 | 23.8 | 0 |
| CBGA (%AR) | 011.99 | Other | 196 | 0.0793 | 0.0809 | 0.0801 | 1 | 0.17 | 5.28 | 0 | 23.8 | 0 |
| CBGA (%AR) | 011.99 | Other | 123 | 0.1060 | 0.1040 | 0.1040 | 1.1 | 0.2 | 5.28 | 0 | 23.8 | 0 |
| CBGA (%AR) | 011.99 | Other | 203 | 0.1030 | 0.1000 | 0.1010 | 1.51 | 0.27 | 5.28 | 0 | 23.8 | 0 |
| CBGA (%AR) | 011.99 | Other | 154 | 0.1201 | 0.1186 | 0.1234 | 2.03 | 0.37 | 5.28 | 0 | 23.8 | 0 |
| CBGA (%AR) | 011.99 | Other | 202 | 0.0909 | 0.0949 | 0.0930 | 2.15 | 0.38 | 5.28 | 0 | 23.8 | 0 |
| CBGA (%AR) | 011.99 | Other | 189 | 0.1192 | 0.1150 | 0.1146 | 2.19 | 0.4 | 5.28 | 0 | 23.8 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%AR) | 011.99 | Other | 186 | 0.1060 | 0.1065 | 0.1013 | 2.74 | 0.49 | 5.28 | 0 | 23.8 | 0 |
| CBGA (%AR) | 011.01 | AOAC 2018.10 | 155 | 0.1200 | 0.1130 | 0.1170 | 3.01 | 0.55 | 5.28 | 0 | 23.8 | 0 |
| CBGA (%AR) | 011.02 | AOAC 2018.11, U | 125 | 0.1184 | 0.1279 | 0.1240 | 3.87 | 0.71 | 5.28 | 0 | 23.8 | 0 |
| CBGA (%AR) | 011.99 | Other | 204 | 0.1175 | 0.1114 | 0.1086 | 4.05 | 0.73 | 5.28 | 0 | 23.8 | 0 |
| CBGA (%AR) | 011.99 | Other | 159 | 0.0598 | 0.0585 | 0.0638 | 4.55 | 0.75 | 5.28 | 0 | 23.8 | 0 |
| CBGA (%AR) | 011.99 | Other | 195 | 0.0993 | 0.0979 | 0.0910 | 4.63 | 0.82 | 5.28 | 0 | 23.8 | 0 |
| CBGA (%AR) | 011.99 | Other | 141 | 0.0538 | 0.0481 | 0.0507 | 5.61 | 0.9 | 5.28 | 0 | 23.8 | 0 |
| CBGA (%AR) | 011.99 | Other | 194 | 0.0913 | 0.0880 | 0.0809 | 6.13 | 1.06 | 5.28 | 0 | 23.8 | 0 |
| CBGA (%AR) | 011.99 | Other | 162 | 0.0928 | 0.1009 | 0.0898 | 6.08 | 1.07 | 5.28 | 0 | 23.8 | 0 |
| CBGA (%AR) | 011.99 | Other | 197 | 0.0930 | 0.0800 | 0.0880 | 7.54 | 1.31 | 5.28 | 0 | 23.8 | 0 |
| CBGA (%AR) | 011.99 | Other | 149 | 0.0890 | 0.0870 | 0.0770 | 7.62 | 1.32 | 5.28 | 0 | 23.8 | 0 |
| CBGA (%AR) | 011.99 | Other | 205 | 0.1070 | 0.0980 | 0.0920 | 7.63 | 1.35 | 5.28 | 0 | 23.8 | 0 |
| CBGA (%AR) | 011.01 | AOAC 2018.10 | 168 | 0.0994 | 0.0994 | 0.1160 | 9.13 | 1.63 | 5.28 | 0 | 23.8 | 0 |
| CBGA (%AR) | 011.02 | AOAC 2018.11, U | 152 | 0.1177 | 0.0956 | 0.1077 | 10.3 | 1.85 | 5.28 | 0 | 23.8 | 0 |
| CBGA (%AR) | 011.99 | Other | 188 | 0.1102 | 0.0906 | 0.0919 | 11.2 | 1.99 | 5.28 | 0 | 23.8 | 0 |
| CBGA (%AR) | 011.99 | Other | 106 | 0.1000 | 0.1000 | 0.0700 | 19.2 | 3.36 | 5.28 | 0 | 23.8 | 0 |
| CBGA (%AR) | 011.99 | Other | 142 | 0.1365 | 0.1111 | 0.1780 | 23.8 | 4.45 | 5.28 | 0 | 23.8 | 0 |
|  |  |  |  | THCV (%AR) |  |  |  |  |  |  |  |  |
| THCV (%AR) | 012.99 | Other | 109 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 |  |  |  |  |  | 1 |
| THCV (%AR) | 012.02 | AOAC 2018.11, U | 108 | <0.0051 | <0.0051 | <0.0051 |  |  |  |  |  | 1 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) Rob Mean min max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| THCV (%AR) | 012.99 | Other | 106 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 149 | <0.0240 | <0.0240 | <0.0240 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 197 | <0.0500 | <0.0500 | <0.0500 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 188 | <0.0500 | <0.0500 | <0.0500 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 123 | <0.0880 | <0.0930 | <0.0880 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 132 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 168 | <0.0250 | <0.0250 | <0.0250 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 186 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 195 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 203 | 0.0000 | 0.0000 | <0.0290 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 142 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 154 | <0.0150 | <0.0150 | <0.0150 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 194 | <0.0239 | <0.0239 | <0.0239 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 200 | 0.0290 | 0.0283 | 0.0292 | 1.64 | 0.24 | 0 |
| THCV (%AR) | 012.99 | Other | 189 | 0.0374 | 0.0358 | 0.0367 | 2.19 | 0.33 | 0 |
| THCV (%AR) | 012.99 | Other | 202 | 0.0150 | 0.0159 | 0.0150 | 3.4 | 0.45 | 0 |
| THCV (%AR) | 012.03 | AOAC 2018.11, M | 159 | 0.0162 | 0.0174 | 0.0169 | 3.58 | 0.49 | 0 |
| THCV (%AR) | 012.99 | Other | 119 | 0.0145 | 0.0142 | 0.0154 | 4.25 | 0.57 | 0 |

Lab

Population of Lab RSDr

|  |  |  |  | Δ8-THC (%AR) |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%AR) | 013.01 | AOAC 2018.10 | 131 | <0.0100 <0.0100 | <0.0100 | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 205 | <0.0070 <0.0070 | <0.0070 | 1 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%AR) | 013.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 |  |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 188 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 162 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 202 | <0.0001 | <0.0001 | <0.0001 |  |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 203 | 0.0000 | 0.0000 | 0.0000 | 0 |  |  |  |  | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 197 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 186 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 159 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 154 | <0.0490 | <0.0490 | <0.0490 |  |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 106 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 149 | <0.0240 | <0.0240 | <0.0240 |  |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.02 | AOAC 2018.11, U | 108 | <0.0045 | <0.0045 | <0.0045 |  |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 194 | <0.0239 | <0.0239 | <0.0239 |  |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 109 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 123 | <0.0880 | <0.0930 | <0.0880 |  |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 200 | <0.0025 | <0.0025 | <0.0025 |  |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 141 | <0.0120 | <0.0120 | <0.0120 |  |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 142 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 195 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 189 | 0.0035 | 0.0154 | 0.0102 | 61.5 | 7.69 |  |  |  | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | CBC (%AR) |  |  |  |  |  |  |  |  |
| CBC (%AR) | 014.99 | Other | 204 | 0.2097 | 0.2095 | 0.2106 | 0.28 | 0.06 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.99 | Other | 203 | 0.1890 | 0.1920 | 0.1880 | 1.1 | 0.21 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.99 | Other | 195 | 0.2072 | 0.2038 | 0.2027 | 1.15 | 0.23 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.99 | Other | 154 | 0.1729 | 0.1721 | 0.1689 | 1.24 | 0.24 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.99 | Other | 186 | 0.1976 | 0.1922 | 0.1952 | 1.39 | 0.27 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.99 | Other | 162 | 0.2236 | 0.2189 | 0.2259 | 1.6 | 0.32 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.99 | Other | 189 | 0.2167 | 0.2089 | 0.2129 | 1.83 | 0.36 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.01 | AOAC 2018.10 | 168 | 0.1880 | 0.1829 | 0.1815 | 1.86 | 0.36 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.99 | Other | 142 | 0.1908 | 0.1835 | 0.1898 | 2.1 | 0.41 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.99 | Other | 196 | 0.1741 | 0.1817 | 0.1791 | 2.17 | 0.42 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.99 | Other | 109 | 0.2320 | 0.2270 | 0.2220 | 2.2 | 0.44 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.99 | Other | 202 | 0.2175 | 0.2095 | 0.2085 | 2.33 | 0.46 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.99 | Other | 200 | 0.2020 | 0.1920 | 0.1952 | 2.6 | 0.51 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.99 | Other | 159 | 0.1460 | 0.1540 | 0.1480 | 2.79 | 0.53 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, U | 125 | 0.2331 | 0.2457 | 0.2384 | 2.65 | 0.54 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.99 | Other | 123 | 0.2110 | 0.2020 | 0.1990 | 3.06 | 0.6 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.99 | Other | 188 | 0.2077 | 0.2024 | 0.2186 | 3.94 | 0.78 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.99 | Other | 119 | 0.2419 | 0.2286 | 0.2487 | 4.26 | 0.86 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, U | 108 | 0.2268 | 0.2079 | 0.2129 | 4.54 | 0.9 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.99 | Other | 149 | 0.2100 | 0.2100 | 0.1900 | 5.68 | 1.12 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.99 | Other | 106 | 0.2100 | 0.2100 | 0.1900 | 5.68 | 1.12 | 3.44 | 0.28 | 12 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%AR) | 014.99 | Other | 205 | 0.2110 | 0.1870 | 0.1900 | 6.67 | 1.31 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.01 | AOAC 2018.10 | 155 | 0.2070 | 0.1810 | 0.1900 | 6.85 | 1.34 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.99 | Other | 197 | 0.2030 | 0.2050 | 0.1800 | 7.09 | 1.39 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.99 | Other | 132 | 0.2300 | 0.2000 | 0.2000 | 8.25 | 1.64 | 3.44 | 0.28 | 12 | 0 |
| CBC (%AR) | 014.01 | AOAC 2018.10 | 131 | 0.2000 | 0.2100 | 0.2500 | 12 | 2.4 | 3.44 | 0.28 | 12 | 0 |
|  |  |  |  | Cd (ug/g AR) |  |  |  |  |  |  |  |  |
| Cd (ug/g AR) | 101.99 | Other | 195 | 21.2200 | 14.5600 | <14.8500 |  |  |  |  |  | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 162 | 0.1723 | 0.1723 | 0.1662 | 2.07 | 0.4 | 5.67 | 2.07 | 13.8 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 196 | 0.1580 | 0.1610 | 0.1530 | 2.57 | 0.49 | 5.67 | 2.07 | 13.8 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 168 | 0.1750 | 0.1690 | 0.1660 | 2.7 | 0.52 | 5.67 | 2.07 | 13.8 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 197 | 0.1780 | 0.1800 | 0.1899 | 3.49 | 0.68 | 5.67 | 2.07 | 13.8 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 142 | 0.1710 | 0.1630 | 0.1780 | 4.4 | 0.85 | 5.67 | 2.07 | 13.8 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 137 | 0.1730 | 0.1850 | 0.1650 | 5.77 | 1.11 | 5.67 | 2.07 | 13.8 | 0 |
| Cd (ug/g AR) | 101.03 | ICP-OES | 132 | 0.1360 | 0.1310 | 0.1480 | 6.32 | 1.18 | 5.67 | 2.07 | 13.8 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 122 | 0.1858 | 0.1504 | 0.1849 | 11.6 | 2.24 | 5.67 | 2.07 | 13.8 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 141 | 0.1680 | 0.1350 | 0.1320 | 13.8 | 2.58 | 5.67 | 2.07 | 13.8 | 0 |
|  |  |  |  | As (ug/g AR) |  |  |  |  |  |  |  |  |
| As (ug/g AR) | 102.99 | Other | 195 | <0.8400 | <0.7880 | 2.2640 |  |  |  |  |  | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 197 | 0.3280 | 0.3400 | 0.3348 | 1.8 | 0.38 | 5.74 | 1.8 | 31.4 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 196 | 0.5150 | 0.5340 | 0.5170 | 2 | 0.45 | 5.74 | 1.8 | 31.4 | 0 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| As (ug/g AR) | 102.04 | ICP-MS | 142 | 0.3050 | 0.3190 | 0.3130 | 2.25 | 0.47 | 5.74 | 1.8 | 31.4 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 168 | 0.3220 | 0.3360 | 0.3250 | 2.25 | 0.48 | 5.74 | 1.8 | 31.4 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 162 | 0.3841 | 0.3780 | 0.3550 | 4.12 | 0.89 | 5.74 | 1.8 | 31.4 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 137 | 0.3690 | 0.3690 | 0.3350 | 5.49 | 1.18 | 5.74 | 1.8 | 31.4 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 122 | 0.3166 | 0.2727 | 0.3164 | 8.38 | 1.75 | 5.74 | 1.8 | 31.4 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 141 | 0.3480 | 0.2900 | 0.2800 | 12 | 2.52 | 5.74 | 1.8 | 31.4 | 0 |
| As (ug/g AR) | 102.03 | ICP-OES | 132 | 0.7200 | 0.4200 | 0.4460 | 31.4 | 7.16 | 5.74 | 1.8 | 31.4 | 0 |

Lab

Population of Lab RSDr

Hg (ug/g AR)

| Hg (ug/g AR) | 103.04 | ICP-MS | 162 | <0.1000 | <0.1000 | <0.1000 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Hg (ug/g AR) | 103.99 | Other | 195 | <1.6560 | <1.6930 | <1.8880 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 137 | 0.0131 | <0.0094 | 0.0096 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 141 | <0.0139 | <0.0139 | 0.0150 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 142 | <0.0005 | <0.0005 | <0.0005 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 197 | <0.0990 | <0.0990 | <0.1025 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 168 | <0.0140 | <0.0130 | <0.0130 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 196 | 0.0083 | 0.0092 | 0.0074 | 10.8 | 1.32 | 0 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 122 | 0.0078 | 0.0059 | 0.0082 | 16.8 | 2.02 | 0 |

Pb (ug/g AR)

| Pb (ug/g AR) | 104.99 | Other | 195 | <2.4000 <2.3620 |  | <6.6700 | 1 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Pb (ug/g AR) | 104.02 | furnace AA | 132 | 0.2120 | 0.2120 | 0.2120 | 0 | 0 | 5.96 | 0 | 11.7 | 0 |

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Pb (ug/g AR) | 104.04 | ICP-MS | 122 | 0.9857 | 0.9709 | 0.9783 | 0.76 | 0.19 | 5.96 | 0 | 11.7 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 141 | 1.1900 | 1.1800 | 1.1100 | 3.76 | 0.96 | 5.96 | 0 | 11.7 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 162 | 1.0702 | 1.1924 | 1.0995 | 5.69 | 1.45 | 5.96 | 0 | 11.7 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 197 | 0.8810 | 0.9830 | 0.8799 | 6.47 | 1.6 | 5.96 | 0 | 11.7 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 137 | 1.0900 | 1.2200 | 1.2500 | 7.17 | 1.84 | 5.96 | 0 | 11.7 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 142 | 0.9100 | 1.0070 | 1.0810 | 8.58 | 2.15 | 5.96 | 0 | 11.7 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 168 | 1.3000 | 1.0800 | 1.1600 | 9.44 | 2.42 | 5.96 | 0 | 11.7 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 196 | 1.1500 | 1.3500 | 1.0800 | 11.7 | 3.02 | 5.96 | 0 | 11.7 | 0 |

Lab

Population of Lab RSDr

Cu (ug/g AR)

| Cu (ug/g AR) | 105.04 | ICP-MS | 122 | 4.2700 | 4.1808 | 4.3554 | 2.05 | 0.64 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cu (ug/g AR) | 105.03 | ICP-OES | 132 | 11.8720 | 11.4480 | 11.8720 | 2.09 | 0.76 | 0 |
| Cu (ug/g AR) | 105.04 | ICP-MS | 142 | 16.0420 | 15.6780 | 14.7280 | 4.38 | 1.66 | 0 |
| Cu (ug/g AR) | 105.99 | Other | 195 | 12.3950 | 14.2450 | 13.7050 | 7.07 | 2.62 | 0 |
| Cu (ug/g AR) | 105.04 | ICP-MS | 196 | 13.0000 | 15.0000 | 13.0000 | 8.45 | 3.13 | 0 |

Ni (ug/g AR)

| Ni (ug/g AR) | 106.04 | ICP-MS | 122 | 4.2700 | 4.1808 | 4.2254 | 1.06 | 0.33 | 3.73 | 1.06 | 22.1 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Ni (ug/g AR) | 106.04 | ICP-MS | 141 | 4.0800 | 4.1000 | 3.9400 | 2.16 | 0.67 | 3.73 | 1.06 | 22.1 | 0 |
| Ni (ug/g AR) | 106.04 | ICP-MS | 142 | 4.0240 | 4.0370 | 4.1910 | 2.27 | 0.7 | 3.73 | 1.06 | 22.1 | 0 |
| Ni (ug/g AR) | 106.04 | ICP-MS | 196 | 4.1200 | 4.2700 | 4.0500 | 2.71 | 0.84 | 3.73 | 1.06 | 22.1 | 0 |
| Ni (ug/g AR) | 106.03 | ICP-OES | 132 | 1.7810 | 1.8660 | 1.9930 | 5.68 | 1.56 | 3.73 | 1.06 | 22.1 | 0 |

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Ni (ug/g AR) | 106.99 | Other | 195 | 3.9160 | 6.1640 | 5.3110 | 22.1 | 7.08 | 3.73 | 1.06 | 22.1 | 0 |

Lab

Population of Lab RSDr

Cr (ug/g AR)

| Cr (ug/g AR) | 107.99 | Other | 195 | <5.7540 | <5.7530 | <5.8150 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cr (ug/g AR) | 107.04 | ICP-MS | 197 | 12.9710 | 13.1250 | 13.0599 | 0.59 | 0.22 | 2.38 | 0.59 | 6.65 | 0 |
| Cr (ug/g AR) | 107.04 | ICP-MS | 141 | 12.4000 | 12.5000 | 12.3000 | 0.81 | 0.29 | 2.38 | 0.59 | 6.65 | 0 |
| Cr (ug/g AR) | 107.03 | ICP-OES | 132 | 1.6540 | 1.6960 | 1.6960 | 1.44 | 0.39 | 2.38 | 0.59 | 6.65 | 0 |
| Cr (ug/g AR) | 107.04 | ICP-MS | 122 | 12.1475 | 11.7718 | 11.9597 | 1.57 | 0.57 | 2.38 | 0.59 | 6.65 | 0 |
| Cr (ug/g AR) | 107.04 | ICP-MS | 196 | 11.6000 | 12.1000 | 11.2000 | 3.88 | 1.4 | 2.38 | 0.59 | 6.65 | 0 |
| Cr (ug/g AR) | 107.04 | ICP-MS | 142 | 5.7510 | 5.4260 | 6.1930 | 6.65 | 2.17 | 2.38 | 0.59 | 6.65 | 0 |

Se (ug/g AR)

| Se (ug/g AR) | 108.99 | Other | 195 | <0.9250 | <0.9880 | <0.1250 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Se (ug/g AR) | 108.04 | ICP-MS | 197 | 0.1630 | 0.1710 | 0.1579 | 4.03 | 0.77 | 0 |
| Se (ug/g AR) | 108.04 | ICP-MS | 141 | 0.3270 | 0.2780 | 0.2990 | 8.16 | 1.71 | 0 |
| Se (ug/g AR) | 108.03 | ICP-OES | 132 | 2.3740 | 2.4590 | 2.1620 | 6.56 | 1.87 | 0 |
| Se (ug/g AR) | 108.04 | ICP-MS | 196 | 7.4000 | 6.4400 | 7.5500 | 8.45 | 2.84 | 0 |
| Se (ug/g AR) | 108.04 | ICP-MS | 142 | 0.1500 | 0.1090 | 0.1420 | 16.3 | 3.01 | 0 |

|  |  |  |  | Ag (ug/g AR) |  |
| --- | --- | --- | --- | --- | --- |
| Ag (ug/g AR) | 109.04 | ICP-MS | 142 | <0.0005 <0.0005 <0.0005 | 1 |
| Ag (ug/g AR) | 109.99 | Other | 195 | 17.2000 <8.8380 <11.4100 | 1 |

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) Rob Mean min max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Ag (ug/g AR) | 109.04 | ICP-MS | 122 | 0.0429 | 0.0577 | 0.0503 | 14.7 | 2.35 | 0 |
| Ag (ug/g AR) | 109.03 | ICP-OES | 132 | 0.0360 | 0.0360 | 0.0000 | 86.6 | 12.4 | 0 |

Lab

Population of Lab RSDr

Ba (ug/g AR)

| Ba (ug/g AR) | 110.99 | Other | 195 | 71.6500 | 71.9700 | <71.8700 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Ba (ug/g AR) | 110.04 | ICP-MS | 122 | 55.4917 | 57.7376 | 57.9350 | 2.38 | 1.09 | 0 |
| Ba (ug/g AR) | 110.03 | ICP-OES | 132 | 37.7360 | 38.5840 | 40.0680 | 3.04 | 1.32 | 0 |

Zn (ug/g AR)

| Zn (ug/g AR) | 111.03 | ICP-OES | 132 | 47.9120 | 47.4880 | 47.0640 | 0.89 | 0.4 | 6.41 | 0.89 | 18.6 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Zn (ug/g AR) | 111.04 | ICP-MS | 196 | 64.1000 | 65.6000 | 64.7000 | 1.17 | 0.55 | 6.41 | 0.89 | 18.6 | 0 |
| Zn (ug/g AR) | 111.04 | ICP-MS | 122 | 52.8297 | 53.9311 | 54.2635 | 1.4 | 0.64 | 6.41 | 0.89 | 18.6 | 0 |
| Zn (ug/g AR) | 111.99 | Other | 195 | 47.2780 | 44.4550 | 45.8970 | 3.08 | 1.37 | 6.41 | 0.89 | 18.6 | 0 |
| Zn (ug/g AR) | 111.04 | ICP-MS | 142 | 52.1710 | 58.1380 | 67.8450 | 13.3 | 6.16 | 6.41 | 0.89 | 18.6 | 0 |
| Zn (ug/g AR) | 111.04 | ICP-MS | 141 | 59.2000 | 44.7000 | 42.5000 | 18.6 | 8.35 | 6.41 | 0.89 | 18.6 | 0 |

Sb (ug/g AR)

| Sb (ug/g AR) | 112.99 | Other | 195 | 38.4200 42.3400 <47.2600 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Sb (ug/g AR) | 112.04 | ICP-MS | 196 | <0.0241 | 0.0250 | 0.0241 | 1 |

Moisture (%AR)

Moisture (%AR)

500.99

Other

178

3.8500	1

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 204 | 6.8100 | 6.8100 | 6.8100 | 0 | 0 | 3.95 | 0 | 11.2 | 0 |
| Moisture (%AR) | 500.99 | Other | 103 | 5.9700 | 5.9700 | 5.9700 | 0 | 0 | 3.95 | 0 | 11.2 | 0 |
| Moisture (%AR) | 500.99 | Other | 142 | 6.2600 | 6.2600 | 6.2600 | 0 | 0 | 3.95 | 0 | 11.2 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 116 | 6.3600 | 6.3800 | 6.3300 | 0.4 | 0.13 | 3.95 | 0 | 11.2 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 182 | 7.0550 | 7.1500 | 7.3600 | 2.17 | 0.73 | 3.95 | 0 | 11.2 | 0 |
| Moisture (%AR) | 500.99 | Other | 165 | 7.9237 | 8.3077 | 8.2353 | 2.5 | 0.86 | 3.95 | 0 | 11.2 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 119 | 6.5900 | 6.9400 | 6.9000 | 2.81 | 0.94 | 3.95 | 0 | 11.2 | 0 |
| Moisture (%AR) | 500.99 | Other | 132 | 5.9400 | 6.1000 | 6.2900 | 2.87 | 0.94 | 3.95 | 0 | 11.2 | 0 |
| Moisture (%AR) | 500.72 | NIR | 138 | 9.1000 | 8.6000 | 8.7000 | 3.01 | 1.04 | 3.95 | 0 | 11.2 | 0 |
| Moisture (%AR) | 500.72 | NIR | 187 | 5.1028 | 5.3895 | 5.4904 | 3.77 | 1.22 | 3.95 | 0 | 11.2 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 199 | 7.9142 | 8.2768 | 8.5228 | 3.72 | 1.28 | 3.95 | 0 | 11.2 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 105 | 8.5400 | 9.2800 | 8.9400 | 4.15 | 1.44 | 3.95 | 0 | 11.2 | 0 |
| Moisture (%AR) | 500.99 | Other | 120 | 5.7660 | 6.1414 | 6.4639 | 5.7 | 1.88 | 3.95 | 0 | 11.2 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 198 | 5.5700 | 5.7400 | 6.2300 | 5.86 | 1.91 | 3.95 | 0 | 11.2 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 192 | 9.0243 | 9.6748 | 10.1495 | 5.87 | 2.07 | 3.95 | 0 | 11.2 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 183 | 9.8600 | 8.9800 | 8.6100 | 7.02 | 2.45 | 3.95 | 0 | 11.2 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 196 | 6.1900 | 7.0200 | 7.1300 | 7.58 | 2.53 | 3.95 | 0 | 11.2 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 186 | 6.7800 | 5.5400 | 6.1600 | 10.1 | 3.31 | 3.95 | 0 | 11.2 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 109 | 6.5500 | 5.6100 | 7.0100 | 11.2 | 3.7 | 3.95 | 0 | 11.2 | 0 |
|  |  |  |  | Δ9-THC (%DW) |  |  |  |  |  |  |  |  |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 178 | 0.2400 |  |  |  |  |  |  |  | 1 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%DW) | 501.60 | NIR | 187 | 0.1845 | 0.1847 | 0.1853 | 0.22 | 0.04 | 2.51 | 0.23 | 16.2 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 118 | 0.2250 | 0.2260 | 0.2250 | 0.26 | 0.05 | 2.51 | 0.23 | 16.2 | 0 |
| Δ9-THC (%DW) | 501.30 | Other LC-MS | 156 | 0.3652 | 0.3618 | 0.3639 | 0.47 | 0.1 | 2.51 | 0.23 | 16.2 | 0 |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, U | 199 | 0.2731 | 0.2720 | 0.2684 | 0.91 | 0.19 | 2.51 | 0.23 | 16.2 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 204 | 0.2597 | 0.2551 | 0.2607 | 1.16 | 0.24 | 2.51 | 0.23 | 16.2 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 186 | 0.2509 | 0.2451 | 0.2455 | 1.31 | 0.27 | 2.51 | 0.23 | 16.2 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 194 | 0.2510 | 0.2494 | 0.2432 | 1.66 | 0.34 | 2.51 | 0.23 | 16.2 | 0 |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, U | 103 | 0.2242 | 0.2151 | 0.2223 | 2.18 | 0.43 | 2.51 | 0.23 | 16.2 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 165 | 0.2586 | 0.2701 | 0.2606 | 2.34 | 0.48 | 2.51 | 0.23 | 16.2 | 0 |
| Δ9-THC (%DW) | 501.30 | Other LC-MS | 188 | 0.2197 | 0.2143 | 0.2254 | 2.53 | 0.5 | 2.51 | 0.23 | 16.2 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 196 | 0.2071 | 0.2179 | 0.2126 | 2.54 | 0.5 | 2.51 | 0.23 | 16.2 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 116 | 0.2769 | 0.2810 | 0.2678 | 2.45 | 0.51 | 2.51 | 0.23 | 16.2 | 0 |
| Δ9-THC (%DW) | 501.30 | Other LC-MS | 105 | 0.2600 | 0.2590 | 0.2710 | 2.53 | 0.52 | 2.51 | 0.23 | 16.2 | 0 |
| Δ9-THC (%DW) | 501.30 | Other LC-MS | 120 | 0.2532 | 0.2508 | 0.2658 | 3.14 | 0.64 | 2.51 | 0.23 | 16.2 | 0 |
| Δ9-THC (%DW) | 501.99 | Other | 183 | 0.2680 | 0.2770 | 0.2591 | 3.34 | 0.69 | 2.51 | 0.23 | 16.2 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 142 | 0.2535 | 0.2353 | 0.2413 | 3.81 | 0.77 | 2.51 | 0.23 | 16.2 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 198 | 0.2611 | 0.2404 | 0.2570 | 4.34 | 0.88 | 2.51 | 0.23 | 16.2 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 109 | 0.2430 | 0.2660 | 0.2470 | 4.88 | 0.99 | 2.51 | 0.23 | 16.2 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 172 | 0.2910 | 0.3032 | 0.2742 | 5.03 | 1.05 | 2.51 | 0.23 | 16.2 | 0 |
| Δ9-THC (%DW) | 501.99 | Other | 172 | 0.2262 | 0.2423 | 0.3054 | 16.2 | 3.32 | 2.51 | 0.23 | 16.2 | 0 |

Population of Lab RSDr

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

RSDr	HorRat(r)

Population of Lab RSDr	 Rob Mean	min	max

Flag

Δ9-THCA (%DW)

| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 178 | 0.1100 |  |  |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%DW) | 502.30 | Other LC-MS | 188 | 0.0613 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 118 | 0.0530 | 0.0530 | 0.0530 | 0 | 0 | 5.26 | 0 | 20.6 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 204 | 0.0733 | 0.0740 | 0.0744 | 0.75 | 0.13 | 5.26 | 0 | 20.6 | 0 |
| Δ9-THCA (%DW) | 502.30 | Other LC-MS | 105 | 0.0492 | 0.0477 | 0.0491 | 1.72 | 0.27 | 5.26 | 0 | 20.6 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 116 | 0.0650 | 0.0676 | 0.0679 | 2.39 | 0.4 | 5.26 | 0 | 20.6 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 196 | 0.0497 | 0.0509 | 0.0523 | 2.55 | 0.41 | 5.26 | 0 | 20.6 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 142 | 0.0548 | 0.0517 | 0.0522 | 3.15 | 0.51 | 5.26 | 0 | 20.6 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 186 | 0.0624 | 0.0591 | 0.0591 | 3.16 | 0.52 | 5.26 | 0 | 20.6 | 0 |
| Δ9-THCA (%DW) | 502.30 | Other LC-MS | 156 | 0.0754 | 0.0703 | 0.0747 | 3.76 | 0.64 | 5.26 | 0 | 20.6 | 0 |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, U | 199 | 0.0576 | 0.0621 | 0.0618 | 4.16 | 0.68 | 5.26 | 0 | 20.6 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 109 | 0.0740 | 0.0690 | 0.0690 | 4.09 | 0.69 | 5.26 | 0 | 20.6 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 165 | 0.0767 | 0.0711 | 0.0714 | 4.31 | 0.73 | 5.26 | 0 | 20.6 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 172 | 0.0658 | 0.0702 | 0.0624 | 5.91 | 0.99 | 5.26 | 0 | 20.6 | 0 |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, U | 103 | 0.0667 | 0.0757 | 0.0714 | 6.32 | 1.07 | 5.26 | 0 | 20.6 | 0 |
| Δ9-THCA (%DW) | 502.60 | NIR | 187 | 0.0336 | 0.0313 | 0.0276 | 9.82 | 1.46 | 5.26 | 0 | 20.6 | 0 |
| Δ9-THCA (%DW) | 502.99 | Other | 183 | 0.0752 | 0.0710 | 0.0628 | 9.05 | 1.52 | 5.26 | 0 | 20.6 | 0 |
| Δ9-THCA (%DW) | 502.30 | Other LC-MS | 120 | 0.0698 | 0.0698 | 0.0832 | 10.4 | 1.77 | 5.26 | 0 | 20.6 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 198 | 0.0475 | 0.0417 | 0.0558 | 14.7 | 2.33 | 5.26 | 0 | 20.6 | 0 |
| Δ9-THCA (%DW) | 502.99 | Other | 172 | 0.0876 | 0.0579 | 0.0805 | 20.6 | 3.5 | 5.26 | 0 | 20.6 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | CBD (%DW) |  |  |  |  |  |  |  |  |
| CBD (%DW) | 503.10 | Other LC-UV | 178 | 3.3300 |  |  |  |  |  |  |  | 1 |
| CBD (%DW) | 503.10 | Other LC-UV | 118 | 3.2500 | 3.2650 | 3.2570 | 0.23 | 0.07 | 1.98 | 0.23 | 5.26 | 0 |
| CBD (%DW) | 503.02 | AOAC 2018.11, U | 199 | 3.9784 | 4.0006 | 4.0147 | 0.46 | 0.14 | 1.98 | 0.23 | 5.26 | 0 |
| CBD (%DW) | 503.60 | NIR | 187 | 2.8754 | 2.8929 | 2.9097 | 0.59 | 0.17 | 1.98 | 0.23 | 5.26 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 204 | 3.9157 | 3.8721 | 3.8840 | 0.58 | 0.18 | 1.98 | 0.23 | 5.26 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 186 | 3.9078 | 3.8774 | 3.8350 | 0.94 | 0.29 | 1.98 | 0.23 | 5.26 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 109 | 4.1340 | 4.0700 | 4.0660 | 0.93 | 0.29 | 1.98 | 0.23 | 5.26 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 142 | 4.0483 | 3.9707 | 4.0132 | 0.97 | 0.3 | 1.98 | 0.23 | 5.26 | 0 |
| CBD (%DW) | 503.30 | Other LC-MS | 188 | 3.9340 | 3.9150 | 3.8250 | 1.5 | 0.46 | 1.98 | 0.23 | 5.26 | 0 |
| CBD (%DW) | 503.30 | Other LC-MS | 105 | 3.4100 | 3.5200 | 3.5100 | 1.75 | 0.53 | 1.98 | 0.23 | 5.26 | 0 |
| CBD (%DW) | 503.30 | Other LC-MS | 156 | 4.6088 | 4.7546 | 4.7324 | 1.67 | 0.53 | 1.98 | 0.23 | 5.26 | 0 |
| CBD (%DW) | 503.99 | Other | 183 | 4.0457 | 4.0379 | 3.8706 | 2.48 | 0.77 | 1.98 | 0.23 | 5.26 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 116 | 3.8900 | 4.1108 | 4.0157 | 2.77 | 0.85 | 1.98 | 0.23 | 5.26 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 196 | 3.4960 | 3.7070 | 3.6630 | 3.07 | 0.93 | 1.98 | 0.23 | 5.26 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 194 | 3.9373 | 3.9142 | 3.7048 | 3.33 | 1.02 | 1.98 | 0.23 | 5.26 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 198 | 3.6192 | 3.3861 | 3.6059 | 3.7 | 1.12 | 1.98 | 0.23 | 5.26 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 172 | 4.4738 | 4.6745 | 4.2653 | 4.58 | 1.44 | 1.98 | 0.23 | 5.26 | 0 |
| CBD (%DW) | 503.99 | Other | 172 | 4.1066 | 4.4138 | 4.5551 | 5.26 | 1.64 | 1.98 | 0.23 | 5.26 | 0 |
|  |  |  |  | CBDA (%DW) |  |  |  |  |  |  |  |  |
| CBDA (%DW) | 504.10 | Other LC-UV | 178 | 4.2100 |  |  |  |  |  |  |  | 1 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDA (%DW) | 504.60 | NIR | 187 | 3.3068 | 3.3204 | 3.3226 | 0.26 | 0.08 | 2.47 | 0.26 | 7.63 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 204 | 4.1731 | 4.1463 | 4.1713 | 0.36 | 0.11 | 2.47 | 0.26 | 7.63 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 118 | 3.6980 | 3.7300 | 3.6930 | 0.54 | 0.17 | 2.47 | 0.26 | 7.63 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 186 | 4.1886 | 4.2155 | 4.1548 | 0.73 | 0.23 | 2.47 | 0.26 | 7.63 | 0 |
| CBDA (%DW) | 504.02 | AOAC 2018.11, U | 199 | 4.3335 | 4.3539 | 4.2907 | 0.75 | 0.23 | 2.47 | 0.26 | 7.63 | 0 |
| CBDA (%DW) | 504.30 | Other LC-MS | 156 | 4.1202 | 4.1905 | 4.1873 | 0.95 | 0.3 | 2.47 | 0.26 | 7.63 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 109 | 4.9970 | 4.9330 | 4.8960 | 1.03 | 0.33 | 2.47 | 0.26 | 7.63 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 116 | 3.8641 | 4.0090 | 3.9751 | 1.92 | 0.59 | 2.47 | 0.26 | 7.63 | 0 |
| CBDA (%DW) | 504.99 | Other | 172 | 4.4251 | 4.5656 | 4.4002 | 2 | 0.63 | 2.47 | 0.26 | 7.63 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 196 | 3.5030 | 3.6980 | 3.5750 | 2.75 | 0.83 | 2.47 | 0.26 | 7.63 | 0 |
| CBDA (%DW) | 504.30 | Other LC-MS | 105 | 3.5200 | 3.6900 | 3.7100 | 2.87 | 0.87 | 2.47 | 0.26 | 7.63 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 198 | 4.2395 | 3.9134 | 4.2061 | 4.35 | 1.35 | 2.47 | 0.26 | 7.63 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 142 | 4.1052 | 3.7725 | 3.7715 | 4.95 | 1.52 | 2.47 | 0.26 | 7.63 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 172 | 4.6589 | 4.9242 | 4.4771 | 4.8 | 1.52 | 2.47 | 0.26 | 7.63 | 0 |
| CBDA (%DW) | 504.99 | Other | 183 | 4.6828 | 4.9974 | 4.4860 | 5.46 | 1.73 | 2.47 | 0.26 | 7.63 | 0 |
| CBDA (%DW) | 504.30 | Other LC-MS | 188 | 4.2270 | 4.0770 | 3.6420 | 7.63 | 2.35 | 2.47 | 0.26 | 7.63 | 0 |

Lab

Population of Lab RSDr

CBN (%DW)

| CBN (%DW) | 505.02 | AOAC 2018.11, U | 199 | <0.0550 | <0.0550 | <0.0550 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%DW) | 505.20 | Other LC-MS | 188 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| CBN (%DW) | 505.10 | Other LC-UV | 109 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| CBN (%DW) | 505.10 | Other LC-UV | 204 | 0.0300 | 0.0298 | 0.0302 | 0.67 | 0.1 | 4.23 | 0.67 | 17.6 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%DW) | 505.20 | Other LC-MS | 105 | 0.0264 | 0.0260 | 0.0263 | 0.79 | 0.12 | 4.23 | 0.67 | 17.6 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 196 | 0.0219 | 0.0226 | 0.0230 | 2.47 | 0.35 | 4.23 | 0.67 | 17.6 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 186 | 0.0314 | 0.0304 | 0.0299 | 2.5 | 0.37 | 4.23 | 0.67 | 17.6 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 142 | 0.0273 | 0.0268 | 0.0283 | 2.78 | 0.41 | 4.23 | 0.67 | 17.6 | 0 |
| CBN (%DW) | 505.30 | GC-FID | 116 | 0.0507 | 0.0494 | 0.0527 | 3.26 | 0.52 | 4.23 | 0.67 | 17.6 | 0 |
| CBN (%DW) | 505.20 | Other LC-MS | 156 | 0.0360 | 0.0394 | 0.0385 | 4.64 | 0.71 | 4.23 | 0.67 | 17.6 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 192 | 0.0329 | 0.0368 | 0.0341 | 5.77 | 0.87 | 4.23 | 0.67 | 17.6 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 198 | 0.0889 | 0.0807 | 0.0875 | 5.12 | 0.89 | 4.23 | 0.67 | 17.6 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 116 | 0.0524 | 0.0420 | 0.0437 | 12.1 | 1.91 | 4.23 | 0.67 | 17.6 | 0 |
| CBN (%DW) | 505.99 | Other | 183 | 0.0241 | 0.0326 | 0.0339 | 17.6 | 2.61 | 4.23 | 0.67 | 17.6 | 0 |
|  |  |  |  | Total Δ9-THC (%DW) |  |  |  |  |  |  |  |  |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 178 | 0.3600 |  |  |  |  |  |  |  | 1 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 118 | 0.2710 | 0.2730 | 0.2720 | 0.37 | 0.08 | 3.14 | 0.37 | 14 | 0 |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, U | 199 | 0.3236 | 0.3265 | 0.3225 | 0.64 | 0.13 | 3.14 | 0.37 | 14 | 0 |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, U | 103 | 0.2828 | 0.2816 | 0.2853 | 0.67 | 0.14 | 3.14 | 0.37 | 14 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 204 | 0.3240 | 0.3200 | 0.3259 | 0.93 | 0.2 | 3.14 | 0.37 | 14 | 0 |
| Total Δ9-THC (%DW) | 506.60 | NIR | 187 | 0.2140 | 0.2121 | 0.2095 | 1.07 | 0.21 | 3.14 | 0.37 | 14 | 0 |
| Total Δ9-THC (%DW) | 506.30 | Other LC-MS | 156 | 0.4308 | 0.4214 | 0.4249 | 1.12 | 0.25 | 3.14 | 0.37 | 14 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 116 | 0.3176 | 0.3100 | 0.3166 | 1.31 | 0.28 | 3.14 | 0.37 | 14 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 165 | 0.3258 | 0.3325 | 0.3232 | 1.47 | 0.31 | 3.14 | 0.37 | 14 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 186 | 0.3056 | 0.2970 | 0.2974 | 1.62 | 0.34 | 3.14 | 0.37 | 14 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 116 | 0.3339 | 0.3403 | 0.3273 | 1.95 | 0.41 | 3.14 | 0.37 | 14 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 192 | 0.2574 | 0.2668 | 0.2569 | 2.14 | 0.44 | 3.14 | 0.37 | 14 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 196 | 0.2507 | 0.2626 | 0.2584 | 2.35 | 0.48 | 3.14 | 0.37 | 14 | 0 |
| Total Δ9-THC (%DW) | 506.30 | Other LC-MS | 105 | 0.3030 | 0.3010 | 0.3140 | 2.29 | 0.48 | 3.14 | 0.37 | 14 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 109 | 0.3080 | 0.3270 | 0.3070 | 3.59 | 0.76 | 3.14 | 0.37 | 14 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 142 | 0.3016 | 0.2807 | 0.2871 | 3.7 | 0.77 | 3.14 | 0.37 | 14 | 0 |
| Total Δ9-THC (%DW) | 506.99 | Other | 183 | 0.3340 | 0.3393 | 0.3142 | 4.02 | 0.85 | 3.14 | 0.37 | 14 | 0 |
| Total Δ9-THC (%DW) | 506.30 | Other LC-MS | 120 | 0.3144 | 0.3120 | 0.3388 | 4.61 | 0.97 | 3.14 | 0.37 | 14 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 172 | 0.3487 | 0.3648 | 0.3290 | 5.16 | 1.1 | 3.14 | 0.37 | 14 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 198 | 0.3028 | 0.2769 | 0.3059 | 5.39 | 1.13 | 3.14 | 0.37 | 14 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 138 | 0.3146 | 0.3515 | 0.3561 | 6.68 | 1.42 | 3.14 | 0.37 | 14 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 182 | 0.3533 | 0.3067 | 0.2974 | 9.39 | 1.98 | 3.14 | 0.37 | 14 | 0 |
| Total Δ9-THC (%DW) | 506.30 | Other LC-MS | 188 | 0.2734 | 0.2143 | 0.2254 | 13.2 | 2.67 | 3.14 | 0.37 | 14 | 0 |
| Total Δ9-THC (%DW) | 506.99 | Other | 172 | 0.3030 | 0.2931 | 0.3760 | 14 | 2.96 | 3.14 | 0.37 | 14 | 0 |
|  |  |  |  | Total CBD (%DW) |  |  |  |  |  |  |  |  |
| Total CBD (%DW) | 507.10 | Other LC-UV | 178 | 7.5400 |  |  |  |  |  |  |  | 1 |
| Total CBD (%DW) | 507.02 | AOAC 2018.11, U | 199 | 7.7788 | 7.8190 | 7.7776 | 0.3 | 0.1 | 2.28 | 0.3 | 5.19 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 118 | 6.4930 | 6.5360 | 6.4960 | 0.37 | 0.12 | 2.28 | 0.3 | 5.19 | 0 |
| Total CBD (%DW) | 507.60 | NIR | 187 | 5.7755 | 5.8049 | 5.8236 | 0.42 | 0.14 | 2.28 | 0.3 | 5.19 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 204 | 7.5755 | 7.5084 | 7.5422 | 0.44 | 0.15 | 2.28 | 0.3 | 5.19 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 186 | 7.5812 | 7.5745 | 7.4787 | 0.76 | 0.26 | 2.28 | 0.3 | 5.19 | 0 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%DW) | 507.10 | Other LC-UV | 109 | 8.5160 | 8.3970 | 8.3590 | 0.97 | 0.34 | 2.28 | 0.3 | 5.19 | 0 |
| Total CBD (%DW) | 507.30 | Other LC-MS | 156 | 8.2212 | 8.4286 | 8.4020 | 1.35 | 0.47 | 2.28 | 0.3 | 5.19 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 192 | 6.1458 | 6.0693 | 5.9663 | 1.49 | 0.49 | 2.28 | 0.3 | 5.19 | 0 |
| Total CBD (%DW) | 507.30 | Other LC-MS | 105 | 6.5000 | 6.7600 | 6.7600 | 2.25 | 0.75 | 2.28 | 0.3 | 5.19 | 0 |
| Total CBD (%DW) | 507.40 | GC-FID | 116 | 7.8976 | 7.5911 | 7.6070 | 2.24 | 0.76 | 2.28 | 0.3 | 5.19 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 116 | 7.2787 | 7.6267 | 7.5019 | 2.36 | 0.8 | 2.28 | 0.3 | 5.19 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 142 | 7.6486 | 7.2791 | 7.3208 | 2.73 | 0.92 | 2.28 | 0.3 | 5.19 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 196 | 6.5680 | 6.9500 | 6.7990 | 2.84 | 0.95 | 2.28 | 0.3 | 5.19 | 0 |
| Total CBD (%DW) | 507.99 | Other | 172 | 7.9874 | 8.4178 | 8.4141 | 2.99 | 1.03 | 2.28 | 0.3 | 5.19 | 0 |
| Total CBD (%DW) | 507.99 | Other | 183 | 8.1525 | 8.4207 | 7.8047 | 3.8 | 1.3 | 2.28 | 0.3 | 5.19 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 198 | 7.3373 | 6.8181 | 7.2946 | 4.03 | 1.36 | 2.28 | 0.3 | 5.19 | 0 |
| Total CBD (%DW) | 507.30 | Other LC-MS | 188 | 7.6410 | 7.4900 | 7.0190 | 4.39 | 1.49 | 2.28 | 0.3 | 5.19 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 172 | 8.5596 | 8.9930 | 8.1919 | 4.67 | 1.62 | 2.28 | 0.3 | 5.19 | 0 |
| Total CBD (%DW) | 507.40 | GC-FID | 138 | 8.4109 | 9.1240 | 9.2808 | 5.19 | 1.81 | 2.28 | 0.3 | 5.19 | 0 |

Lab

Population of Lab RSDr

CBDV (%DW)

| CBDV (%DW) | 508.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDV (%DW) | 508.99 | Other | 142 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| CBDV (%DW) | 508.99 | Other | 188 | <0.0500 | <0.0500 | <0.0500 |  |  | 1 |
| CBDV (%DW) | 508.99 | Other | 186 | 0.0422 | 0.0410 | 0.0406 | 2.02 | 0.31 | 0 |
| CBDV (%DW) | 508.99 | Other | 109 | 0.1180 | 0.1160 | 0.1130 | 2.18 | 0.39 | 0 |
| CBDV (%DW) | 508.99 | Other | 183 | 0.0404 | 0.0468 | 0.0411 | 8.21 | 1.28 | 0 |

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

RSDr	HorRat(r)

Population of Lab RSDr	 Rob Mean	min	max

Flag

CBDVA (%DW)

| CBDVA (%DW) | 509.99 | Other | 186 | 0.0946 | 0.0931 | 0.0917 | 1.56 | 0.27 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDVA (%DW) | 509.99 | Other | 183 | 0.0564 | 0.0562 | 0.0543 | 2.08 | 0.34 | 0 |
| CBDVA (%DW) | 509.99 | Other | 196 | 0.0312 | 0.0334 | 0.0339 | 4.37 | 0.66 | 0 |

CBG (%DW)

| CBG (%DW) | 510.99 | Other | 109 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%DW) | 510.99 | Other | 196 | 0.0576 | 0.0618 | 0.0603 | 3.55 | 0.58 | 19.28 | 3.55 | 55 | 0 |
| CBG (%DW) | 510.99 | Other | 142 | 0.0715 | 0.0651 | 0.0683 | 4.69 | 0.78 | 19.28 | 3.55 | 55 | 0 |
| CBG (%DW) | 510.99 | Other | 204 | 0.1056 | 0.1159 | 0.1096 | 4.7 | 0.85 | 19.28 | 3.55 | 55 | 0 |
| CBG (%DW) | 510.99 | Other | 188 | 0.0793 | 0.0738 | 0.0839 | 6.4 | 1.1 | 19.28 | 3.55 | 55 | 0 |
| CBG (%DW) | 510.99 | Other | 183 | 0.1060 | 0.0905 | 0.0922 | 8.83 | 1.56 | 19.28 | 3.55 | 55 | 0 |
| CBG (%DW) | 510.99 | Other | 186 | 0.0174 | 0.0583 | 0.0454 | 51.8 | 8.02 | 19.28 | 3.55 | 55 | 0 |
| CBG (%DW) | 510.99 | Other | 105 | 0.0677 | 0.0675 | 0.1620 | 55 | 9.75 | 19.28 | 3.55 | 55 | 0 |

CBGA (%DW)

| CBGA (%DW) | 511.99 | Other | 109 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%DW) | 511.99 | Other | 183 | 0.1231 | 0.1242 | 0.1232 | 0.49 | 0.09 | 7.76 | 0.49 | 23.8 | 0 |
| CBGA (%DW) | 511.99 | Other | 196 | 0.0845 | 0.0870 | 0.0863 | 1.5 | 0.26 | 7.76 | 0.49 | 23.8 | 0 |
| CBGA (%DW) | 511.99 | Other | 186 | 0.1130 | 0.1135 | 0.1079 | 2.78 | 0.5 | 7.76 | 0.49 | 23.8 | 0 |
| CBGA (%DW) | 511.99 | Other | 204 | 0.1261 | 0.1195 | 0.1165 | 4.07 | 0.74 | 7.76 | 0.49 | 23.8 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%DW) | 511.99 | Other | 188 | 0.1121 | 0.0922 | 0.0852 | 14.5 | 2.55 | 7.76 | 0.49 | 23.8 | 0 |
| CBGA (%DW) | 511.99 | Other | 142 | 0.1456 | 0.1185 | 0.1899 | 23.8 | 4.5 | 7.76 | 0.49 | 23.8 | 0 |
|  |  |  |  | THCV (%DW) |  |  |  |  |  |  |  |  |
| THCV (%DW) | 512.99 | Other | 186 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| THCV (%DW) | 512.99 | Other | 183 | 0.0000 | 0.0000 | 0.0000 | 0 |  |  |  |  | 0 |
| THCV (%DW) | 512.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 |  |  |  |  |  | 1 |
| THCV (%DW) | 512.99 | Other | 188 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| THCV (%DW) | 512.99 | Other | 109 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| THCV (%DW) | 512.99 | Other | 142 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
|  |  |  |  | Δ8-THC (%DW) |  |  |  |  |  |  |  |  |
| Δ8-THC (%DW) | 513.99 | Other | 183 | 0.0000 | 0.0000 | 0.0000 | 0 |  |  |  |  | 0 |
| Δ8-THC (%DW) | 513.99 | Other | 142 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| Δ8-THC (%DW) | 513.99 | Other | 109 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| Δ8-THC (%DW) | 513.99 | Other | 186 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| Δ8-THC (%DW) | 513.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 |  |  |  |  |  | 1 |
| Δ8-THC (%DW) | 513.99 | Other | 188 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
|  |  |  |  | CBC (%DW) |  |  |  |  |  |  |  |  |
| CBC (%DW) | 514.99 | Other | 204 | 0.2250 | 0.2248 | 0.2260 | 0.28 | 0.06 | 1.8 | 0.29 | 2.77 | 0 |
| CBC (%DW) | 514.99 | Other | 105 | 0.2110 | 0.2110 | 0.2140 | 0.82 | 0.16 | 1.8 | 0.29 | 2.77 | 0 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%DW) | 514.99 | Other | 186 | 0.2105 | 0.2048 | 0.2081 | 1.38 | 0.27 | 1.8 | 0.29 | 2.77 | 0 |
| CBC (%DW) | 514.99 | Other | 109 | 0.2480 | 0.2410 | 0.2390 | 1.95 | 0.39 | 1.8 | 0.29 | 2.77 | 0 |
| CBC (%DW) | 514.99 | Other | 188 | 0.2112 | 0.2059 | 0.2028 | 2.06 | 0.41 | 1.8 | 0.29 | 2.77 | 0 |
| CBC (%DW) | 514.99 | Other | 183 | 0.2379 | 0.2446 | 0.2352 | 2.02 | 0.41 | 1.8 | 0.29 | 2.77 | 0 |
| CBC (%DW) | 514.99 | Other | 142 | 0.2035 | 0.1958 | 0.2025 | 2.09 | 0.41 | 1.8 | 0.29 | 2.77 | 0 |
| CBC (%DW) | 514.99 | Other | 196 | 0.1856 | 0.1954 | 0.1929 | 2.66 | 0.52 | 1.8 | 0.29 | 2.77 | 0 |
| CBC (%DW) | 514.99 | Other | 194 | 0.2243 | 0.2256 | 0.2144 | 2.77 | 0.55 | 1.8 | 0.29 | 2.77 | 0 |

Lab

Population of Lab RSDr

Cd (ug/g DW)

| Cd (ug/g DW) | 601.04 | ICP-MS | 105 | 0.1830 | 0.1850 | 0.1810 | 1.09 | 0.21 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

As (ug/g DW)

| As (ug/g DW) | 602.04 | ICP-MS | 105 | 0.3460 | 0.3400 | 0.3560 | 2.33 | 0.5 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

Hg (ug/g DW)

Hg (ug/g DW)

603.04

ICP-MS

105

<0.0100	<0.0100	<0.1000	1

Pb (ug/g DW)

| Pb (ug/g DW) | 604.04 | ICP-MS | 105 | 1.2900 | 1.3400 | 1.1700 | 6.9 | 1.79 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

Cu (ug/g DW)

Cu (ug/g DW)

605.04

ICP-MS

105

15.9000	15.8000	16.1000	0.96	0.36	0

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

RSDr	HorRat(r)

Population of Lab RSDr	 Rob Mean	min	max

Flag

Ni (ug/g DW)

Ni (ug/g DW)

606.04

ICP-MS

105

4.8100	4.7700	5.2200	5.05	1.61	0

Cr (ug/g DW)

Cr (ug/g DW)

607.04

ICP-MS

105

13.3000	13.8000	13.8000	2.12	0.79	0

Se (ug/g DW)

Se (ug/g DW)

608.04

ICP-MS

105

0.0830	0.0693	0.9200	136	29.28	0

Zn (ug/g DW)

Zn (ug/g DW)

611.04

ICP-MS

105

61.1000	58.1000	63.1000	4.14	1.92	0

- Flag definitions: 1 (Lab RSDr not calculated because there were less than 3 numeric lab results)

- Statistical parameters of robust mean, minimum, and maximum for RSDr only shown only if there were 6 or more lab RSDr values.

- AR refers to concentration on as-received basis. DW refers to concentration on dry weight basis.c