| Homogeneity and Stability |  |  |  |
| --- | --- | --- | --- |
| Sample ID . . . . . . . . . . . . . . . . . . . . . . . . . | HM20SEP-2 |  |  |
| Samples shipped . . . . . . . . . . . . . . . . . . . | 8-Sep-20 |  |  |
| Results due date . . . . . . . . . . . . . . . . . . . . | 15-Oct-20 |  |  |
| Homogeneity 10 sample packets randomly chosen for analysis. Duplicate test portions analyzed for analyte and method shown | in table caption. |  |  |
| Total THC (%AR) via GC-FID (method 006.40) Analysis Date: 9/9/20 | Overall average: | 0.3422 | %RSD |
|  | SD of sample avgs: | 0.005914 | 1.73 |
|  | repeatability SD: | 0.003364 | 0.98 |
|  | between-sample SD: | 0.00541 | 1.58 |
|  | reproducibilty SD: | 0.006374 | 1.86 |
|  | SD used for z scores: | 0.0371 |  |
|  | Check value: | 0.01113 |  |

| Packet # | Replicate 1 | Replicate 2 |
| --- | --- | --- |
| 1 | 0.3262 | 0.3340 |
| 2 | 0.3438 | 0.3403 |
| 3 | 0.3400 | 0.3388 |
| 4 | 0.3371 | 0.3438 |
| 5 | 0.3335 | 0.3400 |
| 6 | 0.3419 | 0.3430 |
| 7 | 0.3502 | 0.3478 |
| 8 | 0.3481 | 0.3494 |
| 9 | 0.3502 | 0.3432 |
| 10 | 0.3481 | 0.3448 |

Total THC (%AR) via GC-FID (method 006.40)

Is between-sample SD less than check value? YES Homogeneity test passed

# Stability

Sample in one packet tested on different days from date sample shipped to results due date. Sample stored at room temperature during period of testing.

| Date of Analysis | Days | Conc. |
| --- | --- | --- |
| 9/9/20 | 0 | 0.3365 |
| 9/12/20 | 3 | 0.3015 |
| 9/16/20 | 7 | 0.3018 |
| 9/20/20 | 11 | 0.2920 |
| 9/20/20 | 11 | 0.2795 |
| 9/23/20 | 14 | 0.3077 |
| 9/26/20 | 17 | 0.2880 |
| 9/30/20 | 21 | 0.2711 |
| 10/6/20 | 27 | 0.3128 |
| 10/6/20 | 27 | 0.2882 |
| 10/7/20 | 28 | 0.2682 |
| 10/8/20 | 29 | 0.2786 |
| 10/13/20 | 34 | 0.2981 |
| 10/13/20 | 34 | 0.3150 |

Total THC (%AR) via GC-FID (method 006.40)	SUMMARY OUTPUT

Days = x variable, Conc = y variable

Regression Statistics

| Multiple R | 0.321357 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| R Square | 0.1032703 |  |  |  |  |  |
| Adjusted R Squ | 0.0285429 |  |  |  |  |  |
| Standard Error | 0.0185259 |  |  |  |  |  |
| Observations | 14 |  |  |  |  |  |
| ANOVA |  |  |  |  |  |  |
|  | df | SS | MS | F | Significance F |  |
| Regression | 1 | 0.0004743 | 0.000474301 | 1.381959 | 0.262564945 |  |
| Residual | 12 | 0.00411851 | 0.000343209 |  |  |  |
| Total | 13 | 0.00459281 |  |  |  |  |
|  | Coefficients | tandard Error | t Stat | P-value | Lower 95% | Upper 95% |
| Intercept | 0.3056247 | 0.00982825 | 31.0965611 | 7.69E-13 | 0.284210833 | 0.327038663 |
| X Variable 1 | -0.0005313 | 0.00045194 | -1.17556767 | 0.262565 | -0.00151597 | 0.000453404 |

y = -0.0005x + 0.3056 R² = 0.1033

0.40

0.38

0.36

0.34

0.32

0.30

0.28

0.26

0.24

0.22

0.20

0

10

20

30

40

Total THC (%AR) via GC-FID (method 006.40)

Coefficient for x variable is not statistically significantly different from 0. No signficant change in result during testing period.

Stability test passed.