Hemp Proficiency Testing All Labs Trueness HM21NOV-1 (hemp)

Sample: HM21NOV-1 (hemp)

All Labs Trueness

Issue Date: December 20, 2021

| Analyte | Code | Method | Lab Num | Result1 Result2 Result3 Δ9-THC (%AR) |  |  | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, | 161 | 0.1513 | 0.1377 | 0.1507 | 0.147 | -2.2 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 212 | 0.1424 | 0.1584 | 0.1524 | 0.151 | -2 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 213 | 0.1431 | 0.1643 | 0.1707 | 0.159 | -1.64 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 182 | 0.1622 | 0.1608 | 0.158 | 0.16 | -1.59 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 141 | 0.1693 | 0.1668 | 0.167 | 0.168 | -1.27 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 197 | 0.1642 | 0.1751 | 0.1673 | 0.169 | -1.21 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, | 103 | 0.186 | 0.166 | 0.172 | 0.175 | -0.96 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 109 | 0.198 | 0.159 | 0.179 | 0.179 | -0.78 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 106 | 0.18 | 0.19 | 0.17 | 0.18 | -0.72 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 124 | 0.1339 | 0.2021 | 0.2061 | 0.181 | -0.69 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 208 | 0.185 | 0.1758 | 0.1818 | 0.181 | -0.68 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 142 | 0.1802 | 0.1836 | 0.1801 | 0.181 | -0.66 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 164 | 0.1849 | 0.1749 | 0.1818 | 0.181 | -0.69 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 155 | 0.187 | 0.175 | 0.184 | 0.182 | -0.63 | 0.1962 | 52 | 0.0225 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 113 | 0.1829 | 0.1812 | 0.1835 | 0.183 | -0.61 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 149 | 0.1902 | 0.1824 | 0.1786 | 0.184 | -0.55 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 144 | 0.1895 | 0.1868 | 0.1749 | 0.184 | -0.55 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 162 | 0.1865 | 0.1927 | 0.1796 | 0.186 | -0.44 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.60 | NIR | 187 | 0.18 | 0.19 | 0.19 | 0.187 | -0.42 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 216 | 0.1888 | 0.1854 | 0.1884 | 0.188 | -0.38 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 196 | 0.1918 | 0.1974 | 0.1791 | 0.189 | -0.3 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 125 | 0.185 | 0.19 | 0.193 | 0.189 | -0.3 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 192 | 0.1874 | 0.1847 | 0.1955 | 0.189 | -0.31 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, | 199 | 0.192 | 0.1875 | 0.1935 | 0.191 | -0.23 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 204 | 0.1927 | 0.1937 | 0.1924 | 0.193 | -0.14 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 115 | 0.18 | 0.2 | 0.2 | 0.193 | -0.13 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 191 | 0.1939 | 0.1909 | 0.1971 | 0.194 | -0.1 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 200 | 0.206 | 0.186 | 0.195 | 0.196 | -0.02 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, | 108 | 0.1825 | 0.1841 | 0.2237 | 0.197 | 0.03 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 217 | 0.2017 | 0.1985 | 0.1948 | 0.198 | 0.1 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 194 | 0.1907 | 0.21 | 0.1971 | 0.199 | 0.14 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.99 | Other | 215 | 0.2028 | 0.2023 | 0.1988 | 0.201 | 0.23 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 184 | 0.2069 | 0.1975 | 0.2043 | 0.203 | 0.3 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 120 | 0.2029 | 0.2062 | 0.1997 | 0.203 | 0.3 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 101 | 0.2002 | 0.1954 | 0.2133 | 0.203 | 0.3 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 163 | 0.207 | 0.204 | 0.201 | 0.204 | 0.35 | 0.1962 | 52 | 0.0225 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 211 | 0.2 | 0.206 | 0.208 | 0.205 | 0.38 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.03 | AOAC 2018.11, | 139 | 0.2016 | 0.2132 | 0.201 | 0.205 | 0.4 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.99 | Other | 132 | 0.213 | 0.205 | 0.203 | 0.207 | 0.48 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 175 | 0.208 | 0.2108 | 0.207 | 0.209 | 0.55 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.99 | Other | 183 | 0.227 | 0.205 | 0.209 | 0.214 | 0.78 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 209 | 0.2182 | 0.2076 | 0.2257 | 0.217 | 0.93 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 189 | 0.2215 | 0.2175 | 0.2178 | 0.219 | 1.01 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 102 | 0.217 | 0.214 | 0.227 | 0.219 | 1.03 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, | 152 | 0.2233 | 0.2102 | 0.2325 | 0.222 | 1.15 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 219 | 0.2246 | 0.2318 | 0.2284 | 0.228 | 1.43 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 116 | 0.2273 | 0.2334 | 0.222 | 0.228 | 1.4 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 148 | 0.232 | 0.236 | 0.231 | 0.233 | 1.64 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 172 | 0.226 | 0.233 | 0.243 | 0.234 | 1.68 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 209 | 0.2458 | 0.2239 | 0.231 | 0.234 | 1.66 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, | 209 | 0.2641 | 0.2372 | 0.2224 | 0.241 | 2 | 0.1962 | 52 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 122 | 0.2498 | 0.2524 | 0.2514 | 0.251 | 2.45 | 0.1962 | 52 | 0.0225 | 0 |
|  |  |  |  | Δ9-THCA (%AR) |  |  |  |  |  |  |  |  |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 192 | <0.0300 | <0.0300 | <0.0300 | <0.0300 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 197 | <0.0698 | <0.0688 | <0.0721 | <0.0698 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, | 108 | <0.0700 | <0.0700 | <0.0700 | <0.0700 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 216 | <0.0701 | <0.0732 | <0.0725 | <0.0701 |  |  |  |  | 1 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 213 | 0.0338 | 0.035 | 0.0353 | 0.0347 | -2.43 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 155 | 0.045 | 0.043 | 0.045 | 0.0443 | -1.82 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 141 | 0.0526 | 0.0486 | 0.0542 | 0.0518 | -1.35 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 124 | 0.0395 | 0.0765 | 0.0448 | 0.0536 | -1.24 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.99 | Other | 215 | 0.0548 | 0.0548 | 0.0535 | 0.0544 | -1.19 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 212 | 0.0666 | 0.048 | 0.048 | 0.0542 | -1.2 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 194 | 0.0515 | 0.0582 | 0.0658 | 0.0585 | -0.93 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 182 | 0.0632 | 0.057 | 0.0535 | 0.0579 | -0.97 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 196 | 0.0599 | 0.0649 | 0.0585 | 0.0611 | -0.77 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 115 | 0.063 | 0.062 | 0.065 | 0.0633 | -0.63 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 211 | 0.058 | 0.062 | 0.07 | 0.0633 | -0.63 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 162 | 0.0666 | 0.061 | 0.0611 | 0.0629 | -0.66 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, | 161 | 0.0639 | 0.0611 | 0.0652 | 0.0634 | -0.62 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 142 | 0.0635 | 0.0637 | 0.0616 | 0.0629 | -0.65 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 208 | 0.0614 | 0.0613 | 0.0701 | 0.0643 | -0.57 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 144 | 0.0708 | 0.0658 | 0.0577 | 0.0648 | -0.54 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 163 | 0.075 | 0.06 | 0.062 | 0.0657 | -0.48 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 113 | 0.0681 | 0.0654 | 0.0649 | 0.0661 | -0.45 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 219 | 0.0644 | 0.068 | 0.0672 | 0.0665 | -0.43 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 102 | 0.067 | 0.066 | 0.068 | 0.067 | -0.4 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 109 | 0.078 | 0.061 | 0.069 | 0.0693 | -0.25 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 200 | 0.073 | 0.065 | 0.068 | 0.0687 | -0.29 | 0.0733 | 47 | 0.0159 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 175 | 0.0711 | 0.0705 | 0.0694 | 0.0703 | -0.19 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, | 199 | 0.0745 | 0.0745 | 0.0735 | 0.0742 | 0.05 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.03 | AOAC 2018.11, | 139 | 0.075 | 0.0812 | 0.0695 | 0.0752 | 0.12 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 120 | 0.0736 | 0.0738 | 0.0761 | 0.0745 | 0.08 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 184 | 0.0751 | 0.0724 | 0.0786 | 0.0754 | 0.13 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 101 | 0.0763 | 0.0695 | 0.0811 | 0.0756 | 0.15 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 149 | 0.0796 | 0.0775 | 0.0751 | 0.0774 | 0.26 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.99 | Other | 132 | 0.091 | 0.073 | 0.069 | 0.0777 | 0.28 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 116 | 0.0754 | 0.082 | 0.0784 | 0.0786 | 0.33 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 189 | 0.0804 | 0.0806 | 0.0801 | 0.0804 | 0.45 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, | 103 | 0.07 | 0.08 | 0.09 | 0.08 | 0.42 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 106 | 0.08 | 0.08 | 0.08 | 0.08 | 0.42 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 217 | 0.0821 | 0.0779 | 0.0792 | 0.0797 | 0.41 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 204 | 0.0832 | 0.0848 | 0.0848 | 0.0843 | 0.69 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 191 | 0.0876 | 0.088 | 0.088 | 0.0879 | 0.92 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, | 152 | 0.0885 | 0.0929 | 0.0851 | 0.0888 | 0.98 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 172 | 0.085 | 0.089 | 0.094 | 0.0893 | 1.01 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.99 | Other | 183 | 0.083 | 0.091 | 0.096 | 0.09 | 1.05 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 122 | 0.091 | 0.0916 | 0.091 | 0.0912 | 1.13 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 164 | 0.0924 | 0.0924 | 0.094 | 0.0929 | 1.24 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 209 | 0.0959 | 0.0896 | 0.098 | 0.0945 | 1.34 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 209 | 0.1071 | 0.0996 | 0.0987 | 0.102 | 1.8 | 0.0733 | 47 | 0.0159 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 125 | 0.111 | 0.107 | 0.093 | 0.104 | 1.91 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, | 209 | 0.1151 | 0.1055 | 0.095 | 0.105 | 2.01 | 0.0733 | 47 | 0.0159 | 0 |
| Δ9-THCA (%AR) | 002.60 | NIR | 187 | 0.14 | 0.12 | 0.12 | 0.127 | 3.36 | 0.0733 | 47 | 0.0159 | 0 |
|  |  |  |  | CBD (%AR) |  |  |  |  |  |  |  |  |
| CBD (%AR) | 003.10 | Other LC-UV | 209 | 2.8866 | 2.6157 | 2.6377 | 2.71 | -2.17 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.01 | AOAC 2018.10 | 209 | 2.7299 | 2.7536 | 2.8659 | 2.78 | -1.95 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, | 161 | 3.0085 | 2.7495 | 2.8772 | 2.88 | -1.66 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.60 | NIR | 187 | 2.81 | 3.03 | 3 | 2.95 | -1.45 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 120 | 3.0847 | 3.0442 | 2.732 | 2.95 | -1.43 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, | 209 | 3.3194 | 2.8296 | 2.8851 | 3.01 | -1.26 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 164 | 3.2571 | 3.0222 | 3.0719 | 3.12 | -0.93 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 141 | 3.0866 | 3.2322 | 3.1319 | 3.15 | -0.83 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 189 | 3.1803 | 3.1728 | 3.1561 | 3.17 | -0.77 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 191 | 3.0976 | 3.1041 | 3.3074 | 3.17 | -0.77 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 106 | 3.2 | 3.32 | 3.06 | 3.19 | -0.7 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 155 | 3.215 | 3.11 | 3.3 | 3.21 | -0.65 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 125 | 3.28 | 3.14 | 3.41 | 3.28 | -0.44 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 149 | 3.4159 | 3.1697 | 3.3304 | 3.31 | -0.35 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 162 | 3.4378 | 3.2839 | 3.2786 | 3.33 | -0.27 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.99 | Other | 132 | 3.355 | 3.298 | 3.338 | 3.33 | -0.28 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 109 | 3.702 | 3.02 | 3.379 | 3.37 | -0.17 | 3.421 | 43 | 0.3263 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%AR) | 003.03 | AOAC 2018.11, | 139 | 3.3519 | 3.5161 | 3.2773 | 3.38 | -0.12 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 142 | 3.3056 | 3.4858 | 3.3872 | 3.39 | -0.09 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.99 | Other | 215 | 3.148 | 3.676 | 3.3463 | 3.39 | -0.09 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 200 | 3.316 | 3.373 | 3.598 | 3.43 | 0.02 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 204 | 3.42 | 3.4494 | 3.4422 | 3.44 | 0.05 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 113 | 3.461 | 3.4454 | 3.4513 | 3.45 | 0.1 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 196 | 3.391 | 3.5942 | 3.424 | 3.47 | 0.15 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 184 | 3.427 | 3.394 | 3.644 | 3.49 | 0.21 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.01 | AOAC 2018.10 | 152 | 3.593 | 3.369 | 3.506 | 3.49 | 0.21 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 208 | 3.5697 | 3.4136 | 3.4838 | 3.49 | 0.21 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, | 199 | 3.546 | 3.479 | 3.555 | 3.53 | 0.32 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 212 | 3.4942 | 3.4942 | 3.5879 | 3.53 | 0.32 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, | 103 | 3.565 | 3.536 | 3.487 | 3.53 | 0.33 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 219 | 3.489 | 3.568 | 3.559 | 3.54 | 0.36 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 122 | 3.5447 | 3.5781 | 3.551 | 3.56 | 0.42 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 211 | 3.574 | 3.614 | 3.556 | 3.58 | 0.49 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 194 | 3.5475 | 3.6789 | 3.6623 | 3.63 | 0.64 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 116 | 3.7383 | 3.782 | 3.6229 | 3.71 | 0.9 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.99 | Other | 183 | 3.428 | 3.732 | 4.049 | 3.74 | 0.97 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 197 | 3.7805 | 3.7972 | 3.7106 | 3.76 | 1.05 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 216 | 3.7485 | 3.8546 | 3.8837 | 3.83 | 1.25 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, | 108 | 3.876 | 3.762 | 3.867 | 3.84 | 1.27 | 3.421 | 43 | 0.3263 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%AR) | 003.10 | Other LC-UV | 144 | 4.1507 | 3.8284 | 3.6661 | 3.88 | 1.41 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 163 | 3.888 | 3.865 | 3.91 | 3.89 | 1.43 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 172 | 4.009 | 4.157 | 4.285 | 4.15 | 2.24 | 3.421 | 43 | 0.3263 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 192 | 4.8424 | 4.2807 | 4.3649 | 4.5 | 3.29 | 3.421 | 43 | 0.3263 | 0 |
|  |  |  |  | CBDA (%AR) |  |  |  |  |  |  |  |  |
| CBDA (%AR) | 004.10 | Other LC-UV | 155 | 4.16 | 4.11 | 4.33 | 4.2 | -1.96 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 189 | 4.4224 | 4.4089 | 4.388 | 4.41 | -1.49 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, | 161 | 4.5873 | 4.2378 | 4.4554 | 4.43 | -1.44 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.60 | NIR | 187 | 4.66 | 4.4 | 4.37 | 4.48 | -1.33 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 106 | 4.5 | 4.65 | 4.45 | 4.53 | -1.2 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 212 | 4.704 | 4.6194 | 4.4948 | 4.61 | -1.03 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 116 | 4.7936 | 4.7111 | 4.5687 | 4.69 | -0.84 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 162 | 4.7881 | 4.6009 | 4.6895 | 4.69 | -0.83 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 141 | 4.6529 | 4.8331 | 4.6231 | 4.7 | -0.81 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.99 | Other | 132 | 4.8 | 4.677 | 4.648 | 4.71 | -0.8 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 142 | 4.7106 | 4.805 | 4.6052 | 4.71 | -0.8 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 125 | 4.92 | 4.73 | 4.78 | 4.81 | -0.56 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.03 | AOAC 2018.11, | 139 | 4.8975 | 4.9611 | 4.6319 | 4.83 | -0.52 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, | 152 | 4.988 | 4.636 | 4.918 | 4.85 | -0.48 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 149 | 5.0996 | 4.6144 | 4.8648 | 4.86 | -0.45 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 200 | 4.833 | 4.901 | 5.097 | 4.94 | -0.25 | 5.0545 | 41 | 0.4351 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDA (%AR) | 004.10 | Other LC-UV | 211 | 4.987 | 4.937 | 4.913 | 4.95 | -0.25 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 194 | 4.8611 | 4.9899 | 5.0197 | 4.96 | -0.22 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 144 | 5.4495 | 4.6242 | 4.8197 | 4.96 | -0.21 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 196 | 4.9704 | 5.2969 | 4.7741 | 5.01 | -0.09 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 204 | 5.0711 | 5.1741 | 5.0943 | 5.11 | 0.13 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 191 | 5.0109 | 4.9959 | 5.346 | 5.12 | 0.15 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 164 | 4.9898 | 5.3372 | 5.0749 | 5.13 | 0.18 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 219 | 5.025 | 5.198 | 5.183 | 5.14 | 0.19 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 208 | 5.3241 | 4.9775 | 5.1426 | 5.15 | 0.22 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 197 | 5.0679 | 5.3018 | 5.1144 | 5.16 | 0.25 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, | 103 | 5.366 | 5.072 | 5.083 | 5.17 | 0.27 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, | 108 | 5.334 | 5.091 | 5.3 | 5.24 | 0.43 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 109 | 5.769 | 4.722 | 5.302 | 5.26 | 0.48 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 122 | 5.3495 | 5.3183 | 5.3521 | 5.34 | 0.66 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 216 | 5.3235 | 5.2956 | 5.4415 | 5.35 | 0.69 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, | 199 | 5.3245 | 5.3595 | 5.4605 | 5.38 | 0.75 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.99 | Other | 215 | 5.4365 | 5.438 | 5.303 | 5.39 | 0.78 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 184 | 5.357 | 5.328 | 5.511 | 5.4 | 0.79 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 163 | 5.548 | 5.287 | 5.421 | 5.42 | 0.84 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 192 | 5.6679 | 5.4077 | 5.5019 | 5.53 | 1.08 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.99 | Other | 183 | 5.581 | 5.314 | 5.851 | 5.58 | 1.21 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.01 | AOAC 2018.10 | 209 | 5.5732 | 5.5884 | 5.9577 | 5.71 | 1.5 | 5.0545 | 41 | 0.4351 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDA (%AR) | 004.10 | Other LC-UV | 172 | 5.601 | 5.778 | 6.116 | 5.83 | 1.79 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 209 | 6.2413 | 5.6841 | 5.6256 | 5.85 | 1.83 | 5.0545 | 41 | 0.4351 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, | 209 | 7.2888 | 6.2739 | 6.2087 | 6.59 | 3.53 | 5.0545 | 41 | 0.4351 | 0 |

Population of Lab Values

CBN (%AR)

| CBN (%AR) | 005.10 | Other LC-UV | 106 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%AR) | 005.10 | Other LC-UV | 196 | <0.0205 | <0.0205 | <0.0205 | <0.0205 |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 149 | <0.0240 | <0.0240 | <0.0240 | <0.0240 |  |  |  |  | 1 |
| CBN (%AR) | 005.02 | AOAC 2018.11, | 199 | <0.0250 | <0.0250 | <0.0250 | <0.0250 |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 194 | <0.0270 | <0.0270 | <0.0270 | <0.0270 |  |  |  |  | 1 |
| CBN (%AR) | 005.30 | GC-FID | 116 | <0.0400 | <0.0400 | <0.0400 | <0.0400 |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 197 | <0.0465 | <0.0459 | <0.0481 | <0.0465 |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 216 | <0.0467 | <0.0488 | <0.0483 | <0.0467 |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 208 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| CBN (%AR) | 005.02 | AOAC 2018.11, | 108 | <0.0700 | <0.0700 | <0.0700 | <0.0700 |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 125 | 0.0071 | 0.0087 | 0.0067 | 0.0075 | -1.56 | 0.0174 | 23 | 0.0064 | 0 |
| CBN (%AR) | 005.01 | AOAC 2018.10 | 209 | 0.01 | 0.01 | 0.0116 | 0.0105 | -1.09 | 0.0174 | 23 | 0.0064 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, | 209 | 0.014 | 0.0118 | 0.0111 | 0.0123 | -0.81 | 0.0174 | 23 | 0.0064 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 209 | 0.013 | 0.0111 | 0.0115 | 0.0119 | -0.88 | 0.0174 | 23 | 0.0064 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 219 | 0.0135 | 0.0135 | 0.0134 | 0.0135 | -0.62 | 0.0174 | 23 | 0.0064 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 162 | 0.014 | 0.0135 | 0.0132 | 0.0136 | -0.61 | 0.0174 | 23 | 0.0064 | 0 |
| CBN (%AR) | 005.99 | Other | 132 | 0.014 | 0.013 | 0.014 | 0.0137 | -0.59 | 0.0174 | 23 | 0.0064 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%AR) | 005.10 | Other LC-UV | 142 | 0.0118 | 0.0153 | 0.0145 | 0.0139 | -0.56 | 0.0174 | 23 | 0.0064 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 109 | 0.015 | 0.013 | 0.015 | 0.0143 | -0.49 | 0.0174 | 23 | 0.0064 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 144 | 0.0152 | 0.015 | 0.0153 | 0.0152 | -0.36 | 0.0174 | 23 | 0.0064 | 0 |
| CBN (%AR) | 005.03 | AOAC 2018.11, | 139 | 0.0153 | 0.0164 | 0.0164 | 0.016 | -0.22 | 0.0174 | 23 | 0.0064 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 192 | 0.0176 | 0.0141 | 0.0158 | 0.0158 | -0.25 | 0.0174 | 23 | 0.0064 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 204 | 0.0172 | 0.0171 | 0.0173 | 0.0172 | -0.04 | 0.0174 | 23 | 0.0064 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 200 | 0.019 | 0.017 | 0.017 | 0.0177 | 0.04 | 0.0174 | 23 | 0.0064 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 211 | 0.017 | 0.017 | 0.019 | 0.0177 | 0.04 | 0.0174 | 23 | 0.0064 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 189 | 0.0178 | 0.018 | 0.0178 | 0.0179 | 0.07 | 0.0174 | 23 | 0.0064 | 0 |
| CBN (%AR) | 005.99 | Other | 183 | 0.018 | 0.018 | 0.019 | 0.0183 | 0.14 | 0.0174 | 23 | 0.0064 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 141 | 0.0233 | 0.0232 | 0.0221 | 0.0229 | 0.86 | 0.0174 | 23 | 0.0064 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, | 107 | 0.021 | 0.026 | 0.029 | 0.0253 | 1.24 | 0.0174 | 23 | 0.0064 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 138 | 0.0216 | 0.0277 |  | 0.0247 | 1.14 | 0.0174 | 23 | 0.0064 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 163 | 0.027 | 0.032 | 0.028 | 0.029 | 1.82 | 0.0174 | 23 | 0.0064 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 116 | 0.0413 | 0.0459 | 0.0417 | 0.043 | 4.02 | 0.0174 | 23 | 0.0064 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 185 | 2.533 | 2.438 | 2.7842 | 2.59 | 404.33 | 0.0174 | 23 | 0.0064 | 0 |
|  |  |  |  | Total Δ9-THC (%AR) |  |  |  |  |  |  |  |  |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 197 | 0.1642 | 0.1751 | 0.1673 | 0.169 | -2.33 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 216 | 0.1888 | 0.1854 | 0.1884 | 0.188 | -1.88 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 192 | 0.1874 | 0.1847 | 0.1955 | 0.189 | -1.84 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 155 | 0.203 | 0.191 | 0.2 | 0.198 | -1.63 | 0.2661 | 60 | 0.0418 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, | 161 | 0.2073 | 0.1913 | 0.2079 | 0.202 | -1.53 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 212 | 0.2008 | 0.2005 | 0.2215 | 0.208 | -1.4 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 182 | 0.2176 | 0.2108 | 0.205 | 0.211 | -1.31 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 141 | 0.2154 | 0.2094 | 0.2145 | 0.213 | -1.27 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 124 | 0.1685 | 0.2692 | 0.2454 | 0.228 | -0.92 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 114 | 0.2426 | 0.2264 | 0.2255 | 0.231 | -0.83 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 142 | 0.2359 | 0.2395 | 0.2341 | 0.236 | -0.71 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 208 | 0.2389 | 0.2295 | 0.2432 | 0.237 | -0.69 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 109 | 0.266 | 0.212 | 0.24 | 0.239 | -0.64 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 144 | 0.2516 | 0.2445 | 0.2255 | 0.241 | -0.61 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 162 | 0.245 | 0.2462 | 0.2332 | 0.241 | -0.59 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 138 | 0.2404 | 0.2449 | 0.2413 | 0.242 | -0.57 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 196 | 0.2443 | 0.2544 | 0.2304 | 0.243 | -0.55 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, | 103 | 0.247 | 0.236 | 0.251 | 0.245 | -0.51 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, | 108 | 0.2384 | 0.2231 | 0.2777 | 0.246 | -0.47 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 115 | 0.24 | 0.26 | 0.25 | 0.25 | -0.39 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 194 | 0.2359 | 0.261 | 0.2547 | 0.251 | -0.37 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.99 | Other | 215 | 0.2508 | 0.2503 | 0.2547 | 0.252 | -0.34 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 153 | 0.2793 | 0.2419 | 0.2362 | 0.252 | -0.33 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 149 | 0.26 | 0.2503 | 0.2445 | 0.252 | -0.35 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 200 | 0.27 | 0.24 | 0.25 | 0.253 | -0.31 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 169 | 0.2476 | 0.2553 | 0.2602 | 0.254 | -0.28 | 0.2661 | 60 | 0.0418 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, | 199 | 0.2573 | 0.2528 | 0.258 | 0.256 | -0.24 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 211 | 0.251 | 0.26 | 0.269 | 0.26 | -0.15 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 116 | 0.2614 | 0.2575 | 0.2645 | 0.261 | -0.12 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 164 | 0.266 | 0.2559 | 0.2642 | 0.262 | -0.1 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 204 | 0.2657 | 0.2681 | 0.2641 | 0.266 | 0 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 120 | 0.2675 | 0.2709 | 0.2664 | 0.268 | 0.05 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 217 | 0.2737 | 0.2668 | 0.2642 | 0.268 | 0.05 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 101 | 0.2671 | 0.2564 | 0.2844 | 0.269 | 0.08 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 175 | 0.2704 | 0.2726 | 0.2679 | 0.27 | 0.1 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.03 | AOAC 2018.11, | 139 | 0.2674 | 0.2844 | 0.262 | 0.271 | 0.12 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 191 | 0.2707 | 0.268 | 0.2742 | 0.271 | 0.12 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.99 | Other | 132 | 0.293 | 0.269 | 0.264 | 0.275 | 0.22 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 102 | 0.276 | 0.272 | 0.287 | 0.278 | 0.29 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 125 | 0.282 | 0.284 | 0.275 | 0.28 | 0.34 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 121 | 0.283 | 0.284 | 0.282 | 0.283 | 0.4 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 219 | 0.281 | 0.2914 | 0.2873 | 0.287 | 0.49 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 189 | 0.292 | 0.2882 | 0.2881 | 0.289 | 0.56 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.60 | NIR | 187 | 0.29 | 0.29 | 0.29 | 0.29 | 0.57 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.99 | Other | 183 | 0.3 | 0.285 | 0.293 | 0.293 | 0.63 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 116 | 0.2934 | 0.3053 | 0.2907 | 0.296 | 0.73 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 128 | 0.3113 | 0.288 | 0.2887 | 0.296 | 0.71 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 163 | 0.294 | 0.299 | 0.302 | 0.298 | 0.77 | 0.2661 | 60 | 0.0418 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 209 | 0.3024 | 0.2862 | 0.3116 | 0.3 | 0.81 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, | 152 | 0.3009 | 0.2926 | 0.3149 | 0.303 | 0.88 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 170 | 0.3125 | 0.2997 | 0.3127 | 0.308 | 1.01 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 172 | 0.3005 | 0.3111 | 0.3254 | 0.312 | 1.1 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, | 107 | 0.286 | 0.332 | 0.332 | 0.317 | 1.21 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 209 | 0.3398 | 0.3112 | 0.3175 | 0.323 | 1.36 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.99 | Other | 181 | 0.34 | 0.32 | 0.32 | 0.327 | 1.45 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 122 | 0.3296 | 0.3328 | 0.3312 | 0.331 | 1.56 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, | 209 | 0.365 | 0.3298 | 0.3058 | 0.334 | 1.61 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 190 | 0.368 | 0.351 | 0.358 | 0.359 | 2.22 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.60 | NIR | 138 | 0.375 | 0.415 | 0.365 | 0.385 | 2.84 | 0.2661 | 60 | 0.0418 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 185 | 0.4071 | 0.5114 | 0.5994 | 0.506 | 5.73 | 0.2661 | 60 | 0.0418 | 0 |
|  |  |  |  | Total CBD (%AR) |  |  |  |  |  |  |  |  |
| Total CBD (%AR) | 007.10 | Other LC-UV | 155 | 6.47 | 6.33 | 6.69 | 6.5 | -2.64 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 116 | 6.1166 | 7.5183 | 6.4532 | 6.7 | -2.28 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, | 161 | 7.0316 | 6.4661 | 6.7846 | 6.76 | -2.16 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.60 | NIR | 187 | 6.89 | 6.89 | 6.83 | 6.87 | -1.96 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 189 | 7.0587 | 7.0393 | 7.0044 | 7.03 | -1.66 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 141 | 7.1672 | 7.4709 | 7.1863 | 7.27 | -1.21 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 162 | 7.6369 | 7.3188 | 7.3913 | 7.45 | -0.89 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.99 | Other | 132 | 7.568 | 7.404 | 7.418 | 7.46 | -0.87 | 7.9354 | 46 | 0.5444 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%AR) | 007.10 | Other LC-UV | 125 | 7.6 | 7.29 | 7.61 | 7.5 | -0.8 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 142 | 7.4368 | 7.6998 | 7.426 | 7.52 | -0.76 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 149 | 7.8883 | 7.2165 | 7.5969 | 7.57 | -0.68 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 212 | 7.6196 | 7.6873 | 7.5298 | 7.61 | -0.59 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.03 | AOAC 2018.11, | 139 | 7.647 | 7.867 | 7.3395 | 7.62 | -0.58 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 164 | 7.6332 | 7.7029 | 7.5226 | 7.62 | -0.58 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 191 | 7.4921 | 7.4855 | 7.9958 | 7.66 | -0.51 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 200 | 7.55 | 7.67 | 8.06 | 7.76 | -0.32 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.01 | AOAC 2018.10 | 209 | 7.6177 | 7.6546 | 8.0907 | 7.79 | -0.27 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 116 | 7.9423 | 7.9137 | 7.6296 | 7.83 | -0.2 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 209 | 8.3603 | 7.6006 | 7.5713 | 7.84 | -0.17 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, | 152 | 7.967 | 7.677 | 7.949 | 7.86 | -0.13 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 196 | 7.7501 | 8.2396 | 7.6109 | 7.87 | -0.13 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 204 | 7.8674 | 7.9871 | 7.9099 | 7.92 | -0.03 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 211 | 7.948 | 7.944 | 7.865 | 7.92 | -0.03 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 194 | 7.8107 | 8.0551 | 8.0645 | 7.98 | 0.08 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 109 | 8.762 | 7.161 | 8.029 | 7.98 | 0.09 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 208 | 8.2389 | 7.7788 | 7.9939 | 8 | 0.13 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 219 | 7.896 | 8.127 | 8.104 | 8.04 | 0.2 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 163 | 8.049 | 7.967 | 8.151 | 8.06 | 0.22 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, | 103 | 8.271 | 7.984 | 7.945 | 8.07 | 0.24 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 170 | 8.0786 | 8.1882 | 8.0403 | 8.1 | 0.31 | 7.9354 | 46 | 0.5444 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%AR) | 007.99 | Other | 215 | 7.9158 | 8.4451 | 7.997 | 8.12 | 0.34 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, | 107 | 7.689 | 8.299 | 8.473 | 8.15 | 0.4 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 169 | 8.0018 | 8.2198 | 8.3313 | 8.18 | 0.46 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 122 | 8.2362 | 8.2422 | 8.2447 | 8.24 | 0.56 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 144 | 8.931 | 7.8847 | 7.8939 | 8.24 | 0.55 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, | 199 | 8.2156 | 8.1793 | 8.3439 | 8.25 | 0.57 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 197 | 8.225 | 8.4469 | 8.1959 | 8.29 | 0.65 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.60 | NIR | 138 | 8.355 | 8.49 | 8.085 | 8.31 | 0.69 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, | 108 | 8.554 | 8.227 | 8.515 | 8.43 | 0.91 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 216 | 8.4172 | 8.4988 | 8.6559 | 8.52 | 1.08 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.99 | Other | 183 | 8.322 | 8.392 | 9.18 | 8.63 | 1.28 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, | 209 | 9.7117 | 8.3318 | 8.3301 | 8.79 | 1.57 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 138 | 9.0779 | 9.0648 | 8.7777 | 8.97 | 1.91 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 172 | 8.9211 | 9.2243 | 9.6487 | 9.26 | 2.44 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 192 | 9.8131 | 9.0232 | 9.19 | 9.34 | 2.58 | 7.9354 | 46 | 0.5444 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 185 | 12.1579 | 12.5202 | 12.5896 | 12.4 | 8.24 | 7.9354 | 46 | 0.5444 | 0 |
|  |  |  |  | CBDV (%AR) |  |  |  |  |  |  |  |  |
| CBDV (%AR) | 008.99 | Other | 194 | <0.0270 | <0.0270 | <0.0270 | <0.0270 |  |  |  |  | 1 |
| CBDV (%AR) | 008.99 | Other | 197 | <0.0465 | <0.0459 | <0.0481 | <0.0465 |  |  |  |  | 1 |
| CBDV (%AR) | 008.99 | Other | 216 | <0.0467 | <0.0488 | <0.0483 | <0.0467 |  |  |  |  | 1 |
| CBDV (%AR) | 008.02 | AOAC 2018.11, | 108 | <0.0700 | <0.0700 | <0.0700 | <0.0700 |  |  |  |  | 1 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDV (%AR) | 008.99 | Other | 189 | 0.0218 | 0.0214 | 0.0208 | 0.0213 | -1.01 | 0.0364 | 12 | 0.0149 | 0 |
| CBDV (%AR) | 008.99 | Other | 162 | 0.0238 | 0.0225 | 0.0218 | 0.0227 | -0.92 | 0.0364 | 12 | 0.0149 | 0 |
| CBDV (%AR) | 008.99 | Other | 196 | 0.0226 | 0.0261 | 0.0235 | 0.0241 | -0.83 | 0.0364 | 12 | 0.0149 | 0 |
| CBDV (%AR) | 008.99 | Other | 132 | 0.026 | 0.026 | 0.035 | 0.029 | -0.5 | 0.0364 | 12 | 0.0149 | 0 |
| CBDV (%AR) | 008.99 | Other | 219 | 0.0275 | 0.0285 | 0.0302 | 0.0287 | -0.52 | 0.0364 | 12 | 0.0149 | 0 |
| CBDV (%AR) | 008.99 | Other | 211 | 0.03 | 0.028 | 0.036 | 0.0313 | -0.34 | 0.0364 | 12 | 0.0149 | 0 |
| CBDV (%AR) | 008.99 | Other | 142 | 0.0347 | 0.0375 | 0.0348 | 0.0357 | -0.05 | 0.0364 | 12 | 0.0149 | 0 |
| CBDV (%AR) | 008.99 | Other | 200 | 0.039 | 0.035 | 0.037 | 0.037 | 0.04 | 0.0364 | 12 | 0.0149 | 0 |
| CBDV (%AR) | 008.99 | Other | 183 | 0.039 | 0.039 | 0.041 | 0.0397 | 0.22 | 0.0364 | 12 | 0.0149 | 0 |
| CBDV (%AR) | 008.99 | Other | 106 | 0.07 | 0.04 | 0.04 | 0.05 | 0.91 | 0.0364 | 12 | 0.0149 | 0 |
| CBDV (%AR) | 008.99 | Other | 208 | 0.0523 | <0.0500 | <0.0500 | 0.0508 |  |  |  |  | 1 |
| CBDV (%AR) | 008.99 | Other | 192 | 0.0817 | 0.0883 | 0.0949 | 0.0883 | 3.47 | 0.0364 | 12 | 0.0149 | 0 |
| CBDV (%AR) | 008.99 | Other | 109 | 0.11 | 0.088 | 0.101 | 0.0997 | 4.23 | 0.0364 | 12 | 0.0149 | 0 |

Population of Lab Values

CBDVA (%AR)

| CBDVA (%AR) | 009.99 | Other | 197 | <0.0698 | <0.0688 | <0.0721 | <0.0698 |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDVA (%AR) | 009.02 | AOAC 2018.11, | 108 | <0.0700 | <0.0700 | <0.0700 | <0.0700 |  |  |  |  | 1 |
| CBDVA (%AR) | 009.99 | Other | 216 | <0.0701 | <0.0732 | <0.0725 | <0.0701 |  |  |  |  | 1 |
| CBDVA (%AR) | 009.99 | Other | 212 | 0.0369 | 0.0369 | 0.0494 | 0.0411 | -1.21 | 0.0552 | 6 | 0.0117 | 0 |
| CBDVA (%AR) | 009.99 | Other | 196 | 0.0462 | 0.0495 | 0.0459 | 0.0472 | -0.68 | 0.0552 | 6 | 0.0117 | 0 |
| CBDVA (%AR) | 009.99 | Other | 208 | <0.0500 | 0.0573 | 0.0516 | 0.053 | -0.19 | 0.0552 | 6 | 0.0117 | 0 |
| CBDVA (%AR) | 009.99 | Other | 219 | 0.0581 | 0.0588 | 0.0591 | 0.0587 | 0.3 | 0.0552 | 6 | 0.0117 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDVA (%AR) | 009.99 | Other | 149 | 0.0623 | 0.061 | 0.0618 | 0.0617 | 0.55 | 0.0552 | 6 | 0.0117 | 0 |
| CBDVA (%AR) | 009.99 | Other | 183 | 0.077 | 0.065 | 0.067 | 0.0697 | 1.24 | 0.0552 | 6 | 0.0117 | 0 |
|  |  |  |  | CBG (%AR) |  |  |  |  |  |  |  |  |
| CBG (%AR) | 010.99 | Other | 212 | 0.0464 | 0.0547 | 0.0471 | 0.0494 | -2.59 | 0.1294 | 24 | 0.0309 | 0 |
| CBG (%AR) | 010.99 | Other | 216 | 0.0831 | 0.0797 | 0.0741 | 0.079 | -1.63 | 0.1294 | 24 | 0.0309 | 0 |
| CBG (%AR) | 010.99 | Other | 197 | 0.0874 | 0.0844 | 0.0865 | 0.0861 | -1.4 | 0.1294 | 24 | 0.0309 | 0 |
| CBG (%AR) | 010.99 | Other | 162 | 0.1003 | 0.1051 | 0.105 | 0.103 | -0.84 | 0.1294 | 24 | 0.0309 | 0 |
| CBG (%AR) | 010.99 | Other | 196 | 0.1031 | 0.1106 | 0.1043 | 0.106 | -0.76 | 0.1294 | 24 | 0.0309 | 0 |
| CBG (%AR) | 010.02 | AOAC 2018.11, | 108 | 0.1065 | 0.106 | 0.1108 | 0.108 | -0.7 | 0.1294 | 24 | 0.0309 | 0 |
| CBG (%AR) | 010.02 | AOAC 2018.11, | 103 | 0.117 | 0.108 | 0.117 | 0.114 | -0.5 | 0.1294 | 24 | 0.0309 | 0 |
| CBG (%AR) | 010.99 | Other | 155 | 0.119 | 0.116 | 0.121 | 0.119 | -0.35 | 0.1294 | 24 | 0.0309 | 0 |
| CBG (%AR) | 010.99 | Other | 109 | 0.125 | 0.114 | 0.118 | 0.119 | -0.34 | 0.1294 | 24 | 0.0309 | 0 |
| CBG (%AR) | 010.99 | Other | 149 | 0.1289 | 0.1245 | 0.106 | 0.12 | -0.31 | 0.1294 | 24 | 0.0309 | 0 |
| CBG (%AR) | 010.99 | Other | 142 | 0.1146 | 0.1283 | 0.1234 | 0.122 | -0.24 | 0.1294 | 24 | 0.0309 | 0 |
| CBG (%AR) | 010.99 | Other | 208 | 0.1235 | 0.1199 | 0.1313 | 0.125 | -0.15 | 0.1294 | 24 | 0.0309 | 0 |
| CBG (%AR) | 010.99 | Other | 211 | 0.121 | 0.13 | 0.143 | 0.131 | 0.06 | 0.1294 | 24 | 0.0309 | 0 |
| CBG (%AR) | 010.99 | Other | 200 | 0.142 | 0.128 | 0.136 | 0.135 | 0.19 | 0.1294 | 24 | 0.0309 | 0 |
| CBG (%AR) | 010.99 | Other | 219 | 0.1355 | 0.1378 | 0.1368 | 0.137 | 0.24 | 0.1294 | 24 | 0.0309 | 0 |
| CBG (%AR) | 010.99 | Other | 106 | 0.14 | 0.15 | 0.13 | 0.14 | 0.34 | 0.1294 | 24 | 0.0309 | 0 |
| CBG (%AR) | 010.99 | Other | 189 | 0.1451 | 0.1456 | 0.1469 | 0.146 | 0.53 | 0.1294 | 24 | 0.0309 | 0 |
| CBG (%AR) | 010.99 | Other | 183 | 0.138 | 0.155 | 0.144 | 0.146 | 0.53 | 0.1294 | 24 | 0.0309 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%AR) | 010.99 | Other | 204 | 0.1411 | 0.1493 | 0.1474 | 0.146 | 0.53 | 0.1294 | 24 | 0.0309 | 0 |
| CBG (%AR) | 010.99 | Other | 125 | 0.157 | 0.149 | 0.164 | 0.157 | 0.88 | 0.1294 | 24 | 0.0309 | 0 |
| CBG (%AR) | 010.99 | Other | 194 | 0.1515 | 0.1627 | 0.1615 | 0.159 | 0.94 | 0.1294 | 24 | 0.0309 | 0 |
| CBG (%AR) | 010.99 | Other | 132 | 0.179 | 0.166 | 0.167 | 0.171 | 1.33 | 0.1294 | 24 | 0.0309 | 0 |
| CBG (%AR) | 010.99 | Other | 192 | 0.2078 | 0.1975 | 0.1978 | 0.201 | 2.32 | 0.1294 | 24 | 0.0309 | 0 |
| CBG (%AR) | 010.99 | Other | 141 | 0.2177 | 0.1925 | 0.2155 | 0.209 | 2.56 | 0.1294 | 24 | 0.0309 | 0 |
|  |  |  |  | CBGA (%AR) |  |  |  |  |  |  |  |  |
| CBGA (%AR) | 011.99 | Other | 155 | 0.16 | 0.156 | 0.153 | 0.156 | -2.92 | 0.2366 | 23 | 0.0275 | 0 |
| CBGA (%AR) | 011.99 | Other | 212 | 0.2107 | 0.1842 | 0.1839 | 0.193 | -1.59 | 0.2366 | 23 | 0.0275 | 0 |
| CBGA (%AR) | 011.99 | Other | 141 | 0.2135 | 0.2098 | 0.2016 | 0.208 | -1.03 | 0.2366 | 23 | 0.0275 | 0 |
| CBGA (%AR) | 011.99 | Other | 106 | 0.21 | 0.22 | 0.2 | 0.21 | -0.97 | 0.2366 | 23 | 0.0275 | 0 |
| CBGA (%AR) | 011.99 | Other | 132 | 0.224 | 0.214 | 0.204 | 0.214 | -0.82 | 0.2366 | 23 | 0.0275 | 0 |
| CBGA (%AR) | 011.99 | Other | 216 | 0.215 | 0.2146 | 0.2164 | 0.215 | -0.78 | 0.2366 | 23 | 0.0275 | 0 |
| CBGA (%AR) | 011.02 | AOAC 2018.11, | 103 | 0.255 | 0.204 | 0.191 | 0.217 | -0.73 | 0.2366 | 23 | 0.0275 | 0 |
| CBGA (%AR) | 011.02 | AOAC 2018.11, | 108 | 0.2132 | 0.2144 | 0.2342 | 0.221 | -0.58 | 0.2366 | 23 | 0.0275 | 0 |
| CBGA (%AR) | 011.99 | Other | 197 | 0.2307 | 0.2392 | 0.2221 | 0.231 | -0.22 | 0.2366 | 23 | 0.0275 | 0 |
| CBGA (%AR) | 011.99 | Other | 192 | 0.2517 | 0.2266 | 0.2279 | 0.235 | -0.04 | 0.2366 | 23 | 0.0275 | 0 |
| CBGA (%AR) | 011.99 | Other | 125 | 0.233 | 0.234 | 0.245 | 0.237 | 0.03 | 0.2366 | 23 | 0.0275 | 0 |
| CBGA (%AR) | 011.99 | Other | 208 | 0.2416 | 0.2304 | 0.2426 | 0.238 | 0.06 | 0.2366 | 23 | 0.0275 | 0 |
| CBGA (%AR) | 011.99 | Other | 196 | 0.2382 | 0.2476 | 0.2322 | 0.239 | 0.1 | 0.2366 | 23 | 0.0275 | 0 |
| CBGA (%AR) | 011.99 | Other | 219 | 0.2358 | 0.2435 | 0.2417 | 0.24 | 0.14 | 0.2366 | 23 | 0.0275 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%AR) | 011.99 | Other | 189 | 0.2468 | 0.2465 | 0.2451 | 0.246 | 0.35 | 0.2366 | 23 | 0.0275 | 0 |
| CBGA (%AR) | 011.99 | Other | 149 | 0.2842 | 0.2605 | 0.1994 | 0.248 | 0.42 | 0.2366 | 23 | 0.0275 | 0 |
| CBGA (%AR) | 011.99 | Other | 211 | 0.244 | 0.252 | 0.257 | 0.251 | 0.52 | 0.2366 | 23 | 0.0275 | 0 |
| CBGA (%AR) | 011.99 | Other | 204 | 0.2458 | 0.2604 | 0.2498 | 0.252 | 0.56 | 0.2366 | 23 | 0.0275 | 0 |
| CBGA (%AR) | 011.99 | Other | 142 | 0.2494 | 0.2702 | 0.2701 | 0.263 | 0.97 | 0.2366 | 23 | 0.0275 | 0 |
| CBGA (%AR) | 011.99 | Other | 194 | 0.2544 | 0.2558 | 0.2807 | 0.264 | 0.98 | 0.2366 | 23 | 0.0275 | 0 |
| CBGA (%AR) | 011.99 | Other | 162 | 0.2915 | 0.2407 | 0.2682 | 0.267 | 1.1 | 0.2366 | 23 | 0.0275 | 0 |
| CBGA (%AR) | 011.99 | Other | 109 | 0.303 | 0.26 | 0.267 | 0.277 | 1.46 | 0.2366 | 23 | 0.0275 | 0 |
| CBGA (%AR) | 011.99 | Other | 183 | 0.311 | 0.265 | 0.287 | 0.288 | 1.86 | 0.2366 | 23 | 0.0275 | 0 |

Population of Lab Values

THCV (%AR)

| THCV (%AR) | 012.99 | Other | 183 | <0.0020 | <0.0020 | <0.0020 | <0.0020 | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| THCV (%AR) | 012.99 | Other | 200 | <0.0050 | <0.0050 | <0.0050 | <0.0050 | 1 |
| THCV (%AR) | 012.99 | Other | 189 | <0.0050 | <0.0050 | <0.0050 | <0.0050 | 1 |
| THCV (%AR) | 012.99 | Other | 219 | <0.0060 | <0.0060 | <0.0060 | <0.0060 | 1 |
| THCV (%AR) | 012.99 | Other | 211 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| THCV (%AR) | 012.99 | Other | 142 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| THCV (%AR) | 012.99 | Other | 109 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| THCV (%AR) | 012.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 | <0.0205 | 1 |
| THCV (%AR) | 012.99 | Other | 149 | <0.0240 | <0.0240 | <0.0240 | <0.0240 | 1 |
| THCV (%AR) | 012.99 | Other | 194 | <0.0270 | <0.0270 | <0.0270 | <0.0270 | 1 |
| THCV (%AR) | 012.99 | Other | 197 | <0.0465 | <0.0459 | <0.0481 | <0.0465 | 1 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| THCV (%AR) | 012.99 | Other | 216 | <0.0467 | <0.0488 | <0.0483 | <0.0467 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 208 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| THCV (%AR) | 012.02 | AOAC 2018.11, | 108 | 0 | 0 | 0 | 0 |  |  |  |  | 2 |
| THCV (%AR) | 012.99 | Other | 132 | 0.023 | 0.005 | 0.015 | 0.0143 |  |  |  |  | 0 |
| THCV (%AR) | 012.99 | Other | 106 | <0.0100 | 0.02 | 0.02 | 0.0167 |  |  |  |  | 0 |
|  |  |  |  | Δ8-THC (%AR) |  |  |  |  |  |  |  |  |
| Δ8-THC (%AR) | 013.99 | Other | 183 | <0.0010 | <0.0010 | <0.0010 | <0.0010 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 219 | <0.0060 | <0.0060 | <0.0060 | <0.0060 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 109 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 142 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 200 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 106 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 211 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 141 | <0.0120 | <0.0120 | <0.0120 | <0.0120 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 102 | <0.0200 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 | <0.0205 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 149 | <0.0240 | <0.0240 | <0.0240 | <0.0240 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 194 | <0.0270 | <0.0270 | <0.0270 | <0.0270 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 192 | <0.0300 | <0.0300 | <0.0300 | <0.0300 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 197 | <0.0465 | <0.0459 | <0.0481 | <0.0465 |  |  |  |  | 1 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%AR) | 013.99 | Other | 216 | <0.0467 | <0.0488 | <0.0483 | <0.0467 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 208 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.02 | AOAC 2018.11, | 108 | 0 | 0 | 0 | 0 |  |  |  |  | 2 |
| Δ8-THC (%AR) | 013.99 | Other | 189 | 0.0084 | 0.0085 | 0.0085 | 0.0085 |  |  |  |  | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 138 | 0.0359 | 0.035 |  | 0.0355 |  |  |  |  | 0 |

Population of Lab Values

CBC (%AR)

| CBC (%AR) | 014.99 | Other | 155 | 0.198 | 0.193 | 0.204 | 0.198 | -2.38 | 0.2403 | 23 | 0.0177 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%AR) | 014.99 | Other | 212 | 0.2043 | 0.1986 | 0.208 | 0.204 | -2.08 | 0.2403 | 23 | 0.0177 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, | 103 | 0.228 | 0.214 | 0.232 | 0.225 | -0.89 | 0.2403 | 23 | 0.0177 | 0 |
| CBC (%AR) | 014.99 | Other | 142 | 0.22 | 0.2311 | 0.223 | 0.225 | -0.88 | 0.2403 | 23 | 0.0177 | 0 |
| CBC (%AR) | 014.99 | Other | 106 | 0.22 | 0.24 | 0.22 | 0.227 | -0.77 | 0.2403 | 23 | 0.0177 | 0 |
| CBC (%AR) | 014.99 | Other | 208 | 0.2291 | 0.2227 | 0.2378 | 0.23 | -0.59 | 0.2403 | 23 | 0.0177 | 0 |
| CBC (%AR) | 014.99 | Other | 132 | 0.236 | 0.226 | 0.228 | 0.23 | -0.58 | 0.2403 | 23 | 0.0177 | 0 |
| CBC (%AR) | 014.99 | Other | 149 | 0.2381 | 0.2308 | 0.2299 | 0.233 | -0.42 | 0.2403 | 23 | 0.0177 | 0 |
| CBC (%AR) | 014.99 | Other | 197 | 0.2326 | 0.2422 | 0.2269 | 0.234 | -0.36 | 0.2403 | 23 | 0.0177 | 0 |
| CBC (%AR) | 014.99 | Other | 125 | 0.227 | 0.227 | 0.249 | 0.234 | -0.34 | 0.2403 | 23 | 0.0177 | 0 |
| CBC (%AR) | 014.99 | Other | 162 | 0.2378 | 0.2405 | 0.2335 | 0.237 | -0.17 | 0.2403 | 23 | 0.0177 | 0 |
| CBC (%AR) | 014.99 | Other | 189 | 0.2389 | 0.2382 | 0.237 | 0.238 | -0.13 | 0.2403 | 23 | 0.0177 | 0 |
| CBC (%AR) | 014.99 | Other | 109 | 0.268 | 0.212 | 0.239 | 0.24 | -0.04 | 0.2403 | 23 | 0.0177 | 0 |
| CBC (%AR) | 014.99 | Other | 216 | 0.2439 | 0.242 | 0.2396 | 0.242 | 0.09 | 0.2403 | 23 | 0.0177 | 0 |
| CBC (%AR) | 014.99 | Other | 196 | 0.2408 | 0.2599 | 0.2352 | 0.245 | 0.28 | 0.2403 | 23 | 0.0177 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%AR) | 014.99 | Other | 194 | 0.239 | 0.2474 | 0.2523 | 0.246 | 0.33 | 0.2403 | 23 | 0.0177 | 0 |
| CBC (%AR) | 014.99 | Other | 219 | 0.246 | 0.2495 | 0.2485 | 0.248 | 0.43 | 0.2403 | 23 | 0.0177 | 0 |
| CBC (%AR) | 014.99 | Other | 183 | 0.247 | 0.251 | 0.267 | 0.255 | 0.83 | 0.2403 | 23 | 0.0177 | 0 |
| CBC (%AR) | 014.99 | Other | 204 | 0.2538 | 0.2555 | 0.2551 | 0.255 | 0.82 | 0.2403 | 23 | 0.0177 | 0 |
| CBC (%AR) | 014.99 | Other | 211 | 0.254 | 0.255 | 0.259 | 0.256 | 0.89 | 0.2403 | 23 | 0.0177 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, | 108 | 0.313 | 0.2531 | 0.2718 | 0.279 | 2.21 | 0.2403 | 23 | 0.0177 | 0 |
| CBC (%AR) | 014.99 | Other | 200 | 0.302 | 0.27 | 0.285 | 0.286 | 2.57 | 0.2403 | 23 | 0.0177 | 0 |
| CBC (%AR) | 014.99 | Other | 192 | 0.3679 | 0.3336 | 0.3363 | 0.346 | 5.98 | 0.2403 | 23 | 0.0177 | 0 |

Population of Lab Values

Cd (ug/g AR)

| Cd (ug/g AR) | 101.04 | ICP-MS | 183 | <0.0200 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cd (ug/g AR) | 101.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 | <0.2000 |  |  |  |  | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 141 | <0.2000 | <0.2000 | <0.2000 | <0.2000 |  |  |  |  | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 196 | 0.0893 | 0.105 | 0.0929 | 0.0957 | -2.82 | 0.15 | 7 | 0.0192 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 197 | 0.1399 | 0.1434 | 0.1384 | 0.141 | -0.49 | 0.15 | 7 | 0.0192 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 106 | 0.147 | 0.139 | 0.143 | 0.143 | -0.36 | 0.15 | 7 | 0.0192 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 194 | 0.1476 | 0.1479 | 0.1501 | 0.149 | -0.07 | 0.15 | 7 | 0.0192 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 142 | 0.1724 | 0.156 | 0.1554 | 0.161 | 0.59 | 0.15 | 7 | 0.0192 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 211 | 0.176 | 0.164 | 0.16 | 0.167 | 0.87 | 0.15 | 7 | 0.0192 | 0 |
| Cd (ug/g AR) | 101.03 | ICP-OES | 132 | 0.168 |  |  | 0.168 |  |  |  |  | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 122 | 0.1695 | 0.1677 | 0.1686 | 0.169 | 0.97 | 0.15 | 7 | 0.0192 | 0 |

|  |  |  | Lab |  |  |  | Lab |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Analyte | Code | Method | Num | Result1 Result2 Result3 |  |  | Value | Z score |  | Rob Mean | n | Rob StDev | Flag |
|  |  |  |  | As (ug/g AR) |  |  |  |  |  |  |  |  |  |
| As (ug/g AR) | 102.04 | ICP-MS | 183 | <0.0200 | <0.0200 | <0.0200 | <0.0200 | 1 |  |  |  |  |  |
| As (ug/g AR) | 102.04 | ICP-MS | 122 | <0.0500 | <0.0500 | <0.0500 | <0.0500 | 1 |  |  |  |  |  |
| As (ug/g AR) | 102.04 | ICP-MS | 197 | <0.0998 | <0.1018 | <0.1002 | <0.0998 | 1 |  |  |  |  |  |
| As (ug/g AR) | 102.04 | ICP-MS | 211 | <0.1000 | <0.1000 | <0.1000 | <0.1000 | 1 |  |  |  |  |  |
| As (ug/g AR) | 102.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 | <0.2000 | 1 |  |  |  |  |  |
| As (ug/g AR) | 102.04 | ICP-MS | 141 | <0.2000 | <0.2000 | <0.2000 | <0.2000 | 1 |  |  |  |  |  |
| As (ug/g AR) | 102.04 | ICP-MS | 196 | 0.0712 | 0.0756 | 0.0747 | 0.0738 | 0 |  |  |  |  |  |
| As (ug/g AR) | 102.04 | ICP-MS | 194 | 0.0779 | 0.0725 | 0.0804 | 0.0769 | 0 |  |  |  |  |  |
| As (ug/g AR) | 102.04 | ICP-MS | 106 | <0.0800 | 0.09 | 0.09 | 0.0867 | 0 |  |  |  |  |  |
| As (ug/g AR) | 102.04 | ICP-MS | 142 | 0.1025 | 0.0858 | 0.0742 | 0.0875 | 0 |  |  |  |  |  |
| As (ug/g AR) | 102.03 | ICP-OES | 132 | 0.4 |  |  | 0.4 | 1 |  |  |  |  |  |
|  |  |  |  | Hg (ug/g AR) |  |  |  |  |  |  |  |  |  |
| Hg (ug/g AR) | 103.04 | ICP-MS | 194 | <0.0020 | <0.0020 | <0.0020 | <0.0020 | 1 |  |  |  |  |  |
| Hg (ug/g AR) | 103.04 | ICP-MS | 183 | <0.0500 | <0.0500 | <0.0500 | <0.0500 | 1 |  |  |  |  |  |
| Hg (ug/g AR) | 103.04 | ICP-MS | 122 | <0.0500 | <0.0500 | <0.0500 | <0.0500 | 1 |  |  |  |  |  |
| Hg (ug/g AR) | 103.04 | ICP-MS | 197 | <0.0998 | <0.1018 | <0.1002 | <0.0998 | 1 |  |  |  |  |  |
| Hg (ug/g AR) | 103.04 | ICP-MS | 141 | <0.1000 | <0.1000 | <0.1000 | <0.1000 | 1 |  |  |  |  |  |
| Hg (ug/g AR) | 103.04 | ICP-MS | 211 | <0.1000 | <0.1000 | <0.1000 | <0.1000 | 1 |  |  |  |  |  |
| Hg (ug/g AR) | 103.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 | <0.2000 | 1 |  |  |  |  |  |
| Hg (ug/g AR) | 103.04 | ICP-MS | 106 | 0.007 | 0.007 | 0.008 | 0.0073 | 0 |  |  |  |  |  |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Hg (ug/g AR) | 103.04 | ICP-MS | 196 | 0.008 | 0.0081 | 0.0084 | 0.0082 |  |  |  |  | 0 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 142 | 0.0104 | 0.0092 | 0.0093 | 0.0096 |  |  |  |  | 0 |

Population of Lab Values

Pb (ug/g AR)

| Pb (ug/g AR) | 104.04 | ICP-MS | 141 | <0.5000 | <0.5000 | <0.5000 | <0.5000 |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Pb (ug/g AR) | 104.04 | ICP-MS | 211 | 0.171 | 0.186 | 0.19 | 0.182 | -1.17 | 0.2073 | 9 | 0.0213 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 183 | 0.1835 | 0.1945 | 0.192 | 0.19 | -0.81 | 0.2073 | 9 | 0.0213 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 194 | 0.1866 | 0.2009 | 0.1931 | 0.194 | -0.65 | 0.2073 | 9 | 0.0213 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 162 | 0.1886 | 0.2075 | 0.1986 | 0.198 | -0.43 | 0.2073 | 9 | 0.0213 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 197 | 0.211 | 0.1883 | 0.1962 | 0.199 | -0.41 | 0.2073 | 9 | 0.0213 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 142 | 0.2219 | 0.2159 | 0.2175 | 0.218 | 0.52 | 0.2073 | 9 | 0.0213 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 196 | 0.208 | 0.255 | 0.196 | 0.22 | 0.58 | 0.2073 | 9 | 0.0213 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 106 | 0.245 | 0.213 | 0.232 | 0.23 | 1.06 | 0.2073 | 9 | 0.0213 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 122 | 0.2349 | 0.2351 | 0.235 | 0.235 | 1.3 | 0.2073 | 9 | 0.0213 | 0 |
| Pb (ug/g AR) | 104.02 | furnace AA | 132 | 0.4 |  |  | 0.4 |  |  |  |  | 1 |

Cu (ug/g AR)

| Cu (ug/g AR) | 105.04 | ICP-MS | 196 | 12.4 | 9.72 | 11 | 11 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cu (ug/g AR) | 105.03 | ICP-OES | 132 | 14.4 |  |  | 14.4 | 1 |
| Cu (ug/g AR) | 105.04 | ICP-MS | 142 | 16.2952 | 14.9811 | 15.3548 | 15.5 | 0 |

|  |  |  | Lab |  |  |  | Lab |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Analyte | Code | Method | Num | Result1 Result2 Result3 |  |  | Value | Z score |  | Rob Mean | n | Rob StDev | Flag |
|  |  |  |  | Ni (ug/g AR) |  |  |  |  |  |  |  |  |  |
| Ni (ug/g AR) | 106.04 | ICP-MS | 196 | 1.15 | 0.971 | 1.12 | 1.08 | 0 |  |  |  |  |  |
| Ni (ug/g AR) | 106.03 | ICP-OES | 132 | 1.2 |  |  | 1.2 | 1 |  |  |  |  |  |
| Ni (ug/g AR) | 106.04 | ICP-MS | 141 | 1.43 | 1.41 | 1.27 | 1.37 | 0 |  |  |  |  |  |
|  |  |  |  | Cr (ug/g AR) |  |  |  |  |  |  |  |  |  |
| Cr (ug/g AR) | 107.04 | ICP-MS | 141 | <4.0000 | <4.0000 | <4.0000 | <4.0000 | 1 |  |  |  |  |  |
| Cr (ug/g AR) | 107.03 | ICP-OES | 132 | 0.24 |  |  | 0.24 | 1 |  |  |  |  |  |
| Cr (ug/g AR) | 107.04 | ICP-MS | 142 | 0.3696 | 0.3134 | 0.3374 | 0.34 | 0 |  |  |  |  |  |
| Cr (ug/g AR) | 107.04 | ICP-MS | 197 | 0.4076 | 0.3543 | 0.4253 | 0.396 | 0 |  |  |  |  |  |
| Cr (ug/g AR) | 107.04 | ICP-MS | 196 | 0.496 | 0.438 | 0.525 | 0.486 | 0 |  |  |  |  |  |
|  |  |  |  | Se (ug/g AR) |  |  |  |  |  |  |  |  |  |
| Se (ug/g AR) | 108.04 | ICP-MS | 197 | <0.0998 | <0.1018 | <0.1002 | <0.0998 | 1 |  |  |  |  |  |
| Se (ug/g AR) | 108.04 | ICP-MS | 141 | <1.0000 | <1.0000 | <1.0000 | <1.0000 | 1 |  |  |  |  |  |
| Se (ug/g AR) | 108.04 | ICP-MS | 142 | 0.0709 | 0.1085 | 0.0875 | 0.089 | 0 |  |  |  |  |  |
| Se (ug/g AR) | 108.04 | ICP-MS | 196 | 0.219 | 0.159 | 0.184 | 0.187 | 0 |  |  |  |  |  |
| Se (ug/g AR) | 108.03 | ICP-OES | 132 | 1.2 |  |  | 1.2 | 1 |  |  |  |  |  |
|  |  |  |  | Ag (ug/g AR) |  |  |  |  |  |  |  |  |  |
| Ag (ug/g AR) | 109.03 | ICP-OES | 132 | 0.104 |  |  | 0.104 | 1 |  |  |  |  |  |

Population of Lab Values

Ba (ug/g AR)

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

Value	Z score

Population of Lab Values Rob Mean	n	Rob StDev

Flag

Ba (ug/g AR)

110.03	ICP-OES

132

29.2	29.2	1

Zn (ug/g AR)

| Zn (ug/g AR) | 111.04 | ICP-MS | 141 | 54.2 | 48.1 | 51.4 | 51.2 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Zn (ug/g AR) | 111.03 | ICP-OES | 132 | 59.2 |  |  | 59.2 | 1 |
| Zn (ug/g AR) | 111.04 | ICP-MS | 142 | 64.0727 | 63.5184 | 63.2279 | 63.6 | 0 |
| Zn (ug/g AR) | 111.04 | ICP-MS | 196 | 72 | 59 | 62.3 | 64.4 | 0 |

Sb (ug/g AR)

Sb (ug/g AR)

112.04	ICP-MS

196

<0.0241  <0.0241	<0.0241	<0.0241	1

(-)-alpha-Bisabolol (%AR)

| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 142 | 0.0226 | 0.0178 | 0 | 0.0202 | -1 | 0.0551 | 7 | 0.0347 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 162 | 0.0293 | 0.0248 | 0.024 | 0.026 | -0.84 | 0.0551 | 7 | 0.0347 | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 196 | 0.0264 | 0.0311 | 0.0365 | 0.0313 | -0.68 | 0.0551 | 7 | 0.0347 | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 194 | 0.0475 | 0.0502 | 0.0489 | 0.0489 | -0.18 | 0.0551 | 7 | 0.0347 | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 108 | 0.0862 | 0.0844 | 0.0787 | 0.0831 | 0.81 | 0.0551 | 7 | 0.0347 | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 200 | 0.085 | 0.0907 | 0.0824 | 0.086 | 0.89 | 0.0551 | 7 | 0.0347 | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 212 | 0.0919 | 0.0895 | 0.0885 | 0.09 | 1.01 | 0.0551 | 7 | 0.0347 | 0 |

| Analyte | Code |  | Method | Lab Num | Result1 Result2 Result3 |  |  | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | beta-Caryophyllene (%AR) |  |  |  |  |  |  |  |  |  |
| beta-Caryophyllene (% |  | 202.02 | GC,MS | 162 | 0.0687 | 0.0605 | 0.063 | 0.0641 | -0.85 | 0.0736 | 7 | 0.0112 | 0 |
| beta-Caryophyllene (% |  | 202.02 | GC,MS | 194 | 0.0709 | 0.0661 | 0.0616 | 0.0662 | -0.66 | 0.0736 | 7 | 0.0112 | 0 |
| beta-Caryophyllene (% |  | 202.02 | GC,MS | 108 | 0.0694 | 0.0693 | 0.066 | 0.0682 | -0.48 | 0.0736 | 7 | 0.0112 | 0 |
| beta-Caryophyllene (% |  | 202.02 | GC,MS | 200 | 0.0685 | 0.0724 | 0.068 | 0.0696 | -0.35 | 0.0736 | 7 | 0.0112 | 0 |
| beta-Caryophyllene (% |  | 202.02 | GC,MS | 196 | 0.0666 | 0.0706 | 0.08 | 0.0724 | -0.11 | 0.0736 | 7 | 0.0112 | 0 |
| beta-Caryophyllene (% |  | 202.02 | GC,MS | 212 | 0.0812 | 0.0802 | 0.0911 | 0.0842 | 0.95 | 0.0736 | 7 | 0.0112 | 0 |
| beta-Caryophyllene (% |  | 202.02 | GC,MS | 142 | 0.0908 | 0.0813 | 0.1324 | 0.101 | 2.49 | 0.0736 | 7 | 0.0112 | 0 |

Population of Lab Values

(-)-Caryophyllene oxide (%AR)

| (-)-Caryophyllene oxide | 203.02 | GC,MS | 142 | 0.0062 | 0.0071 | 0.008 | 0.0071 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| (-)-Caryophyllene oxide | 203.02 | GC,MS | 196 | 0.0083 | 0.0102 | 0.0083 | 0.0089 | 0 |
| (-)-Caryophyllene oxide | 203.02 | GC,MS | 200 | 0.0148 | 0.0134 | 0.0123 | 0.0135 | 0 |
| (-)-Caryophyllene oxide | 203.02 | GC,MS | 194 | 0.0152 | 0.0156 | 0.0151 | 0.0153 | 0 |
| (-)-Caryophyllene oxide | 203.02 | GC,MS | 212 | 0.0163 | 0.0186 | 0.0212 | 0.0187 | 0 |

Geraniol (%AR)

| Geraniol (%AR) | 204.02 | GC,MS | 142 | <0.0001 | <0.0001 | <0.0001 | <0.0001 | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Geraniol (%AR) | 204.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 | <0.0004 | 1 |
| Geraniol (%AR) | 204.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 | 1 |
| Geraniol (%AR) | 204.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| Geraniol (%AR) | 204.02 | GC,MS | 200 | <0.0400 | <0.0400 | <0.0400 | <0.0400 | 1 |

| Analyte | Code | Method | Lab Num | Result1 Result2 Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

Population of Lab Values

alpha-Humulene (%AR)

| alpha-Humulene (%AR) | 205.02 | GC,MS | 162 | 0.0225 | 0.019 | 0.0201 | 0.0205 | -1.38 | 0.0264 | 6 | 0.0043 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 196 | 0.0223 | 0.0239 | 0.0277 | 0.0246 | -0.42 | 0.0264 | 6 | 0.0043 | 0 |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 194 | 0.0272 | 0.0252 | 0.0236 | 0.0253 | -0.25 | 0.0264 | 6 | 0.0043 | 0 |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 142 | 0.025 | 0.0219 | 0.036 | 0.0276 | 0.29 | 0.0264 | 6 | 0.0043 | 0 |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 200 | 0.0309 | 0.0291 | 0.0274 | 0.0291 | 0.64 | 0.0264 | 6 | 0.0043 | 0 |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 212 | 0.0295 | 0.0298 | 0.0343 | 0.0312 | 1.12 | 0.0264 | 6 | 0.0043 | 0 |

d-Limonene (%AR)

| d-Limonene (%AR) | 206.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| d-Limonene (%AR) | 206.02 | GC,MS | 142 | 0.0007 | 0.001 | 0.0006 | 0.0008 | 0 |
| d-Limonene (%AR) | 206.02 | GC,MS | 196 | 0.0015 | 0.0015 | 0.0015 | 0.0015 | 0 |
| d-Limonene (%AR) | 206.02 | GC,MS | 194 | 0.0023 | 0.0022 | 0.0022 | 0.0022 | 0 |
| d-Limonene (%AR) | 206.02 | GC,MS | 200 | 0.0278 | 0.0205 | 0.0249 | 0.0244 | 0 |

Linalool (%AR)

| Linalool (%AR) | 207.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Linalool (%AR) | 207.02 | GC,MS | 142 | 0.001 | 0.001 | 0.0014 | 0.0011 | 0 |
| Linalool (%AR) | 207.02 | GC,MS | 196 | 0.0017 | 0.0019 | <0.0013 | 0.0016 | 0 |
| Linalool (%AR) | 207.02 | GC,MS | 212 | 0.0034 | 0.0043 | 0.0032 | 0.0036 | 0 |
| Linalool (%AR) | 207.02 | GC,MS | 194 | 0.004 | 0.004 | 0.004 | 0.004 | 0 |

|  |  |  | Lab |  |  |  | Lab |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Analyte | Code | Method | Num | Result1 Result2 Result3 |  |  | Value | Z score |  | Rob Mean | n | Rob StDev | Flag |
| Linalool (%AR) | 207.02 | GC,MS | 200 | 0.0056 | 0.0045 | 0.0052 | 0.0051 | 0 |  |  |  |  |  |
|  |  |  |  | beta-Myrcene (%AR) |  |  |  |  |  |  |  |  |  |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |  |  |  |  |  |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 142 | 0.0019 | 0.0024 | 0.0016 | 0.002 | 0 |  |  |  |  |  |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 200 | 0.0025 | 0.0028 | 0.0026 | 0.0026 | 0 |  |  |  |  |  |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 196 | 0.0042 | 0.0043 | 0.0043 | 0.0043 | 0 |  |  |  |  |  |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 212 | 0.0032 | 0.0037 | 0.006 | 0.0043 | 0 |  |  |  |  |  |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 194 | 0.0056 | 0.0052 | 0.0051 | 0.0053 | 0 |  |  |  |  |  |
|  |  |  |  | Ocimene (%AR) |  |  |  |  |  |  |  |  |  |
| Ocimene (%AR) | 210.04 | Other | 142 | <0.0001 | <0.0001 | <0.0001 | <0.0001 | 1 |  |  |  |  |  |
| Ocimene (%AR) | 210.02 | GC,MS | 200 | <0.0009 | <0.0009 | <0.0009 | <0.0009 | 1 |  |  |  |  |  |
| Ocimene (%AR) | 210.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 | 1 |  |  |  |  |  |
| Ocimene (%AR) | 210.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |  |  |  |  |  |
| Ocimene (%AR) | 210.02 | GC,MS | 194 | 0.0016 | 0.0016 | 0.0016 | 0.0016 | 0 |  |  |  |  |  |
|  |  |  |  | alpha-Pinene (%AR) |  |  |  |  |  |  |  |  |  |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 | 1 |  |  |  |  |  |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |  |  |  |  |  |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 200 | 0.0009 | 0.0011 | 0.001 | 0.001 | 0 |  |  |  |  |  |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 142 | 0.0006 | 0.0008 | 0.0004 | 0.0006 | 0 |  |  |  |  |  |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 212 | 0.0009 | 0.0015 | 0.001 | 0.0011 |  |  |  |  | 0 |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 194 | 0.0011 | 0.0011 | 0.001 | 0.0011 |  |  |  |  | 0 |
|  |  |  |  | (-)-beta-Pinene (%AR) |  |  |  |  |  |  |  |  |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 |  |  |  |  | 1 |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 200 | <0.0013 | <0.0013 | <0.0013 | <0.0013 |  |  |  |  | 1 |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 142 | 0.0002 | 0.0003 | 0.0002 | 0.0002 |  |  |  |  | 0 |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 194 | 0.0011 | 0.0011 | 0.0011 | 0.0011 |  |  |  |  | 0 |
|  |  |  |  | alpha-Terpinene (%AR) |  |  |  |  |  |  |  |  |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 200 | <0.0010 | <0.0010 | <0.0010 | <0.0010 |  |  |  |  | 1 |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 |  |  |  |  | 1 |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 142 | 0.0001 | <0.0001 | 0.0001 | 0.0001 |  |  |  |  | 0 |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 194 | 0.0009 | 0.0009 | 0.0009 | 0.0009 |  |  |  |  | 0 |
|  |  |  |  | gamma-Terpinene (%AR) |  |  |  |  |  |  |  |  |
| gamma-Terpinene (%A | 214.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 | <0.0004 |  |  |  |  | 1 |
| gamma-Terpinene (%A | 214.02 | GC,MS | 200 | <0.0004 | <0.0004 | <0.0004 | <0.0004 |  |  |  |  | 1 |
| gamma-Terpinene (%A | 214.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 |  |  |  |  | 1 |
| gamma-Terpinene (%A | 214.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |

Population of Lab Values

|  |  |  | Lab |  |  |  | Lab |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Analyte | Code | Method | Num | Result1 Result2 Result3 |  |  | Value | Z score |  | Rob Mean | n | Rob StDev | Flag |
| gamma-Terpinene (%A | 214.02 | GC,MS | 142 | 0.0001 | 0.0001 | 0.0001 | 0.0001 |  |  |  |  |  | 0 |
|  |  |  |  | alpha-Terpinolene (%AR) |  |  |  |  |  |  |  |  |  |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 200 | <0.0008 | <0.0008 | <0.0008 | <0.0008 |  |  |  |  |  | 1 |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 |  |  |  |  |  | 1 |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 142 | 0.0001 | 0.0001 | 0.0001 | 0.0001 |  |  |  |  |  | 0 |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 194 | 0.0009 | 0.0009 | 0.0009 | 0.0009 |  |  |  |  |  | 0 |
|  |  |  |  | Moisture (%AR) |  |  |  |  |  |  |  |  |  |
| Moisture (%AR) | 500.99 | Other | 142 | 0.0444 | 0.0409 | 0.042 | 0.0424 | -3.37 | 6.5236 |  | 25 | 1.9224 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 153 | 3.04 | 2.63 | 2.98 | 2.88 | -1.89 | 6.5236 |  | 25 | 1.9224 | 0 |
| Moisture (%AR) | 500.99 | Other | 189 | 4.351 | 4.351 | 4.351 | 4.35 | -1.13 | 6.5236 |  | 25 | 1.9224 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 109 | 4.99 | 4.5 | 5.61 | 5.03 | -0.78 | 6.5236 |  | 25 | 1.9224 | 0 |
| Moisture (%AR) | 500.72 | NIR | 187 | 5.22 | 5.18 | 5.13 | 5.18 | -0.7 | 6.5236 |  | 25 | 1.9224 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 116 | 5.22 | 5.35 | 5.19 | 5.25 | -0.66 | 6.5236 |  | 25 | 1.9224 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 212 | 5.29 | 5.17 | 5.33 | 5.26 | -0.66 | 6.5236 |  | 25 | 1.9224 | 0 |
| Moisture (%AR) | 500.99 | Other | 103 | 5.22 | 5.37 | 5.54 | 5.38 | -0.6 | 6.5236 |  | 25 | 1.9224 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 204 | 5.65 | 5.65 | 5.65 | 5.65 | -0.45 | 6.5236 |  | 25 | 1.9224 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 106 | 5.88 | 5.46 | 5.69 | 5.68 | -0.44 | 6.5236 |  | 25 | 1.9224 | 0 |
| Moisture (%AR) | 500.99 | Other | 132 | 6.89 | 6.67 | 4.52 | 6.03 | -0.26 | 6.5236 |  | 25 | 1.9224 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 196 | 6.11 | 6.47 | 5.66 | 6.08 | -0.23 | 6.5236 |  | 25 | 1.9224 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Moisture (%AR) | 500.72 | NIR | 138 | 6.35 | 6.7 | 6.5 | 6.52 | 0 | 6.5236 | 25 | 1.9224 | 0 |
| Moisture (%AR) | 500.99 | Other | 215 | 7.0185 | 6.6819 | 6.0198 | 6.57 | 0.03 | 6.5236 | 25 | 1.9224 | 0 |
| Moisture (%AR) | 500.99 | Other | 102 | 6.3 | 6.8 | 7.3 | 6.8 | 0.14 | 6.5236 | 25 | 1.9224 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 182 | 6.485 | 7.075 | 7.28 | 6.95 | 0.22 | 6.5236 | 25 | 1.9224 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 211 | 6.84 | 7.23 | 7.01 | 7.03 | 0.26 | 6.5236 | 25 | 1.9224 | 0 |
| Moisture (%AR) | 500.99 | Other | 120 | 7.2233 | 7.1596 | 7.0398 | 7.14 | 0.32 | 6.5236 | 25 | 1.9224 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 210 | 7.91 | 7.27 | 7.09 | 7.42 | 0.47 | 6.5236 | 25 | 1.9224 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 199 | 7.9819 | 7.7946 | 8.7431 | 8.17 | 0.86 | 6.5236 | 25 | 1.9224 | 0 |
| Moisture (%AR) | 500.99 | Other | 105 | 8.69 | 8.62 | 8.63 | 8.65 | 1.1 | 6.5236 | 25 | 1.9224 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 208 | 9.25 | 8.65 | 8.85 | 8.92 | 1.24 | 6.5236 | 25 | 1.9224 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 192 | 8.9921 | 8.9609 | 8.881 | 8.94 | 1.26 | 6.5236 | 25 | 1.9224 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 172 | 9.778 | 9.272 | 9.63 | 9.56 | 1.58 | 6.5236 | 25 | 1.9224 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 209 | 9.9081 | 10.6212 | 10.7182 | 10.4 | 2.02 | 6.5236 | 25 | 1.9224 | 0 |
|  |  |  |  | Δ9-THC (%DW) |  |  |  |  |  |  |  |  |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 154 | 0.1603 | 0.1561 | 0.1342 | 0.15 | -2.2 | 0.2128 | 23 | 0.0284 | 0 |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, | 103 | 0.196 | 0.175 | 0.182 | 0.184 | -1 | 0.2128 | 23 | 0.0284 | 0 |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, | 218 | 0.184 | 0.184 | 0.191 | 0.186 | -0.93 | 0.2128 | 23 | 0.0284 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 118 | 0.192 | 0.185 | 0.181 | 0.186 | -0.94 | 0.2128 | 23 | 0.0284 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 109 | 0.208 | 0.166 | 0.19 | 0.188 | -0.87 | 0.2128 | 23 | 0.0284 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 142 | 0.1886 | 0.1914 | 0.188 | 0.189 | -0.83 | 0.2128 | 23 | 0.0284 | 0 |
| Δ9-THC (%DW) | 501.30 | Other LC-MS | 105 | 0.203 | 0.187 | 0.197 | 0.196 | -0.6 | 0.2128 | 23 | 0.0284 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%DW) | 501.60 | NIR | 187 | 0.19 | 0.2 | 0.2 | 0.197 | -0.57 | 0.2128 | 23 | 0.0284 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 208 | 0.2039 | 0.1924 | 0.1994 | 0.199 | -0.5 | 0.2128 | 23 | 0.0284 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 196 | 0.2043 | 0.2111 | 0.1899 | 0.202 | -0.39 | 0.2128 | 23 | 0.0284 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 204 | 0.2042 | 0.2053 | 0.2039 | 0.204 | -0.29 | 0.2128 | 23 | 0.0284 | 0 |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, | 199 | 0.2087 | 0.2034 | 0.212 | 0.208 | -0.17 | 0.2128 | 23 | 0.0284 | 0 |
| Δ9-THC (%DW) | 501.99 | Other | 215 | 0.2181 | 0.2167 | 0.2127 | 0.216 | 0.11 | 0.2128 | 23 | 0.0284 | 0 |
| Δ9-THC (%DW) | 501.30 | Other LC-MS | 120 | 0.2187 | 0.2221 | 0.2149 | 0.219 | 0.2 | 0.2128 | 23 | 0.0284 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 211 | 0.215 | 0.222 | 0.224 | 0.22 | 0.26 | 0.2128 | 23 | 0.0284 | 0 |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, | 210 | 0.222 | 0.225 | 0.22 | 0.222 | 0.34 | 0.2128 | 23 | 0.0284 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 189 | 0.2316 | 0.2274 | 0.2278 | 0.229 | 0.57 | 0.2128 | 23 | 0.0284 | 0 |
| Δ9-THC (%DW) | 501.30 | Other LC-MS | 102 | 0.232 | 0.23 | 0.245 | 0.236 | 0.8 | 0.2128 | 23 | 0.0284 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 116 | 0.2401 | 0.2463 | 0.2343 | 0.24 | 0.96 | 0.2128 | 23 | 0.0284 | 0 |
| Δ9-THC (%DW) | 501.01 | AOAC 2018.10 | 209 | 0.2436 | 0.2317 | 0.2519 | 0.242 | 1.04 | 0.2128 | 23 | 0.0284 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 172 | 0.2499 | 0.2576 | 0.243 | 0.25 | 1.31 | 0.2128 | 23 | 0.0284 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 209 | 0.2744 | 0.2499 | 0.2579 | 0.261 | 1.69 | 0.2128 | 23 | 0.0284 | 0 |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, | 209 | 0.2948 | 0.2648 | 0.2483 | 0.269 | 1.99 | 0.2128 | 23 | 0.0284 | 0 |
|  |  |  |  | Δ9-THCA (%DW) |  |  |  |  |  |  |  |  |
| Δ9-THCA (%DW) | 502.99 | Other | 215 | 0.0589 | 0.0587 | 0.0573 | 0.0583 | -1.22 | 0.0794 | 23 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 118 | 0.063 | 0.06 | 0.06 | 0.061 | -1.07 | 0.0794 | 23 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.30 | Other LC-MS | 105 | 0.064 | 0.0576 | 0.0655 | 0.0624 | -0.99 | 0.0794 | 23 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 196 | 0.0638 | 0.0694 | 0.062 | 0.0651 | -0.83 | 0.0794 | 23 | 0.0173 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 142 | 0.0665 | 0.0664 | 0.0643 | 0.0657 | -0.79 | 0.0794 | 23 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 154 | 0.0712 | 0.069 | 0.061 | 0.0671 | -0.72 | 0.0794 | 23 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 211 | 0.062 | 0.067 | 0.075 | 0.068 | -0.66 | 0.0794 | 23 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 208 | 0.0676 | 0.0671 | 0.0769 | 0.0705 | -0.51 | 0.0794 | 23 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, | 218 | 0.076 | 0.069 | 0.069 | 0.0713 | -0.47 | 0.0794 | 23 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.30 | Other LC-MS | 102 | 0.072 | 0.071 | 0.073 | 0.072 | -0.43 | 0.0794 | 23 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 109 | 0.082 | 0.063 | 0.073 | 0.0727 | -0.39 | 0.0794 | 23 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.03 | AOAC 2018.11, | 210 | 0.077 | 0.073 | 0.075 | 0.075 | -0.26 | 0.0794 | 23 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.30 | Other LC-MS | 120 | 0.0793 | 0.0795 | 0.0818 | 0.0802 | 0.05 | 0.0794 | 23 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, | 199 | 0.081 | 0.0808 | 0.0805 | 0.0808 | 0.08 | 0.0794 | 23 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 116 | 0.0797 | 0.0865 | 0.0827 | 0.083 | 0.21 | 0.0794 | 23 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 189 | 0.084 | 0.0843 | 0.0837 | 0.084 | 0.27 | 0.0794 | 23 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, | 103 | 0.074 | 0.085 | 0.095 | 0.0847 | 0.3 | 0.0794 | 23 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 204 | 0.0882 | 0.0899 | 0.0867 | 0.0883 | 0.51 | 0.0794 | 23 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 172 | 0.094 | 0.0984 | 0.094 | 0.0955 | 0.93 | 0.0794 | 23 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.01 | AOAC 2018.10 | 209 | 0.1071 | 0.1001 | 0.1094 | 0.106 | 1.51 | 0.0794 | 23 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 209 | 0.1196 | 0.1112 | 0.1101 | 0.114 | 1.98 | 0.0794 | 23 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, | 209 | 0.1285 | 0.1178 | 0.1061 | 0.117 | 2.2 | 0.0794 | 23 | 0.0173 | 0 |
| Δ9-THCA (%DW) | 502.60 | NIR | 187 | 0.15 | 0.13 | 0.13 | 0.137 | 3.32 | 0.0794 | 23 | 0.0173 | 0 |
|  |  |  |  | CBD (%DW) |  |  |  |  |  |  |  |  |
| CBD (%DW) | 503.10 | Other LC-UV | 154 | 3.062 | 2.8749 | 2.5553 | 2.83 | -1.89 | 3.5164 | 22 | 0.3633 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%DW) | 503.10 | Other LC-UV | 209 | 3.2223 | 2.9198 | 2.9443 | 3.03 | -1.34 | 3.5164 | 22 | 0.3633 | 0 |
| CBD (%DW) | 503.60 | NIR | 187 | 2.96 | 3.2 | 3.16 | 3.11 | -1.13 | 3.5164 | 22 | 0.3633 | 0 |
| CBD (%DW) | 503.01 | AOAC 2018.10 | 209 | 3.0473 | 3.0738 | 3.1991 | 3.11 | -1.13 | 3.5164 | 22 | 0.3633 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 118 | 3.246 | 3.123 | 3.048 | 3.14 | -1.04 | 3.5164 | 22 | 0.3633 | 0 |
| CBD (%DW) | 503.30 | Other LC-MS | 120 | 3.3248 | 3.279 | 2.9389 | 3.18 | -0.92 | 3.5164 | 22 | 0.3633 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 189 | 3.325 | 3.3171 | 3.2997 | 3.31 | -0.56 | 3.5164 | 22 | 0.3633 | 0 |
| CBD (%DW) | 503.02 | AOAC 2018.11, | 209 | 3.7053 | 3.1586 | 3.2206 | 3.36 | -0.43 | 3.5164 | 22 | 0.3633 | 0 |
| CBD (%DW) | 503.30 | Other LC-MS | 105 | 3.81 | 3.41 | 3.26 | 3.49 | -0.06 | 3.5164 | 22 | 0.3633 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 142 | 3.4592 | 3.6344 | 3.5357 | 3.54 | 0.07 | 3.5164 | 22 | 0.3633 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 109 | 3.897 | 3.162 | 3.58 | 3.55 | 0.08 | 3.5164 | 22 | 0.3633 | 0 |
| CBD (%DW) | 503.02 | AOAC 2018.11, | 218 | 3.617 | 3.553 | 3.553 | 3.57 | 0.16 | 3.5164 | 22 | 0.3633 | 0 |
| CBD (%DW) | 503.99 | Other | 215 | 3.3856 | 3.9392 | 3.5606 | 3.63 | 0.31 | 3.5164 | 22 | 0.3633 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 204 | 3.6248 | 3.656 | 3.6483 | 3.64 | 0.35 | 3.5164 | 22 | 0.3633 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 196 | 3.6117 | 3.8429 | 3.6294 | 3.69 | 0.49 | 3.5164 | 22 | 0.3633 | 0 |
| CBD (%DW) | 503.02 | AOAC 2018.11, | 103 | 3.761 | 3.737 | 3.692 | 3.73 | 0.59 | 3.5164 | 22 | 0.3633 | 0 |
| CBD (%DW) | 503.02 | AOAC 2018.11, | 210 | 3.807 | 3.851 | 3.718 | 3.79 | 0.76 | 3.5164 | 22 | 0.3633 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 208 | 3.9335 | 3.7368 | 3.822 | 3.83 | 0.87 | 3.5164 | 22 | 0.3633 | 0 |
| CBD (%DW) | 503.02 | AOAC 2018.11, | 199 | 3.8536 | 3.7731 | 3.8956 | 3.84 | 0.89 | 3.5164 | 22 | 0.3633 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 211 | 3.844 | 3.887 | 3.825 | 3.85 | 0.92 | 3.5164 | 22 | 0.3633 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 116 | 3.9496 | 3.9903 | 3.8232 | 3.92 | 1.11 | 3.5164 | 22 | 0.3633 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 172 | 4.4328 | 4.5964 | 4.285 | 4.44 | 2.54 | 3.5164 | 22 | 0.3633 | 0 |

Population of Lab Values

|  |  |  | Lab |  |  |  | Lab |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Analyte | Code | Method | Num | Result1 Result2 Result3 |  |  | Value | Z score |  | Rob Mean | n | Rob StDev | Flag |
|  |  |  |  | CBDA (%DW) |  |  |  |  |  |  |  |  |  |
| CBDA (%DW) | 504.10 | Other LC-UV | 154 | 4.5735 | 4.299 | 3.7693 | 4.21 | -1.56 | 5.4051 |  | 21 | 0.7655 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 118 | 4.574 | 4.436 | 4.412 | 4.47 | -1.22 | 5.4051 |  | 21 | 0.7655 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 189 | 4.6235 | 4.6094 | 4.5876 | 4.61 | -1.04 | 5.4051 |  | 21 | 0.7655 | 0 |
| CBDA (%DW) | 504.60 | NIR | 187 | 4.92 | 4.64 | 4.61 | 4.72 | -0.89 | 5.4051 |  | 21 | 0.7655 | 0 |
| CBDA (%DW) | 504.30 | Other LC-MS | 105 | 4.9 | 4.5 | 4.93 | 4.78 | -0.82 | 5.4051 |  | 21 | 0.7655 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 142 | 4.9295 | 5.0099 | 4.8071 | 4.92 | -0.64 | 5.4051 |  | 21 | 0.7655 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 116 | 5.0646 | 4.9706 | 4.8213 | 4.95 | -0.59 | 5.4051 |  | 21 | 0.7655 | 0 |
| CBDA (%DW) | 504.02 | AOAC 2018.11, | 218 | 5.071 | 4.911 | 4.885 | 4.96 | -0.59 | 5.4051 |  | 21 | 0.7655 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 211 | 5.364 | 5.31 | 5.285 | 5.32 | -0.11 | 5.4051 |  | 21 | 0.7655 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 196 | 5.2938 | 5.6633 | 5.0605 | 5.34 | -0.09 | 5.4051 |  | 21 | 0.7655 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 204 | 5.3748 | 5.4839 | 5.394 | 5.42 | 0.02 | 5.4051 |  | 21 | 0.7655 | 0 |
| CBDA (%DW) | 504.02 | AOAC 2018.11, | 103 | 5.662 | 5.36 | 5.381 | 5.47 | 0.08 | 5.4051 |  | 21 | 0.7655 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 109 | 6.072 | 4.944 | 5.617 | 5.54 | 0.18 | 5.4051 |  | 21 | 0.7655 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 208 | 5.8668 | 5.4488 | 5.642 | 5.65 | 0.32 | 5.4051 |  | 21 | 0.7655 | 0 |
| CBDA (%DW) | 504.99 | Other | 215 | 5.8469 | 5.8274 | 5.6427 | 5.77 | 0.48 | 5.4051 |  | 21 | 0.7655 | 0 |
| CBDA (%DW) | 504.02 | AOAC 2018.11, | 210 | 5.826 | 5.883 | 5.649 | 5.79 | 0.5 | 5.4051 |  | 21 | 0.7655 | 0 |
| CBDA (%DW) | 504.02 | AOAC 2018.11, | 199 | 5.7864 | 5.8126 | 5.9837 | 5.86 | 0.6 | 5.4051 |  | 21 | 0.7655 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 172 | 6.1931 | 6.3888 | 6.116 | 6.23 | 1.08 | 5.4051 |  | 21 | 0.7655 | 0 |
| CBDA (%DW) | 504.01 | AOAC 2018.10 | 209 | 6.2212 | 6.2382 | 6.6503 | 6.37 | 1.26 | 5.4051 |  | 21 | 0.7655 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 209 | 6.967 | 6.345 | 6.2796 | 6.53 | 1.47 | 5.4051 |  | 21 | 0.7655 | 0 |
| CBDA (%DW) | 504.02 | AOAC 2018.11, | 209 | 8.1363 | 7.0034 | 6.9305 | 7.36 | 2.55 | 5.4051 |  | 21 | 0.7655 | 0 |

Population of Lab Values

|  |  |  | Lab |  |  |  | Lab |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Analyte | Code | Method | Num | Result1 Result2 Result3 |  |  | Value | Z score |  | Rob Mean | n | Rob StDev | Flag |
|  |  |  |  | CBN (%DW) |  |  |  |  |  |  |  |  |  |
| CBN (%DW) | 505.10 | Other LC-UV | 154 | <0.0128 | <0.0128 | <0.0128 | <0.0128 |  |  |  |  |  | 1 |
| CBN (%DW) | 505.02 | AOAC 2018.11, | 210 | <0.0140 | <0.0140 | <0.0140 | <0.0140 |  |  |  |  |  | 1 |
| CBN (%DW) | 505.02 | AOAC 2018.11, | 199 | <0.0250 | <0.0250 | <0.0250 | <0.0250 |  |  |  |  |  | 1 |
| CBN (%DW) | 505.30 | GC-FID | 116 | <0.0400 | <0.0400 | <0.0400 | <0.0400 |  |  |  |  |  | 1 |
| CBN (%DW) | 505.10 | Other LC-UV | 208 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| CBN (%DW) | 505.01 | AOAC 2018.10 | 209 | 0.0112 | 0.0111 | 0.0129 | 0.0117 | -1.45 | 0.0163 |  | 13 | 0.0031 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 209 | 0.0145 | 0.0124 | 0.0128 | 0.0132 | -0.97 | 0.0163 |  | 13 | 0.0031 | 0 |
| CBN (%DW) | 505.02 | AOAC 2018.11, | 209 | 0.0156 | 0.0132 | 0.0123 | 0.0137 | -0.82 | 0.0163 |  | 13 | 0.0031 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 196 | 0.0136 | 0.0136 | 0.0137 | 0.0136 | -0.84 | 0.0163 |  | 13 | 0.0031 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 109 | 0.016 | 0.014 | 0.016 | 0.0153 | -0.3 | 0.0163 |  | 13 | 0.0031 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 142 | 0.0123 | 0.016 | 0.0151 | 0.0145 | -0.58 | 0.0163 |  | 13 | 0.0031 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 192 | 0.0194 | 0.0154 | 0.0174 | 0.0174 | 0.36 | 0.0163 |  | 13 | 0.0031 | 0 |
| CBN (%DW) | 505.20 | Other LC-MS | 105 | 0.0171 | 0.0183 | 0.0163 | 0.0172 | 0.31 | 0.0163 |  | 13 | 0.0031 | 0 |
| CBN (%DW) | 505.02 | AOAC 2018.11, | 218 | 0.019 | 0.018 | 0.018 | 0.0183 | 0.66 | 0.0163 |  | 13 | 0.0031 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 204 | 0.0182 | 0.0181 | 0.0183 | 0.0182 | 0.61 | 0.0163 |  | 13 | 0.0031 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 189 | 0.0186 | 0.0188 | 0.0187 | 0.0187 | 0.77 | 0.0163 |  | 13 | 0.0031 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 211 | 0.018 | 0.018 | 0.02 | 0.0187 | 0.76 | 0.0163 |  | 13 | 0.0031 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 116 | 0.0436 | 0.0484 | 0.044 | 0.0453 | 9.28 | 0.0163 |  | 13 | 0.0031 | 0 |

Population of Lab Values

|  |  |  | Lab |  |  |  | Lab |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Analyte | Code | Method | Num | Result1 Result2 Result3 |  |  | Value | Z score |  | Rob Mean | n | Rob StDev | Flag |
|  |  |  |  | Total Δ9-THC (%DW) |  |  |  |  |  |  |  |  |  |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 192 | 0.2059 | 0.2029 | 0.2146 | 0.208 | -1.57 | 0.2758 |  | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 154 | 0.2227 | 0.2166 | 0.1877 | 0.209 | -1.54 | 0.2758 |  | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 182 | 0.2328 | 0.226 | 0.2214 | 0.227 | -1.13 | 0.2758 |  | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 118 | 0.247 | 0.238 | 0.234 | 0.24 | -0.84 | 0.2758 |  | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 142 | 0.2468 | 0.2497 | 0.2444 | 0.247 | -0.67 | 0.2758 |  | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, | 218 | 0.25 | 0.244 | 0.251 | 0.248 | -0.64 | 0.2758 |  | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.30 | Other LC-MS | 105 | 0.259 | 0.238 | 0.254 | 0.25 | -0.59 | 0.2758 |  | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 109 | 0.28 | 0.222 | 0.254 | 0.252 | -0.55 | 0.2758 |  | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 138 | 0.2572 | 0.262 | 0.2581 | 0.259 | -0.39 | 0.2758 |  | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 196 | 0.2602 | 0.272 | 0.2442 | 0.259 | -0.39 | 0.2758 |  | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, | 103 | 0.261 | 0.25 | 0.266 | 0.259 | -0.39 | 0.2758 |  | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 208 | 0.2632 | 0.2513 | 0.2669 | 0.26 | -0.35 | 0.2758 |  | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 153 | 0.2881 | 0.2484 | 0.2474 | 0.261 | -0.34 | 0.2758 |  | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 116 | 0.2758 | 0.2721 | 0.2789 | 0.276 | 0 | 0.2758 |  | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, | 199 | 0.2797 | 0.2742 | 0.2826 | 0.279 | 0.07 | 0.2758 |  | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 211 | 0.27 | 0.28 | 0.29 | 0.28 | 0.1 | 0.2758 |  | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 204 | 0.2816 | 0.2841 | 0.28 | 0.282 | 0.14 | 0.2758 |  | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.30 | Other LC-MS | 120 | 0.2883 | 0.2918 | 0.2866 | 0.289 | 0.3 | 0.2758 |  | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.30 | Other LC-MS | 102 | 0.295 | 0.292 | 0.309 | 0.299 | 0.53 | 0.2758 |  | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 189 | 0.3053 | 0.3013 | 0.3012 | 0.303 | 0.62 | 0.2758 |  | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.60 | NIR | 187 | 0.31 | 0.31 | 0.31 | 0.31 | 0.79 | 0.2758 |  | 26 | 0.0433 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 116 | 0.31 | 0.3222 | 0.3068 | 0.313 | 0.86 | 0.2758 | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.01 | AOAC 2018.10 | 209 | 0.3375 | 0.3195 | 0.3479 | 0.335 | 1.37 | 0.2758 | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 172 | 0.3323 | 0.3439 | 0.3598 | 0.345 | 1.61 | 0.2758 | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 209 | 0.3793 | 0.3474 | 0.3545 | 0.36 | 1.95 | 0.2758 | 26 | 0.0433 | 0 |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, | 209 | 0.4075 | 0.3681 | 0.3413 | 0.372 | 2.23 | 0.2758 | 26 | 0.0433 | 0 |
|  |  |  |  | Total CBD (%DW) |  |  |  |  |  |  |  |  |
| Total CBD (%DW) | 507.10 | Other LC-UV | 154 | 7.0729 | 6.6451 | 5.861 | 6.53 | -1.71 | 8.3644 | 22 | 1.0746 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 118 | 7.258 | 7.014 | 6.918 | 7.06 | -1.21 | 8.3644 | 22 | 1.0746 | 0 |
| Total CBD (%DW) | 507.40 | GC-FID | 116 | 6.4534 | 7.9432 | 6.8065 | 7.07 | -1.21 | 8.3644 | 22 | 1.0746 | 0 |
| Total CBD (%DW) | 507.60 | NIR | 187 | 7.28 | 7.27 | 7.2 | 7.25 | -1.04 | 8.3644 | 22 | 1.0746 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 189 | 7.3798 | 7.3596 | 7.323 | 7.35 | -0.94 | 8.3644 | 22 | 1.0746 | 0 |
| Total CBD (%DW) | 507.30 | Other LC-MS | 105 | 8.11 | 7.36 | 7.58 | 7.68 | -0.63 | 8.3644 | 22 | 1.0746 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 142 | 7.7823 | 8.0281 | 7.7515 | 7.85 | -0.47 | 8.3644 | 22 | 1.0746 | 0 |
| Total CBD (%DW) | 507.02 | AOAC 2018.11, | 218 | 8.064 | 7.86 | 7.837 | 7.92 | -0.41 | 8.3644 | 22 | 1.0746 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 116 | 8.3913 | 8.3495 | 8.0515 | 8.26 | -0.09 | 8.3644 | 22 | 1.0746 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 196 | 8.2544 | 8.8096 | 8.0675 | 8.38 | 0.01 | 8.3644 | 22 | 1.0746 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 204 | 8.3385 | 8.4654 | 8.3836 | 8.4 | 0.03 | 8.3644 | 22 | 1.0746 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 109 | 9.222 | 7.498 | 8.506 | 8.41 | 0.04 | 8.3644 | 22 | 1.0746 | 0 |
| Total CBD (%DW) | 507.02 | AOAC 2018.11, | 103 | 8.727 | 8.437 | 8.411 | 8.52 | 0.15 | 8.3644 | 22 | 1.0746 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 211 | 8.549 | 8.544 | 8.459 | 8.52 | 0.14 | 8.3644 | 22 | 1.0746 | 0 |
| Total CBD (%DW) | 507.01 | AOAC 2018.10 | 209 | 8.5034 | 8.5446 | 9.0314 | 8.69 | 0.31 | 8.3644 | 22 | 1.0746 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%DW) | 507.10 | Other LC-UV | 209 | 9.3323 | 8.4843 | 8.4516 | 8.76 | 0.36 | 8.3644 | 22 | 1.0746 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 208 | 9.0787 | 8.5154 | 8.77 | 8.79 | 0.39 | 8.3644 | 22 | 1.0746 | 0 |
| Total CBD (%DW) | 507.02 | AOAC 2018.11, | 199 | 8.9283 | 8.8708 | 9.1433 | 8.98 | 0.57 | 8.3644 | 22 | 1.0746 | 0 |
| Total CBD (%DW) | 507.40 | GC-FID | 138 | 9.7107 | 9.6967 | 9.3896 | 9.6 | 1.15 | 8.3644 | 22 | 1.0746 | 0 |
| Total CBD (%DW) | 507.02 | AOAC 2018.11, | 209 | 10.8408 | 9.3005 | 9.2986 | 9.81 | 1.35 | 8.3644 | 22 | 1.0746 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 172 | 9.8641 | 10.1994 | 10.6687 | 10.2 | 1.75 | 8.3644 | 22 | 1.0746 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 192 | 10.787 | 9.9114 | 10.0857 | 10.3 | 1.77 | 8.3644 | 22 | 1.0746 | 0 |
|  |  |  |  | CBDV (%DW) |  |  |  |  |  |  |  |  |
| CBDV (%DW) | 508.99 | Other | 154 | <0.0118 | <0.0118 | <0.0118 | <0.0118 |  |  |  |  | 1 |
| CBDV (%DW) | 508.99 | Other | 189 | 0.0228 | 0.0224 | 0.0218 | 0.0223 | -0.7 | 0.0502 | 7 | 0.0397 | 0 |
| CBDV (%DW) | 508.99 | Other | 196 | 0.024 | 0.0279 | 0.0249 | 0.0256 | -0.62 | 0.0502 | 7 | 0.0397 | 0 |
| CBDV (%DW) | 508.02 | AOAC 2018.11, | 210 | 0.03 | 0.031 | 0.032 | 0.031 | -0.48 | 0.0502 | 7 | 0.0397 | 0 |
| CBDV (%DW) | 508.99 | Other | 211 | 0.032 | 0.03 | 0.039 | 0.0337 | -0.42 | 0.0502 | 7 | 0.0397 | 0 |
| CBDV (%DW) | 508.99 | Other | 142 | 0.0363 | 0.0391 | 0.0363 | 0.0372 | -0.33 | 0.0502 | 7 | 0.0397 | 0 |
| CBDV (%DW) | 508.99 | Other | 208 | 0.0576 | <0.0500 | <0.0500 | 0.0525 |  |  |  |  | 1 |
| CBDV (%DW) | 508.99 | Other | 192 | 0.0898 | 0.097 | 0.1042 | 0.097 | 1.18 | 0.0502 | 7 | 0.0397 | 0 |
| CBDV (%DW) | 508.99 | Other | 109 | 0.115 | 0.092 | 0.107 | 0.105 | 1.37 | 0.0502 | 7 | 0.0397 | 0 |
|  |  |  |  | CBDVA (%DW) |  |  |  |  |  |  |  |  |
| CBDVA (%DW) | 509.99 | Other | 154 | 0.0259 | 0.0264 | 0.022 | 0.0248 |  |  |  |  | 0 |
| CBDVA (%DW) | 509.99 | Other | 196 | 0.0492 | 0.0529 | 0.0486 | 0.0502 |  |  |  |  | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDVA (%DW) | 509.02 | AOAC 2018.11, | 210 | 0.057 | 0.058 | 0.056 | 0.057 |  |  |  |  | 0 |
| CBDVA (%DW) | 509.99 | Other | 208 | 0.0539 | 0.0627 | 0.0566 | 0.0577 |  |  |  |  | 0 |

Population of Lab Values

CBG (%DW)

| CBG (%DW) | 510.99 | Other | 154 | 0.079 | 0.0668 | 0.0712 | 0.0723 | -3.11 | 0.1321 | 13 | 0.0192 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%DW) | 510.99 | Other | 196 | 0.1098 | 0.1183 | 0.1105 | 0.113 | -1 | 0.1321 | 13 | 0.0192 | 0 |
| CBG (%DW) | 510.02 | AOAC 2018.11, | 103 | 0.123 | 0.114 | 0.124 | 0.12 | -0.61 | 0.1321 | 13 | 0.0192 | 0 |
| CBG (%DW) | 510.99 | Other | 105 | 0.134 | 0.115 | 0.118 | 0.122 | -0.51 | 0.1321 | 13 | 0.0192 | 0 |
| CBG (%DW) | 510.99 | Other | 118 | 0.127 | 0.125 | 0.119 | 0.124 | -0.44 | 0.1321 | 13 | 0.0192 | 0 |
| CBG (%DW) | 510.99 | Other | 109 | 0.131 | 0.12 | 0.125 | 0.125 | -0.35 | 0.1321 | 13 | 0.0192 | 0 |
| CBG (%DW) | 510.99 | Other | 142 | 0.1199 | 0.1338 | 0.1288 | 0.127 | -0.24 | 0.1321 | 13 | 0.0192 | 0 |
| CBG (%DW) | 510.02 | AOAC 2018.11, | 210 | 0.134 | 0.134 | 0.137 | 0.135 | 0.15 | 0.1321 | 13 | 0.0192 | 0 |
| CBG (%DW) | 510.99 | Other | 208 | 0.136 | 0.1312 | 0.144 | 0.137 | 0.26 | 0.1321 | 13 | 0.0192 | 0 |
| CBG (%DW) | 510.99 | Other | 211 | 0.13 | 0.14 | 0.154 | 0.141 | 0.48 | 0.1321 | 13 | 0.0192 | 0 |
| CBG (%DW) | 510.99 | Other | 189 | 0.1517 | 0.1522 | 0.1536 | 0.153 | 1.06 | 0.1321 | 13 | 0.0192 | 0 |
| CBG (%DW) | 510.99 | Other | 204 | 0.1495 | 0.1582 | 0.1562 | 0.155 | 1.18 | 0.1321 | 13 | 0.0192 | 0 |
| CBG (%DW) | 510.99 | Other | 192 | 0.2284 | 0.217 | 0.2171 | 0.221 | 4.62 | 0.1321 | 13 | 0.0192 | 0 |

CBGA (%DW)

| CBGA (%DW) | 511.99 | Other | 154 | 0.2235 | 0.2132 | 0.1955 | 0.211 | -2.45 | 0.2574 | 12 | 0.019 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%DW) | 511.02 | AOAC 2018.11, | 103 | 0.269 | 0.216 | 0.202 | 0.229 | -1.49 | 0.2574 | 12 | 0.019 | 0 |
| CBGA (%DW) | 511.02 | AOAC 2018.11, | 210 | 0.251 | 0.252 | 0.248 | 0.25 | -0.37 | 0.2574 | 12 | 0.019 | 0 |

| Analyte Code |  | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%DW) | 511.99 | Other | 118 | 0.334 | 0.213 | 0.206 | 0.251 | -0.34 | 0.2574 | 12 | 0.019 | 0 |
| CBGA (%DW) | 511.99 | Other | 196 | 0.2537 | 0.2647 | 0.2461 | 0.255 | -0.14 | 0.2574 | 12 | 0.019 | 0 |
| CBGA (%DW) | 511.99 | Other | 189 | 0.258 | 0.2577 | 0.2563 | 0.257 | 0 | 0.2574 | 12 | 0.019 | 0 |
| CBGA (%DW) | 511.99 | Other | 192 | 0.2766 | 0.2489 | 0.2501 | 0.259 | 0.06 | 0.2574 | 12 | 0.019 | 0 |
| CBGA (%DW) | 511.99 | Other | 208 | 0.2663 | 0.2522 | 0.2661 | 0.262 | 0.22 | 0.2574 | 12 | 0.019 | 0 |
| CBGA (%DW) | 511.99 | Other | 204 | 0.2605 | 0.276 | 0.2648 | 0.267 | 0.51 | 0.2574 | 12 | 0.019 | 0 |
| CBGA (%DW) | 511.99 | Other | 211 | 0.262 | 0.271 | 0.276 | 0.27 | 0.64 | 0.2574 | 12 | 0.019 | 0 |
| CBGA (%DW) | 511.99 | Other | 142 | 0.261 | 0.2817 | 0.2819 | 0.275 | 0.92 | 0.2574 | 12 | 0.019 | 0 |
| CBGA (%DW) | 511.99 | Other | 109 | 0.319 | 0.272 | 0.283 | 0.291 | 1.78 | 0.2574 | 12 | 0.019 | 0 |

Population of Lab Values

THCV (%DW)

| THCV (%DW) | 512.99 | Other | 189 | <0.0050 | <0.0050 | <0.0050 | <0.0050 | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| THCV (%DW) | 512.99 | Other | 142 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| THCV (%DW) | 512.99 | Other | 109 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| THCV (%DW) | 512.99 | Other | 211 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| THCV (%DW) | 512.02 | AOAC 2018.11, | 210 | <0.0140 | <0.0140 | <0.0140 | <0.0140 | 1 |
| THCV (%DW) | 512.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 | <0.0205 | 1 |
| THCV (%DW) | 512.99 | Other | 208 | <0.0500 | <0.0500 | <0.0500 | <0.0500 | 1 |

|  |  |  | Δ8-THC (%DW) |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%DW) | 513.99 | Other | 118 | <0.0050 <0.0050 <0.0050 | <0.0050 | 1 |
| Δ8-THC (%DW) | 513.99 | Other | 211 | <0.0100 <0.0100 <0.0100 | <0.0100 | 1 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%DW) | 513.99 | Other | 142 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Δ8-THC (%DW) | 513.99 | Other | 109 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Δ8-THC (%DW) | 513.02 | AOAC 2018.11, | 210 | <0.0140 | <0.0140 | <0.0140 | <0.0140 |  |  |  |  | 1 |
| Δ8-THC (%DW) | 513.99 | Other | 102 | <0.0200 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  | 1 |
| Δ8-THC (%DW) | 513.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 | <0.0205 |  |  |  |  | 1 |
| Δ8-THC (%DW) | 513.99 | Other | 192 | <0.0300 | <0.0300 | <0.0300 | <0.0300 |  |  |  |  | 1 |
| Δ8-THC (%DW) | 513.99 | Other | 208 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| Δ8-THC (%DW) | 513.99 | Other | 189 | 0.0088 | 0.0089 | 0.0089 | 0.0089 |  |  |  |  | 0 |

Population of Lab Values

CBC (%DW)

| CBC (%DW) | 514.99 | Other | 154 | 0.1984 | 0.1879 | 0.1626 | 0.183 | -3.35 | 0.2536 | 11 | 0.0211 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%DW) | 514.99 | Other | 142 | 0.2302 | 0.241 | 0.2328 | 0.235 | -0.9 | 0.2536 | 11 | 0.0211 | 0 |
| CBC (%DW) | 514.02 | AOAC 2018.11, | 103 | 0.241 | 0.226 | 0.246 | 0.238 | -0.76 | 0.2536 | 11 | 0.0211 | 0 |
| CBC (%DW) | 514.99 | Other | 189 | 0.2498 | 0.2491 | 0.2478 | 0.249 | -0.22 | 0.2536 | 11 | 0.0211 | 0 |
| CBC (%DW) | 514.99 | Other | 105 | 0.259 | 0.245 | 0.245 | 0.25 | -0.19 | 0.2536 | 11 | 0.0211 | 0 |
| CBC (%DW) | 514.99 | Other | 208 | 0.2525 | 0.2437 | 0.2609 | 0.252 | -0.06 | 0.2536 | 11 | 0.0211 | 0 |
| CBC (%DW) | 514.99 | Other | 109 | 0.282 | 0.222 | 0.254 | 0.253 | -0.05 | 0.2536 | 11 | 0.0211 | 0 |
| CBC (%DW) | 514.99 | Other | 196 | 0.2565 | 0.2778 | 0.2493 | 0.261 | 0.36 | 0.2536 | 11 | 0.0211 | 0 |
| CBC (%DW) | 514.99 | Other | 204 | 0.269 | 0.2708 | 0.2704 | 0.27 | 0.78 | 0.2536 | 11 | 0.0211 | 0 |
| CBC (%DW) | 514.99 | Other | 211 | 0.273 | 0.274 | 0.279 | 0.275 | 1.03 | 0.2536 | 11 | 0.0211 | 0 |
| CBC (%DW) | 514.99 | Other | 192 | 0.4043 | 0.3664 | 0.3691 | 0.38 | 6 | 0.2536 | 11 | 0.0211 | 0 |

|  |  |  | Lab |  | Lab |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Analyte | Code | Method | Num | Result1 Result2 Result3 | Value | Z score | Rob Mean | n | Rob StDev Flag |

Population of Lab Values

- Flag definitions: 1 (Lab Value not used because there were less than 2 numeric lab results), 2 (Lab Value not used because there were less than 2 numeric lab results greater than 0).

- Robust statistics performed to determine robust mean and Z scores only if there were 6 or more numeric lab values.

- AR refers to concentration on as-received basis. DW refers to concentration on dry weight basis.