Hemp Proficiency Testing All Labs Precision HM21SEP-1 (hemp)

Sample: HM21SEP-1 (hemp)

All Labs Precision

Issue Date:

October 24, 2021

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

Δ9-THC (%AR)

RSDr	HorRat(r)

Population of Lab RSDr Rob Mean	min	max

Flag

| Δ9-THC (%AR) | 001.10 | Other LC-UV | 204 | 0.2173 | 0.2157 | 0.2167 | 0.37 | 0.07 | 2.6 | 0.37 | 17.1 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 148 | 0.172 | 0.174 | 0.172 | 0.67 | 0.13 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 149 | 0.213 | 0.214 | 0.216 | 0.71 | 0.14 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 211 | 0.223 | 0.221 | 0.219 | 0.9 | 0.18 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 155 | 0.183 | 0.186 | 0.183 | 0.94 | 0.18 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.03 | AOAC 2018.11, | 139 | 0.2436 | 0.2411 | 0.2389 | 0.98 | 0.2 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 149 | 0.21 | 0.209 | 0.206 | 1 | 0.2 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 194 | 0.2253 | 0.2221 | 0.2265 | 1.01 | 0.2 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 216 | 0.1958 | 0.1921 | 0.1969 | 1.29 | 0.25 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 122 | 0.2295 | 0.2276 | 0.2231 | 1.45 | 0.29 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 196 | 0.2361 | 0.2375 | 0.2431 | 1.55 | 0.31 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.99 | Other | 183 | 0.2214 | 0.2275 | 0.2282 | 1.66 | 0.33 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 197 | 0.1932 | 0.1904 | 0.1867 | 1.72 | 0.34 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 217 | 0.2358 | 0.2365 | 0.229 | 1.77 | 0.36 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 162 | 0.2245 | 0.2198 | 0.2284 | 1.92 | 0.38 | 2.6 | 0.37 | 17.1 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 142 | 0.2176 | 0.2111 | 0.2099 | 1.95 | 0.39 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 191 | 0.2103 | 0.2102 | 0.2174 | 1.94 | 0.39 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 113 | 0.2193 | 0.219 | 0.2268 | 1.99 | 0.4 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 212 | 0.1811 | 0.1765 | 0.1839 | 2.07 | 0.4 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 163 | 0.223 | 0.229 | 0.233 | 2.2 | 0.44 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, | 108 | 0.1907 | 0.1877 | 0.1822 | 2.31 | 0.45 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 101 | 0.2346 | 0.2442 | 0.2441 | 2.29 | 0.46 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, | 199 | 0.2175 | 0.2145 | 0.2245 | 2.34 | 0.47 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 109 | 0.223 | 0.22 | 0.213 | 2.35 | 0.47 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 215 | 0.2159 | 0.2162 | 0.2072 | 2.4 | 0.48 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.99 | Other | 132 | 0.224 | 0.231 | 0.235 | 2.42 | 0.49 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 106 | 0.23 | 0.22 | 0.22 | 2.59 | 0.52 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 207 | 0.245 | 0.237 | 0.233 | 2.56 | 0.52 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 115 | 0.21 | 0.22 | 0.21 | 2.71 | 0.54 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 144 | 0.2182 | 0.2079 | 0.2088 | 2.7 | 0.54 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 200 | 0.225 | 0.222 | 0.213 | 2.84 | 0.57 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.60 | NIR | 187 | 0.2 | 0.2 | 0.19 | 2.94 | 0.58 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, | 214 | 0.2 | 0.2 | 0.19 | 2.94 | 0.58 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, | 161 | 0.1826 | 0.1944 | 0.1865 | 3.2 | 0.62 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 182 | 0.195 | 0.2002 | 0.1876 | 3.26 | 0.64 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 165 | 0.2226 | 0.2128 | 0.2266 | 3.22 | 0.64 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 192 | 0.1723 | 0.1701 | 0.1813 | 3.4 | 0.66 | 2.6 | 0.37 | 17.1 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 116 | 0.2434 | 0.2537 | 0.2378 | 3.29 | 0.67 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 124 | 0.1863 | 0.2001 | 0.1948 | 3.59 | 0.7 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 184 | 0.2275 | 0.2248 | 0.24 | 3.51 | 0.71 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 164 | 0.2004 | 0.1983 | 0.1871 | 3.66 | 0.72 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, | 152 | 0.2632 | 0.2728 | 0.2828 | 3.59 | 0.74 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 172 | 0.266 | 0.271 | 0.252 | 3.74 | 0.77 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 209 | 0.2228 | 0.2405 | 0.2394 | 4.23 | 0.85 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 102 | 0.2108 | 0.1944 | 0.208 | 4.29 | 0.85 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, | 125 | 0.193 | 0.201 | 0.212 | 4.72 | 0.93 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 208 | 0.198 | 0.219 | 0.2266 | 6.91 | 1.37 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 175 | 0.1997 | 0.2277 | 0.227 | 7.32 | 1.46 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 120 | 0.267 | 0.2313 | 0.2438 | 7.32 | 1.49 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, | 103 | 0.205 | 0.204 | 0.238 | 8.97 | 1.79 | 2.6 | 0.37 | 17.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 141 | 0.3259 | 0.2413 | 0.3343 | 17.1 | 3.58 | 2.6 | 0.37 | 17.1 | 0 |
|  |  |  |  | Δ9-THCA (%AR) |  |  |  |  |  |  |  |  |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, | 214 | 0.11 | 0.11 | 0.11 | 0 | 0 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 106 | 0.13 | 0.13 | 0.13 | 0 | 0 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 209 | 0.1145 | 0.1148 | 0.1145 | 0.15 | 0.03 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 115 | 0.095 | 0.095 | 0.094 | 0.61 | 0.11 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 122 | 0.1092 | 0.1106 | 0.1094 | 0.69 | 0.12 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, | 199 | 0.1105 | 0.109 | 0.109 | 0.79 | 0.14 | 3.23 | 0 | 23.2 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, | 103 | 0.118 | 0.119 | 0.12 | 0.84 | 0.15 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 163 | 0.107 | 0.109 | 0.108 | 0.93 | 0.17 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 116 | 0.1086 | 0.1106 | 0.1103 | 0.98 | 0.18 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 204 | 0.1215 | 0.1187 | 0.1207 | 1.2 | 0.22 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.03 | AOAC 2018.11, | 139 | 0.1101 | 0.1105 | 0.1127 | 1.26 | 0.23 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 162 | 0.0956 | 0.0934 | 0.0955 | 1.31 | 0.23 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 200 | 0.113 | 0.116 | 0.114 | 1.34 | 0.24 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 216 | 0.0934 | 0.0936 | 0.0913 | 1.37 | 0.24 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 197 | 0.0959 | 0.0988 | 0.099 | 1.77 | 0.31 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 148 | 0.105 | 0.109 | 0.106 | 1.95 | 0.35 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 149 | 0.104 | 0.107 | 0.108 | 1.96 | 0.35 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 192 | 0.1776 | 0.1754 | 0.1819 | 1.85 | 0.36 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 184 | 0.1065 | 0.1049 | 0.1092 | 2.03 | 0.36 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 109 | 0.12 | 0.121 | 0.116 | 2.22 | 0.4 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 113 | 0.1138 | 0.1174 | 0.1122 | 2.33 | 0.42 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 149 | 0.112 | 0.113 | 0.108 | 2.38 | 0.43 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, | 125 | 0.119 | 0.125 | 0.122 | 2.46 | 0.45 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 211 | 0.099 | 0.097 | 0.094 | 2.6 | 0.46 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 191 | 0.1263 | 0.1247 | 0.1311 | 2.62 | 0.48 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 164 | 0.0885 | 0.0883 | 0.0928 | 2.83 | 0.49 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 120 | 0.1139 | 0.1091 | 0.1082 | 2.78 | 0.5 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 194 | 0.1218 | 0.1152 | 0.1169 | 2.9 | 0.53 | 3.23 | 0 | 23.2 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 101 | 0.1142 | 0.1178 | 0.1212 | 2.97 | 0.54 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, | 152 | 0.1091 | 0.1024 | 0.1075 | 3.29 | 0.59 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 196 | 0.1052 | 0.1074 | 0.1005 | 3.38 | 0.6 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, | 108 | 0.0864 | 0.0858 | 0.081 | 3.51 | 0.61 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 208 | 0.0979 | 0.105 | 0.1034 | 3.65 | 0.65 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 142 | 0.1013 | 0.1014 | 0.095 | 3.69 | 0.65 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, | 210 | 0.124 | 0.117 | 0.116 | 3.66 | 0.67 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 172 | 0.136 | 0.144 | 0.134 | 3.83 | 0.71 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.99 | Other | 132 | 0.098 | 0.102 | 0.094 | 4.08 | 0.72 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.60 | NIR | 187 | 0.13 | 0.14 | 0.13 | 4.33 | 0.8 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.99 | Other | 183 | 0.1204 | 0.1287 | 0.1311 | 4.43 | 0.81 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 217 | 0.1094 | 0.0986 | 0.1038 | 5.2 | 0.93 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 175 | 0.1063 | 0.0963 | 0.0983 | 5.28 | 0.94 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 144 | 0.1126 | 0.1179 | 0.1054 | 5.6 | 1.01 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, | 161 | 0.0823 | 0.0919 | 0.0833 | 6.15 | 1.07 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 165 | 0.1236 | 0.1147 | 0.1298 | 6.19 | 1.13 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 212 | 0.1037 | 0.117 | 0.1158 | 6.56 | 1.18 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 207 | 0.104 | 0.1 | 0.091 | 6.77 | 1.2 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 182 | 0.1155 | 0.101 | 0.1022 | 7.58 | 1.36 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 124 | 0.0724 | 0.0816 | 0.0869 | 9.14 | 1.57 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 155 | 0.073 | 0.072 | 0.06 | 10.6 | 1.77 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 215 | 0.0964 | 0.1248 | 0.1078 | 13 | 2.34 | 3.23 | 0 | 23.2 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 213 | 0.0582 | 0.0925 | 0.0809 | 22.6 | 3.86 | 3.23 | 0 | 23.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 141 | 0.0778 | 0.0561 | 0.0904 | 23.2 | 3.94 | 3.23 | 0 | 23.2 | 0 |
|  |  |  |  | CBD (%AR) |  |  |  |  |  |  |  |  |
| CBD (%AR) | 003.03 | AOAC 2018.11, | 139 | 3.4059 | 3.375 | 3.3557 | 0.75 | 0.23 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 122 | 3.3617 | 3.3476 | 3.312 | 0.77 | 0.23 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 196 | 3.6931 | 3.7455 | 3.768 | 1.03 | 0.31 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 216 | 3.8279 | 3.8266 | 3.7559 | 1.08 | 0.33 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 142 | 3.4957 | 3.4493 | 3.4186 | 1.12 | 0.34 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 204 | 3.4322 | 3.3529 | 3.4049 | 1.19 | 0.36 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 194 | 3.5104 | 3.4094 | 3.4526 | 1.47 | 0.44 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.60 | NIR | 187 | 2.98 | 3.06 | 2.97 | 1.64 | 0.49 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 197 | 3.6742 | 3.7263 | 3.6009 | 1.72 | 0.52 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 162 | 3.5422 | 3.417 | 3.4879 | 1.8 | 0.54 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 192 | 3.5633 | 3.6338 | 3.6919 | 1.77 | 0.54 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, | 103 | 3.313 | 3.292 | 3.415 | 1.97 | 0.59 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 149 | 3.247 | 3.319 | 3.187 | 2.03 | 0.61 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 113 | 3.503 | 3.417 | 3.556 | 2.01 | 0.61 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 109 | 3.798 | 3.67 | 3.669 | 2 | 0.61 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, | 199 | 3.544 | 3.484 | 3.629 | 2.05 | 0.62 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.99 | Other | 132 | 3.206 | 3.195 | 3.322 | 2.17 | 0.65 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 191 | 3.2825 | 3.3017 | 3.1708 | 2.17 | 0.65 | 2.57 | 0.75 | 10.7 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%AR) | 003.02 | AOAC 2018.11, | 108 | 3.575 | 3.408 | 3.479 | 2.4 | 0.73 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 184 | 3.1882 | 3.0407 | 3.1587 | 2.49 | 0.74 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, | 214 | 3.21 | 3.29 | 3.12 | 2.65 | 0.79 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, | 125 | 3.23 | 3.37 | 3.4 | 2.72 | 0.82 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 141 | 3.3957 | 3.3703 | 3.2182 | 2.88 | 0.87 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 172 | 4.094 | 4.185 | 3.956 | 2.83 | 0.87 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 163 | 3.737 | 3.582 | 3.784 | 2.86 | 0.87 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 164 | 3.5642 | 3.38 | 3.4077 | 2.88 | 0.87 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.01 | AOAC 2018.10 | 152 | 3.0991 | 3.2387 | 3.2784 | 2.94 | 0.88 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 208 | 3.3419 | 3.5562 | 3.4315 | 3.13 | 0.94 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 211 | 3.514 | 3.503 | 3.323 | 3.11 | 0.94 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 106 | 3.52 | 3.6 | 3.38 | 3.18 | 0.96 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 116 | 3.58 | 3.6966 | 3.4562 | 3.36 | 1.02 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, | 161 | 2.9369 | 3.0876 | 3.1419 | 3.48 | 1.03 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 215 | 3.1381 | 3.0498 | 3.2684 | 3.49 | 1.04 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 144 | 3.9277 | 3.6896 | 3.9032 | 3.41 | 1.05 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 209 | 1.9615 | 2.0375 | 2.1181 | 3.84 | 1.07 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 155 | 2.95 | 2.93 | 3.13 | 3.67 | 1.08 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.99 | Other | 183 | 3.5996 | 3.8952 | 3.6903 | 4.06 | 1.24 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 200 | 3.434 | 3.475 | 3.208 | 4.26 | 1.28 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 212 | 3.6206 | 3.7664 | 3.416 | 4.89 | 1.48 | 2.57 | 0.75 | 10.7 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 207 | 3.834 | 3.72 | 3.382 | 6.45 | 1.96 | 2.57 | 0.75 | 10.7 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%AR) | 003.30 | Other LC-MS | 149 | 4.002 | 3.502 | 3.25 | 10.7 | 3.24 | 2.57 | 0.75 | 10.7 | 0 |
|  |  |  |  | CBDA (%AR) |  |  |  |  |  |  |  |  |
| CBDA (%AR) | 004.30 | Other LC-MS | 122 | 6.3821 | 6.3352 | 6.3522 | 0.37 | 0.12 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 192 | 5.0382 | 4.9881 | 5.0445 | 0.62 | 0.2 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 204 | 6.5097 | 6.4233 | 6.5087 | 0.76 | 0.25 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 196 | 6.7663 | 6.877 | 6.7911 | 0.85 | 0.28 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 106 | 6.6 | 6.46 | 6.49 | 1.13 | 0.38 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, | 152 | 5.6985 | 5.8386 | 5.7258 | 1.29 | 0.42 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 197 | 6.5946 | 6.7194 | 6.7581 | 1.28 | 0.43 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 163 | 7.026 | 6.838 | 6.924 | 1.36 | 0.46 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.03 | AOAC 2018.11, | 139 | 5.7256 | 5.884 | 5.862 | 1.47 | 0.48 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 109 | 7.31 | 7.163 | 7.105 | 1.47 | 0.49 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.99 | Other | 132 | 5.353 | 5.419 | 5.532 | 1.67 | 0.54 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, | 214 | 6.08 | 6.18 | 5.97 | 1.73 | 0.57 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 194 | 6.3755 | 6.1359 | 6.2672 | 1.92 | 0.63 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 142 | 6.4132 | 6.3125 | 6.1426 | 2.17 | 0.72 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 184 | 5.7678 | 5.5804 | 5.8364 | 2.31 | 0.75 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 208 | 6.4238 | 6.7206 | 6.702 | 2.51 | 0.84 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 216 | 7.0611 | 7.2591 | 6.8995 | 2.55 | 0.86 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 191 | 6.9012 | 6.9721 | 6.6362 | 2.59 | 0.87 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, | 108 | 6.733 | 6.411 | 6.464 | 2.64 | 0.88 | 2.92 | 0.37 | 8.23 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDA (%AR) | 004.10 | Other LC-UV | 149 | 5.872 | 6.192 | 5.98 | 2.71 | 0.89 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 162 | 6.3685 | 6.0312 | 6.1884 | 2.72 | 0.9 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, | 199 | 6.3185 | 6.556 | 6.732 | 3.18 | 1.05 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.99 | Other | 183 | 7.0305 | 7.5002 | 7.2254 | 3.25 | 1.1 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, | 161 | 5.3622 | 5.6893 | 5.7335 | 3.62 | 1.18 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 164 | 5.8501 | 5.5106 | 5.4777 | 3.67 | 1.19 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 155 | 5.36 | 5.2 | 5.6 | 3.74 | 1.21 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, | 125 | 5.87 | 6.3 | 5.95 | 3.79 | 1.24 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.60 | NIR | 187 | 5.51 | 5.95 | 5.76 | 3.84 | 1.25 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 200 | 5.914 | 6.112 | 5.653 | 3.91 | 1.28 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 144 | 6.848 | 6.3151 | 6.5373 | 4.08 | 1.35 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 172 | 7.518 | 7.778 | 7.18 | 4 | 1.36 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 211 | 6.268 | 6.252 | 5.793 | 4.42 | 1.45 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 141 | 6.2229 | 6.0153 | 5.6982 | 4.42 | 1.45 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 116 | 5.8607 | 6.1597 | 5.5844 | 4.9 | 1.6 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 212 | 5.9201 | 6.1689 | 5.5875 | 4.95 | 1.62 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 149 | 6.099 | 5.969 | 6.689 | 6.14 | 2.02 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 207 | 6.607 | 6.441 | 5.869 | 6.14 | 2.03 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 215 | 6.0759 | 6.722 | 6.992 | 7.14 | 2.37 | 2.92 | 0.37 | 8.23 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 209 | 4.097 | 4.8303 | 4.5483 | 8.23 | 2.59 | 2.92 | 0.37 | 8.23 | 0 |

Population of Lab RSDr

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

CBN (%AR)

RSDr	HorRat(r)

Population of Lab RSDr Rob Mean	min	max

Flag

| CBN (%AR) | 005.10 | Other LC-UV | 197 | <0.0495 | <0.0485 | <0.0476 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%AR) | 005.10 | Other LC-UV | 211 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 208 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 196 | <0.0205 | <0.0205 | <0.0205 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 116 | <0.0400 | <0.0400 | <0.0400 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.30 | GC-FID | 116 | <0.0400 | <0.0400 | <0.0400 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 194 | <0.0270 | <0.0270 | <0.0270 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 149 | <0.0250 | <0.0250 | <0.0250 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 216 | <0.0437 | <0.0493 | <0.0513 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.02 | AOAC 2018.11, | 108 | 0 | 0 | 0 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.02 | AOAC 2018.11, | 199 | <0.0250 | <0.0250 | <0.0250 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 200 | 0.016 | 0.016 | 0.016 | 0 | 0 | 8.65 | 0 | 91 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 106 | 0.02 | 0.02 | 0.02 | 0 | 0 | 8.65 | 0 | 91 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, | 214 | 0.02 | 0.02 | 0.02 | 0 | 0 | 8.65 | 0 | 91 | 0 |
| CBN (%AR) | 005.20 | Other LC-MS | 122 | 0.0129 | 0.013 | 0.0129 | 0.45 | 0.06 | 8.65 | 0 | 91 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 204 | 0.0159 | 0.0157 | 0.0158 | 0.63 | 0.09 | 8.65 | 0 | 91 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 144 | 0.0136 | 0.0134 | 0.0133 | 1.14 | 0.15 | 8.65 | 0 | 91 | 0 |
| CBN (%AR) | 005.99 | Other | 183 | 0.0155 | 0.0161 | 0.0156 | 2.04 | 0.27 | 8.65 | 0 | 91 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 142 | 0.0135 | 0.0133 | 0.0129 | 2.31 | 0.3 | 8.65 | 0 | 91 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 162 | 0.014 | 0.0147 | 0.0136 | 3.95 | 0.52 | 8.65 | 0 | 91 | 0 |
| CBN (%AR) | 005.99 | Other | 132 | 0.011 | 0.012 | 0.012 | 4.95 | 0.64 | 8.65 | 0 | 91 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%AR) | 005.10 | Other LC-UV | 207 | 0.018 | 0.017 | 0.016 | 5.88 | 0.8 | 8.65 | 0 | 91 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 163 | 0.025 | 0.026 | 0.022 | 8.55 | 1.23 | 8.65 | 0 | 91 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 109 | 0.015 | 0.014 | 0.017 | 9.96 | 1.33 | 8.65 | 0 | 91 | 0 |
| CBN (%AR) | 005.03 | AOAC 2018.11, | 139 | 0.0147 | 0.0146 | 0.0121 | 10.7 | 1.41 | 8.65 | 0 | 91 | 0 |
| CBN (%AR) | 005.20 | Other LC-MS | 149 | 0.015 | 0.012 | 0.015 | 12.4 | 1.63 | 8.65 | 0 | 91 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 209 | 0.0156 | 0.0119 | 0.013 | 14.1 | 1.85 | 8.65 | 0 | 91 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 192 | 0.0409 | 0.0404 | 0.0525 | 15.4 | 2.41 | 8.65 | 0 | 91 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 141 | 0.1004 | 0.0699 | 0.1055 | 20.9 | 3.67 | 8.65 | 0 | 91 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 185 | 0.0448 | 0.0511 | 0.0784 | 30.7 | 5.03 | 8.65 | 0 | 91 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 138 | 0.0396 | 0.0347 | 0.0118 | 51.7 | 7.61 | 8.65 | 0 | 91 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, | 125 | 0.0058 | 0.0071 | 0.0278 | 91 | 11.96 | 8.65 | 0 | 91 | 0 |
|  |  |  |  | Total Δ9-THC (%AR) |  |  |  |  |  |  |  |  |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, | 214 | 0.29 | 0.29 | 0.29 | 0 | 0 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 149 | 0.301 | 0.303 | 0.301 | 0.38 | 0.08 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.03 | AOAC 2018.11, | 139 | 0.3402 | 0.338 | 0.3377 | 0.4 | 0.09 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 149 | 0.312 | 0.314 | 0.311 | 0.49 | 0.1 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 196 | 0.3284 | 0.3317 | 0.3312 | 0.54 | 0.11 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 204 | 0.3239 | 0.3198 | 0.3226 | 0.65 | 0.14 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 216 | 0.2777 | 0.2742 | 0.277 | 0.67 | 0.14 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 197 | 0.2773 | 0.277 | 0.2735 | 0.77 | 0.16 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 122 | 0.3253 | 0.3246 | 0.319 | 1.07 | 0.23 | 2.97 | 0 | 47.4 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 116 | 0.3111 | 0.3063 | 0.3046 | 1.1 | 0.23 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 148 | 0.264 | 0.27 | 0.265 | 1.21 | 0.25 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 194 | 0.3321 | 0.3231 | 0.3291 | 1.4 | 0.3 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 211 | 0.31 | 0.306 | 0.301 | 1.48 | 0.31 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, | 199 | 0.3144 | 0.3101 | 0.3201 | 1.59 | 0.34 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 164 | 0.289 | 0.2867 | 0.2799 | 1.66 | 0.34 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 162 | 0.3083 | 0.3017 | 0.3121 | 1.71 | 0.36 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.99 | Other | 132 | 0.31 | 0.321 | 0.318 | 1.8 | 0.38 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 200 | 0.32 | 0.32 | 0.31 | 1.82 | 0.38 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 217 | 0.3317 | 0.323 | 0.3201 | 1.86 | 0.39 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 115 | 0.29 | 0.3 | 0.29 | 1.97 | 0.41 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 163 | 0.357 | 0.361 | 0.372 | 2.14 | 0.46 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 191 | 0.3211 | 0.3195 | 0.3324 | 2.17 | 0.46 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 142 | 0.3064 | 0.3 | 0.2932 | 2.2 | 0.46 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 212 | 0.272 | 0.2791 | 0.2855 | 2.42 | 0.5 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 101 | 0.3347 | 0.3475 | 0.3504 | 2.43 | 0.52 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 109 | 0.329 | 0.326 | 0.314 | 2.46 | 0.52 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 116 | 0.3387 | 0.3507 | 0.3345 | 2.46 | 0.53 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.99 | Other | 183 | 0.327 | 0.3403 | 0.3431 | 2.55 | 0.54 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, | 108 | 0.2665 | 0.263 | 0.2533 | 2.62 | 0.54 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 128 | 0.3505 | 0.3477 | 0.3644 | 2.53 | 0.54 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 144 | 0.317 | 0.3113 | 0.3013 | 2.56 | 0.54 | 2.97 | 0 | 47.4 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, | 152 | 0.3589 | 0.3751 | 0.3771 | 2.69 | 0.58 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 155 | 0.247 | 0.238 | 0.233 | 2.96 | 0.6 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 209 | 0.3232 | 0.3412 | 0.3399 | 3 | 0.64 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 170 | 0.385 | 0.3653 | 0.366 | 3 | 0.65 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.60 | NIR | 187 | 0.31 | 0.32 | 0.3 | 3.23 | 0.68 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 192 | 0.321 | 0.324 | 0.3409 | 3.26 | 0.69 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, | 125 | 0.298 | 0.311 | 0.318 | 3.28 | 0.69 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 182 | 0.2963 | 0.2888 | 0.2772 | 3.35 | 0.7 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 207 | 0.337 | 0.325 | 0.313 | 3.69 | 0.78 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 175 | 0.2929 | 0.3122 | 0.3132 | 3.74 | 0.78 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 172 | 0.3853 | 0.3973 | 0.3695 | 3.63 | 0.79 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, | 161 | 0.2548 | 0.275 | 0.2596 | 4.01 | 0.82 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 165 | 0.331 | 0.3134 | 0.3404 | 4.18 | 0.89 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 190 | 0.51 | 0.5 | 0.47 | 4.22 | 0.95 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 169 | 0.2558 | 0.241 | 0.2645 | 4.68 | 0.96 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 185 | 0.2054 | 0.2261 | 0.2142 | 4.83 | 0.96 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 215 | 0.3003 | 0.3256 | 0.3018 | 4.59 | 0.96 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 124 | 0.2498 | 0.2717 | 0.271 | 4.71 | 0.97 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 153 | 0.2861 | 0.2841 | 0.3128 | 5.44 | 1.14 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 208 | 0.2838 | 0.311 | 0.3173 | 5.86 | 1.23 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 121 | 0.302 | 0.334 | 0.302 | 5.91 | 1.24 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 120 | 0.3669 | 0.327 | 0.3387 | 5.96 | 1.27 | 2.97 | 0 | 47.4 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, | 103 | 0.3085 | 0.3084 | 0.3432 | 6.27 | 1.32 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 114 | 0.3124 | 0.2818 | 0.2738 | 7.04 | 1.46 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.99 | Other | 181 | 0.28 | 0.32 | 0.32 | 7.53 | 1.58 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 138 | 0.3243 | 0.2703 | 0.3311 | 10.8 | 2.27 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 141 | 0.3941 | 0.2905 | 0.4136 | 18.1 | 3.9 | 2.97 | 0 | 47.4 | 0 |
| Total Δ9-THC (%AR) | 006.60 | NIR | 138 | 0.17 | 0.41 | 0.5 | 47.4 | 10.19 | 2.97 | 0 | 47.4 | 0 |
|  |  |  |  | Total CBD (%AR) |  |  |  |  |  |  |  |  |
| Total CBD (%AR) | 007.40 | GC-FID | 185 | 7.304 | 7.336 | 7.305 | 0.25 | 0.08 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 122 | 8.9588 | 8.9036 | 8.8829 | 0.44 | 0.15 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.03 | AOAC 2018.11, | 139 | 8.4273 | 8.5353 | 8.4967 | 0.64 | 0.22 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 196 | 9.6272 | 9.7766 | 9.7238 | 0.78 | 0.27 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 192 | 7.9819 | 8.0084 | 8.116 | 0.88 | 0.3 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 197 | 9.4577 | 9.6192 | 9.5278 | 0.85 | 0.3 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 204 | 9.1412 | 8.9861 | 9.113 | 0.91 | 0.32 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, | 152 | 8.0967 | 8.3591 | 8.2999 | 1.67 | 0.57 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 109 | 10.209 | 9.953 | 9.9 | 1.65 | 0.58 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 142 | 9.1201 | 8.9854 | 8.8057 | 1.76 | 0.61 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 194 | 9.1018 | 8.7906 | 8.949 | 1.74 | 0.61 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.99 | Other | 132 | 7.905 | 7.951 | 8.178 | 1.82 | 0.62 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 216 | 10.0205 | 10.1928 | 9.8068 | 1.93 | 0.68 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, | 214 | 3.44 | 3.51 | 3.34 | 2.49 | 0.75 | 2.85 | 0.25 | 26.7 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%AR) | 007.10 | Other LC-UV | 149 | 8.397 | 8.75 | 8.431 | 2.28 | 0.79 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 163 | 9 | 9.381 | 9.339 | 2.26 | 0.79 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 170 | 9.917 | 9.48 | 9.776 | 2.29 | 0.81 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 162 | 9.1274 | 8.7063 | 8.9152 | 2.36 | 0.82 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 191 | 9.3348 | 9.4163 | 8.9907 | 2.44 | 0.85 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, | 199 | 9.0853 | 9.2336 | 9.533 | 2.46 | 0.86 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, | 108 | 9.48 | 9.03 | 9.148 | 2.53 | 0.88 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 208 | 8.9756 | 9.4501 | 9.3091 | 2.64 | 0.92 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.60 | NIR | 187 | 7.81 | 8.27 | 8.02 | 2.87 | 0.98 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, | 125 | 8.37 | 8.9 | 8.62 | 3.07 | 1.06 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 155 | 7.65 | 7.49 | 8 | 3.38 | 1.15 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 164 | 9.4143 | 8.8906 | 8.8853 | 3.35 | 1.17 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 149 | 9.352 | 8.736 | 9.116 | 3.43 | 1.2 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, | 161 | 7.6396 | 8.0771 | 8.1702 | 3.56 | 1.22 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.99 | Other | 183 | 9.7653 | 10.4729 | 10.027 | 3.55 | 1.26 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 172 | 10.6873 | 11.0063 | 10.2529 | 3.55 | 1.27 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 144 | 9.9348 | 9.2292 | 9.6377 | 3.69 | 1.3 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 141 | 8.8533 | 8.6457 | 8.2156 | 3.79 | 1.31 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 211 | 9.011 | 8.986 | 8.404 | 3.9 | 1.36 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 200 | 8.62 | 8.84 | 8.17 | 4 | 1.38 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 116 | 8.7198 | 9.0986 | 8.3538 | 4.27 | 1.48 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 169 | 7.9868 | 7.4644 | 8.1455 | 4.53 | 1.55 | 2.85 | 0.25 | 26.7 | 0 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%AR) | 007.10 | Other LC-UV | 212 | 8.8126 | 9.1765 | 8.3163 | 4.92 | 1.71 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 215 | 8.4667 | 8.9449 | 9.4004 | 5.22 | 1.82 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 116 | 8.468 | 7.8427 | 8.8616 | 6.12 | 2.11 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 207 | 9.628 | 9.369 | 8.529 | 6.26 | 2.19 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 138 | 8.3248 | 9.6343 | 9.2657 | 7.44 | 2.6 | 2.85 | 0.25 | 26.7 | 0 |
| Total CBD (%AR) | 007.60 | NIR | 138 | 5.87 | 9.32 | 10.09 | 26.7 | 9.2 | 2.85 | 0.25 | 26.7 | 0 |

Lab

Population of Lab RSDr

CBDV (%AR)

| CBDV (%AR) | 008.99 | Other | 109 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDV (%AR) | 008.99 | Other | 211 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| CBDV (%AR) | 008.99 | Other | 162 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| CBDV (%AR) | 008.99 | Other | 142 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| CBDV (%AR) | 008.99 | Other | 216 | <0.0437 | <0.0493 | <0.0513 |  |  |  |  |  | 1 |
| CBDV (%AR) | 008.99 | Other | 194 | <0.0270 | <0.0270 | <0.0270 |  |  |  |  |  | 1 |
| CBDV (%AR) | 008.02 | AOAC 2018.11, | 108 | 0 | 0 | 0 |  |  |  |  |  | 1 |
| CBDV (%AR) | 008.99 | Other | 197 | <0.0495 | <0.0485 | <0.0476 |  |  |  |  |  | 1 |
| CBDV (%AR) | 008.02 | AOAC 2018.11, | 214 | 0.02 | 0.02 | 0.02 | 0 | 0 | 4.82 | 0 | 41.9 | 0 |
| CBDV (%AR) | 008.99 | Other | 106 | 0.05 | 0.05 | 0.05 | 0 | 0 | 4.82 | 0 | 41.9 | 0 |
| CBDV (%AR) | 008.99 | Other | 149 | 0.027 | 0.028 | 0.028 | 2.09 | 0.31 | 4.82 | 0 | 41.9 | 0 |
| CBDV (%AR) | 008.99 | Other | 207 | 0.024 | 0.023 | 0.023 | 2.47 | 0.35 | 4.82 | 0 | 41.9 | 0 |
| CBDV (%AR) | 008.99 | Other | 183 | 0.034 | 0.0362 | 0.0366 | 3.93 | 0.6 | 4.82 | 0 | 41.9 | 0 |
| CBDV (%AR) | 008.99 | Other | 200 | 0.032 | 0.031 | 0.029 | 4.98 | 0.74 | 4.82 | 0 | 41.9 | 0 |

Lab

CBDVA (%AR)

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDV (%AR) | 008.99 | Other | 196 | 0.0241 | 0.0246 | 0.0267 | 5.49 | 0.79 | 4.82 | 0 | 41.9 | 0 |
| CBDV (%AR) | 008.99 | Other | 208 | 0.0736 | 0.0932 | 0.0834 | 11.8 | 2.03 | 4.82 | 0 | 41.9 | 0 |
| CBDV (%AR) | 008.99 | Other | 132 | 0.024 | 0.023 | 0.046 | 41.9 | 6.24 | 4.82 | 0 | 41.9 | 0 |

Lab

Population of Lab RSDr

| CBDVA (%AR) | 009.99 | Other | 197 | <0.0743 | <0.0728 | <0.0714 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDVA (%AR) | 009.99 | Other | 216 | <0.0655 | <0.0739 | <0.0769 |  |  |  |  |  | 1 |
| CBDVA (%AR) | 009.02 | AOAC 2018.11, | 214 | 0.07 | 0.07 | 0.07 | 0 | 0 | 3.31 | 0 | 8.01 | 0 |
| CBDVA (%AR) | 009.99 | Other | 149 | 0.074 | 0.075 | 0.074 | 0.78 | 0.13 | 3.31 | 0 | 8.01 | 0 |
| CBDVA (%AR) | 009.99 | Other | 196 | 0.0648 | 0.0667 | 0.0652 | 1.53 | 0.25 | 3.31 | 0 | 8.01 | 0 |
| CBDVA (%AR) | 009.99 | Other | 183 | 0.0811 | 0.0853 | 0.0846 | 2.69 | 0.46 | 3.31 | 0 | 8.01 | 0 |
| CBDVA (%AR) | 009.99 | Other | 208 | 0.109 | 0.1172 | 0.114 | 3.64 | 0.66 | 3.31 | 0 | 8.01 | 0 |
| CBDVA (%AR) | 009.99 | Other | 212 | 0.0621 | 0.0625 | 0.0555 | 6.55 | 1.08 | 3.31 | 0 | 8.01 | 0 |
| CBDVA (%AR) | 009.02 | AOAC 2018.11, | 108 | 0.0519 | 0.0529 | 0.0455 | 8.01 | 1.28 | 3.31 | 0 | 8.01 | 0 |

| CBG (%AR) |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%AR) | 010.02 | AOAC 2018.11, | 214 | 0.12 | 0.12 | 0.12 | 0 | 0 | 3.71 | 0 | 13.9 | 0 |
| CBG (%AR) | 010.99 | Other | 106 | 0.14 | 0.14 | 0.14 | 0 | 0 | 3.71 | 0 | 13.9 | 0 |
| CBG (%AR) | 010.99 | Other | 216 | 0.0803 | 0.0816 | 0.0812 | 0.82 | 0.14 | 3.71 | 0 | 13.9 | 0 |
| CBG (%AR) | 010.99 | Other | 183 | 0.1313 | 0.1322 | 0.1341 | 1.08 | 0.2 | 3.71 | 0 | 13.9 | 0 |
| CBG (%AR) | 010.99 | Other | 194 | 0.1385 | 0.1368 | 0.1397 | 1.05 | 0.2 | 3.71 | 0 | 13.9 | 0 |
| CBG (%AR) | 010.02 | AOAC 2018.11, | 108 | 0.1012 | 0.1037 | 0.1034 | 1.33 | 0.24 | 3.71 | 0 | 13.9 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%AR) | 010.99 | Other | 208 | 0.1271 | 0.1289 | 0.125 | 1.54 | 0.28 | 3.71 | 0 | 13.9 | 0 |
| CBG (%AR) | 010.99 | Other | 142 | 0.1266 | 0.1238 | 0.1228 | 1.58 | 0.29 | 3.71 | 0 | 13.9 | 0 |
| CBG (%AR) | 010.99 | Other | 196 | 0.1135 | 0.1156 | 0.1177 | 1.82 | 0.33 | 3.71 | 0 | 13.9 | 0 |
| CBG (%AR) | 010.99 | Other | 132 | 0.164 | 0.17 | 0.167 | 1.8 | 0.34 | 3.71 | 0 | 13.9 | 0 |
| CBG (%AR) | 010.99 | Other | 204 | 0.1499 | 0.1435 | 0.1503 | 2.58 | 0.49 | 3.71 | 0 | 13.9 | 0 |
| CBG (%AR) | 010.99 | Other | 200 | 0.162 | 0.16 | 0.153 | 2.98 | 0.57 | 3.71 | 0 | 13.9 | 0 |
| CBG (%AR) | 010.99 | Other | 207 | 0.123 | 0.114 | 0.121 | 3.96 | 0.72 | 3.71 | 0 | 13.9 | 0 |
| CBG (%AR) | 010.02 | AOAC 2018.11, | 125 | 0.152 | 0.163 | 0.164 | 4.17 | 0.79 | 3.71 | 0 | 13.9 | 0 |
| CBG (%AR) | 010.99 | Other | 149 | 0.105 | 0.113 | 0.103 | 4.95 | 0.89 | 3.71 | 0 | 13.9 | 0 |
| CBG (%AR) | 010.99 | Other | 109 | 0.123 | 0.116 | 0.128 | 4.93 | 0.9 | 3.71 | 0 | 13.9 | 0 |
| CBG (%AR) | 010.99 | Other | 197 | 0.0845 | 0.0805 | 0.0746 | 6.24 | 1.07 | 3.71 | 0 | 13.9 | 0 |
| CBG (%AR) | 010.99 | Other | 162 | 0.1073 | 0.1214 | 0.1171 | 6.27 | 1.14 | 3.71 | 0 | 13.9 | 0 |
| CBG (%AR) | 010.99 | Other | 212 | 0.0326 | 0.039 | 0.0351 | 9.07 | 1.38 | 3.71 | 0 | 13.9 | 0 |
| CBG (%AR) | 010.99 | Other | 155 | 0.116 | 0.109 | 0.128 | 8.17 | 1.48 | 3.71 | 0 | 13.9 | 0 |
| CBG (%AR) | 010.99 | Other | 141 | 0.0827 | 0.0894 | 0.107 | 13.5 | 2.37 | 3.71 | 0 | 13.9 | 0 |
| CBG (%AR) | 010.99 | Other | 211 | 0.126 | 0.126 | 0.098 | 13.9 | 2.52 | 3.71 | 0 | 13.9 | 0 |
|  |  |  |  | CBGA (%AR) |  |  |  |  |  |  |  |  |
| CBGA (%AR) | 011.99 | Other | 194 | 0.3044 | 0.3049 | 0.3028 | 0.36 | 0.08 | 4.75 | 0.36 | 14 | 0 |
| CBGA (%AR) | 011.99 | Other | 196 | 0.297 | 0.3025 | 0.3032 | 1.13 | 0.24 | 4.75 | 0.36 | 14 | 0 |
| CBGA (%AR) | 011.99 | Other | 204 | 0.3031 | 0.2964 | 0.3062 | 1.66 | 0.35 | 4.75 | 0.36 | 14 | 0 |
| CBGA (%AR) | 011.99 | Other | 216 | 0.2541 | 0.2631 | 0.2626 | 1.95 | 0.4 | 4.75 | 0.36 | 14 | 0 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%AR) | 011.02 | AOAC 2018.11, | 108 | 0.2647 | 0.2607 | 0.2544 | 2 | 0.41 | 4.75 | 0.36 | 14 | 0 |
| CBGA (%AR) | 011.99 | Other | 132 | 0.229 | 0.237 | 0.228 | 2.13 | 0.43 | 4.75 | 0.36 | 14 | 0 |
| CBGA (%AR) | 011.02 | AOAC 2018.11, | 214 | 0.26 | 0.25 | 0.25 | 2.28 | 0.46 | 4.75 | 0.36 | 14 | 0 |
| CBGA (%AR) | 011.99 | Other | 142 | 0.3073 | 0.307 | 0.2935 | 2.6 | 0.55 | 4.75 | 0.36 | 14 | 0 |
| CBGA (%AR) | 011.99 | Other | 208 | 0.2891 | 0.3005 | 0.3093 | 3.38 | 0.71 | 4.75 | 0.36 | 14 | 0 |
| CBGA (%AR) | 011.99 | Other | 197 | 0.2667 | 0.2883 | 0.279 | 3.9 | 0.81 | 4.75 | 0.36 | 14 | 0 |
| CBGA (%AR) | 011.99 | Other | 109 | 0.365 | 0.342 | 0.338 | 4.18 | 0.89 | 4.75 | 0.36 | 14 | 0 |
| CBGA (%AR) | 011.99 | Other | 106 | 0.26 | 0.24 | 0.24 | 4.68 | 0.95 | 4.75 | 0.36 | 14 | 0 |
| CBGA (%AR) | 011.99 | Other | 212 | 0.2177 | 0.2433 | 0.217 | 6.63 | 1.33 | 4.75 | 0.36 | 14 | 0 |
| CBGA (%AR) | 011.99 | Other | 141 | 0.0894 | 0.0773 | 0.0902 | 8.44 | 1.46 | 4.75 | 0.36 | 14 | 0 |
| CBGA (%AR) | 011.02 | AOAC 2018.11, | 125 | 0.24 | 0.281 | 0.263 | 7.86 | 1.61 | 4.75 | 0.36 | 14 | 0 |
| CBGA (%AR) | 011.99 | Other | 207 | 0.263 | 0.254 | 0.223 | 8.51 | 1.73 | 4.75 | 0.36 | 14 | 0 |
| CBGA (%AR) | 011.99 | Other | 162 | 0.2705 | 0.3021 | 0.2539 | 8.89 | 1.84 | 4.75 | 0.36 | 14 | 0 |
| CBGA (%AR) | 011.99 | Other | 183 | 0.2903 | 0.3067 | 0.3478 | 9.41 | 1.98 | 4.75 | 0.36 | 14 | 0 |
| CBGA (%AR) | 011.99 | Other | 149 | 0.222 | 0.293 | 0.275 | 14 | 2.88 | 4.75 | 0.36 | 14 | 0 |

Lab

Population of Lab RSDr

THCV (%AR)

| THCV (%AR) | 012.99 | Other | 183 | 0 | 0 | 0 | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| THCV (%AR) | 012.99 | Other | 106 | <0.0100 | <0.0100 | <0.0100 | 1 |
| THCV (%AR) | 012.99 | Other | 200 | <0.0050 | <0.0050 | <0.0050 | 1 |
| THCV (%AR) | 012.99 | Other | 216 | <0.0437 | <0.0493 | <0.0513 | 1 |
| THCV (%AR) | 012.99 | Other | 197 | <0.0495 | <0.0485 | <0.0476 | 1 |

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) Rob Mean min max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| THCV (%AR) | 012.99 | Other | 142 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 208 | <0.0500 | <0.0500 | <0.0500 |  |  | 1 |
| THCV (%AR) | 012.02 | AOAC 2018.11, | 108 | 0 | 0 | 0 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 109 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 194 | <0.0270 | <0.0270 | <0.0270 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 149 | <0.0250 | <0.0250 | <0.0250 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 211 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 |  |  | 1 |
| THCV (%AR) | 012.02 | AOAC 2018.11, | 214 | 0.02 | 0.02 | 0.02 | 0 | 0 | 0 |
| THCV (%AR) | 012.99 | Other | 132 | 0.017 | 0.018 | 0.019 | 5.56 | 0.76 | 0 |

Lab

Population of Lab RSDr

Δ8-THC (%AR)

| Δ8-THC (%AR) | 013.99 | Other | 183 | 0 | 0 | 0 | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%AR) | 013.99 | Other | 207 | <0.0130 | <0.0130 | <0.0130 | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 197 | <0.0495 | <0.0485 | <0.0476 | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 208 | <0.0500 | <0.0500 | <0.0500 | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 200 | <0.0100 | <0.0100 | <0.0100 | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 141 | <0.0120 | <0.0120 | <0.0120 | 1 |
| Δ8-THC (%AR) | 013.02 | AOAC 2018.11, | 214 | <0.0100 | <0.0100 | <0.0100 | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 194 | <0.0270 | <0.0270 | <0.0270 | 1 |
| Δ8-THC (%AR) | 013.02 | AOAC 2018.11, | 108 | 0 | 0 | 0 | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 211 | <0.0100 | <0.0100 | <0.0100 | 1 |

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) Rob Mean min max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%AR) | 013.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 216 | <0.0437 | <0.0493 | <0.0513 |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 109 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 142 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 149 | <0.0250 | <0.0250 | <0.0250 |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 162 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 106 | 0.04 | 0.03 | 0.03 | 17.3 | 2.61 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 138 | 0.0434 | 0.0457 | 0.0327 | 17.1 | 2.65 | 0 |

Lab

Population of Lab RSDr

CBC (%AR)

| CBC (%AR) | 014.02 | AOAC 2018.11, | 214 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%AR) | 014.99 | Other | 106 | 0.23 | 0.23 | 0.23 | 0 | 0 | 2.86 | 0 | 14.2 | 0 |
| CBC (%AR) | 014.99 | Other | 194 | 0.2526 | 0.2526 | 0.2498 | 0.64 | 0.13 | 2.86 | 0 | 14.2 | 0 |
| CBC (%AR) | 014.99 | Other | 196 | 0.2764 | 0.2776 | 0.2729 | 0.89 | 0.18 | 2.86 | 0 | 14.2 | 0 |
| CBC (%AR) | 014.99 | Other | 204 | 0.2658 | 0.2612 | 0.2626 | 0.9 | 0.18 | 2.86 | 0 | 14.2 | 0 |
| CBC (%AR) | 014.99 | Other | 216 | 0.2498 | 0.2522 | 0.2544 | 0.91 | 0.19 | 2.86 | 0 | 14.2 | 0 |
| CBC (%AR) | 014.99 | Other | 162 | 0.2547 | 0.2499 | 0.2582 | 1.64 | 0.33 | 2.86 | 0 | 14.2 | 0 |
| CBC (%AR) | 014.99 | Other | 149 | 0.24 | 0.244 | 0.236 | 1.67 | 0.34 | 2.86 | 0 | 14.2 | 0 |
| CBC (%AR) | 014.99 | Other | 142 | 0.2446 | 0.2383 | 0.2351 | 2.02 | 0.41 | 2.86 | 0 | 14.2 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, | 103 | 0.226 | 0.228 | 0.236 | 2.3 | 0.46 | 2.86 | 0 | 14.2 | 0 |
| CBC (%AR) | 014.99 | Other | 132 | 0.227 | 0.23 | 0.238 | 2.45 | 0.49 | 2.86 | 0 | 14.2 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, | 125 | 0.232 | 0.245 | 0.239 | 2.73 | 0.55 | 2.86 | 0 | 14.2 | 0 |

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%AR) | 014.99 | Other | 197 | 0.2544 | 0.2583 | 0.2448 | 2.75 | 0.56 | 2.86 | 0 | 14.2 | 0 |
| CBC (%AR) | 014.99 | Other | 109 | 0.285 | 0.274 | 0.269 | 2.97 | 0.61 | 2.86 | 0 | 14.2 | 0 |
| CBC (%AR) | 014.99 | Other | 183 | 0.2509 | 0.2649 | 0.2671 | 3.37 | 0.69 | 2.86 | 0 | 14.2 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, | 108 | 0.2607 | 0.2635 | 0.2452 | 3.84 | 0.79 | 2.86 | 0 | 14.2 | 0 |
| CBC (%AR) | 014.99 | Other | 200 | 0.281 | 0.272 | 0.26 | 3.89 | 0.8 | 2.86 | 0 | 14.2 | 0 |
| CBC (%AR) | 014.99 | Other | 155 | 0.218 | 0.2 | 0.216 | 4.67 | 0.93 | 2.86 | 0 | 14.2 | 0 |
| CBC (%AR) | 014.99 | Other | 208 | 0.2216 | 0.2441 | 0.2297 | 4.92 | 0.99 | 2.86 | 0 | 14.2 | 0 |
| CBC (%AR) | 014.99 | Other | 207 | 0.272 | 0.266 | 0.244 | 5.66 | 1.16 | 2.86 | 0 | 14.2 | 0 |
| CBC (%AR) | 014.99 | Other | 211 | 0.256 | 0.249 | 0.223 | 7.17 | 1.45 | 2.86 | 0 | 14.2 | 0 |
| CBC (%AR) | 014.99 | Other | 212 | 0.1622 | 0.2126 | 0.1756 | 14.2 | 2.76 | 2.86 | 0 | 14.2 | 0 |

Lab

Population of Lab RSDr

Cd (ug/g AR)

| Cd (ug/g AR) | 101.04 | ICP-MS | 141 | <0.2000 | <0.2000 | <0.2000 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cd (ug/g AR) | 101.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 |  |  |  |  |  | 1 |
| Cd (ug/g AR) | 101.03 | ICP-OES | 132 | 0.104 |  |  |  |  |  |  |  | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 122 | 0.1228 | 0.1219 | 0.1235 | 0.65 | 0.12 | 4.58 | 0.65 | 10.4 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 197 | 0.1256 | 0.1288 | 0.1284 | 1.37 | 0.25 | 4.58 | 0.65 | 10.4 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 106 | 0.132 | 0.128 | 0.132 | 1.77 | 0.33 | 4.58 | 0.65 | 10.4 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 196 | 0.0991 | 0.0947 | 0.0969 | 2.27 | 0.4 | 4.58 | 0.65 | 10.4 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 194 | 0.1302 | 0.1377 | 0.1324 | 2.89 | 0.54 | 4.58 | 0.65 | 10.4 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 214 | 0.128 | 0.125 | 0.139 | 5.64 | 1.04 | 4.58 | 0.65 | 10.4 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 142 | 0.134 | 0.13 | 0.146 | 6.09 | 1.13 | 4.58 | 0.65 | 10.4 | 0 |

Lab

As (ug/g AR)

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cd (ug/g AR) | 101.04 | ICP-MS | 183 | 0.11 | 0.1285 | 0.107 | 10.1 | 1.83 | 4.58 | 0.65 | 10.4 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 211 | 0.144 | 0.122 | 0.149 | 10.4 | 1.93 | 4.58 | 0.65 | 10.4 | 0 |

Lab

Population of Lab RSDr

Hg (ug/g AR)

| As (ug/g AR) | 102.03 | ICP-OES | 132 | 0.4 |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| As (ug/g AR) | 102.04 | ICP-MS | 141 | <0.2000 | <0.2000 | <0.2000 |  |  | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 211 | <0.1000 | <0.1000 | <0.1000 |  |  | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 |  |  | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 197 | <0.0992 | <0.0980 | <0.1004 |  |  | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 142 | 0.054 | <0.0010 | <0.0010 |  |  | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 183 | <0.0020 | <0.0020 | <0.0020 |  |  | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 122 | 0.0422 | 0.0426 | 0.0424 | 0.47 | 0.07 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 106 | 0.047 | 0.045 | 0.046 | 2.17 | 0.34 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 196 | 0.0672 | 0.0712 | 0.0771 | 6.93 | 1.17 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 194 | 0.0613 | 0.0736 | 0.0686 | 9.12 | 1.53 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 214 | 0.059 | 0.065 | 0.083 | 18.1 | 3.04 | 0 |

| Hg (ug/g AR) | 103.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Hg (ug/g AR) | 103.04 | ICP-MS | 183 | <0.0120 | <0.0120 | <0.0120 | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 197 | <0.0992 | <0.0980 | <0.1004 | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 141 | <0.1000 | <0.1000 | <0.1000 | 1 |

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) Rob Mean min max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Hg (ug/g AR) | 103.04 | ICP-MS | 214 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 194 | <0.0200 | <0.0200 | <0.0200 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 142 | <0.0010 | <0.0010 | <0.0010 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 196 | <0.0072 | <0.0072 | 0.008 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 211 | <0.1000 | <0.1000 | <0.1000 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 106 | 0.007 | 0.007 | 0.007 | 0 | 0 | 0 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 122 | 0.0171 | 0.0175 | 0.0172 | 1.21 | 0.16 | 0 |

Lab

Population of Lab RSDr

Pb (ug/g AR)

| Pb (ug/g AR) | 104.02 | furnace AA | 132 | 0.4 |  |  |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Pb (ug/g AR) | 104.04 | ICP-MS | 141 | <0.5000 | <0.5000 | <0.5000 |  |  |  |  |  | 1 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 122 | 0.29 | 0.2901 | 0.2869 | 0.63 | 0.13 | 6.64 | 0.63 | 27.2 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 194 | 0.2677 | 0.2626 | 0.2598 | 1.52 | 0.31 | 6.64 | 0.63 | 27.2 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 196 | 0.282 | 0.275 | 0.292 | 3.02 | 0.63 | 6.64 | 0.63 | 27.2 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 106 | 0.276 | 0.26 | 0.28 | 3.89 | 0.8 | 6.64 | 0.63 | 27.2 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 197 | 0.2332 | 0.2624 | 0.2439 | 5.99 | 1.22 | 6.64 | 0.63 | 27.2 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 214 | 0.264 | 0.248 | 0.29 | 7.93 | 1.63 | 6.64 | 0.63 | 27.2 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 162 | 0.2906 | 0.2562 | 0.2459 | 8.86 | 1.82 | 6.64 | 0.63 | 27.2 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 142 | 0.31 | 0.261 | 0.277 | 8.84 | 1.83 | 6.64 | 0.63 | 27.2 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 211 | 0.304 | 0.241 | 0.277 | 11.5 | 2.38 | 6.64 | 0.63 | 27.2 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 183 | 0.2235 | 0.369 | 0.254 | 27.2 | 5.64 | 6.64 | 0.63 | 27.2 | 0 |

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

Cu (ug/g AR)

RSDr	HorRat(r)

Population of Lab RSDr Rob Mean	min	max

Flag

Cr (ug/g AR)

Ni (ug/g AR)

| Cu (ug/g AR) | 105.03 | ICP-OES | 132 | 10.8 |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cu (ug/g AR) | 105.04 | ICP-MS | 214 | 13.562 | 13.642 | 14.288 | 2.88 | 1.07 | 0 |
| Cu (ug/g AR) | 105.04 | ICP-MS | 196 | 14.4 | 12.7 | 14.5 | 7.3 | 2.71 | 0 |
| Cu (ug/g AR) | 105.04 | ICP-MS | 142 | 18.364 | 17.272 | 34.987 | 42.2 | 16.97 | 0 |

| Ni (ug/g AR) | 106.03 | ICP-OES | 132 | 0.8 |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Ni (ug/g AR) | 106.04 | ICP-MS | 196 | 1.11 | 1.09 | 1.12 | 1.38 | 0.35 | 0 |
| Ni (ug/g AR) | 106.04 | ICP-MS | 214 | 1.021 | 1.036 | 1.102 | 4.09 | 1.03 | 0 |
| Ni (ug/g AR) | 106.04 | ICP-MS | 142 | 1.152 | 1.082 | 1.035 | 5.4 | 1.37 | 0 |
| Ni (ug/g AR) | 106.04 | ICP-MS | 141 | 1.24 | 1.34 | 1.2 | 5.72 | 1.48 | 0 |

| Cr (ug/g AR) | 107.04 | ICP-MS | 141 | <4.0000 | <4.0000 | <4.0000 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cr (ug/g AR) | 107.03 | ICP-OES | 132 | 0.2 |  |  |  |  | 1 |
| Cr (ug/g AR) | 107.04 | ICP-MS | 197 | 0.4301 | 0.4431 | 0.4328 | 1.58 | 0.35 | 0 |
| Cr (ug/g AR) | 107.04 | ICP-MS | 196 | 0.517 | 0.477 | 0.525 | 5.08 | 1.15 | 0 |
| Cr (ug/g AR) | 107.04 | ICP-MS | 214 | 0.521 | 0.445 | 0.556 | 11.2 | 2.53 | 0 |
| Cr (ug/g AR) | 107.04 | ICP-MS | 142 | 0.392 | 0.246 | 0.279 | 25 | 5.25 | 0 |

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

Se (ug/g AR)

RSDr	HorRat(r)

Population of Lab RSDr Rob Mean	min	max

Flag

Ba (ug/g AR)

| Se (ug/g AR) | 108.04 | ICP-MS | 214 | <0.1000 | 0.101 | <0.1000 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Se (ug/g AR) | 108.03 | ICP-OES | 132 | 0.4 |  |  |  |  | 1 |
| Se (ug/g AR) | 108.04 | ICP-MS | 141 | <1.0000 | <1.0000 | <1.0000 |  |  | 1 |
| Se (ug/g AR) | 108.04 | ICP-MS | 197 | <0.0992 | <0.0980 | <0.1004 |  |  | 1 |
| Se (ug/g AR) | 108.04 | ICP-MS | 142 | 0.06 | 0.058 | 0.057 | 2.62 | 0.43 | 0 |
| Se (ug/g AR) | 108.04 | ICP-MS | 196 | 0.241 | 0.24 | 0.227 | 3.31 | 0.67 | 0 |

|  |  |  |  | Ag (ug/g AR) |  |
| --- | --- | --- | --- | --- | --- |
| Ag (ug/g AR) | 109.04 | ICP-MS | 214 | <0.0500 <0.0500 <0.0500 | 1 |
| Ag (ug/g AR) | 109.03 | ICP-OES | 132 | 0.088 | 1 |

| Ba (ug/g AR) | 110.03 | ICP-OES | 132 | 25.6 |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Ba (ug/g AR) | 110.04 | ICP-MS | 214 | 41.977 | 41.918 | 43.107 | 1.58 | 0.7 | 0 |

|  |  |  | Zn (ug/g AR) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Zn (ug/g AR) | 111.03 | ICP-OES | 132 | 44.8 |  |  |  |  | 1 |
| Zn (ug/g AR) | 111.04 | ICP-MS | 196 | 66.3 | 66.9 | 67.6 | 0.97 | 0.46 | 0 |
| Zn (ug/g AR) | 111.04 | ICP-MS | 141 | 54 | 43.4 | 59 | 15.3 | 6.93 | 0 |
| Zn (ug/g AR) | 111.04 | ICP-MS | 142 | 61.901 | 59.056 | 128.004 | 47 | 22.86 | 0 |

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

|  |  |  |  | Sb (ug/g AR) |  |
| --- | --- | --- | --- | --- | --- |
| Sb (ug/g AR) | 112.04 | ICP-MS | 196 | <0.0241 <0.0241 <0.0241 | 1 |
| Sb (ug/g AR) | 112.04 | ICP-MS | 214 | <0.0100 <0.0100 0.013 | 1 |

RSDr	HorRat(r)

Population of Lab RSDr Rob Mean	min	max

Flag

(-)-alpha-Bisabolol (%AR)

beta-Caryophyllene (%AR)

| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 108 | 0.1083 | 0.1106 | 0.108 | 1.31 | 0.23 | 6.28 | 1.31 | 14.5 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 211 | 0.122 | 0.126 | 0.123 | 1.68 | 0.31 | 6.28 | 1.31 | 14.5 | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 162 | 0.0325 | 0.0355 | 0.0355 | 5.02 | 0.76 | 6.28 | 1.31 | 14.5 | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 194 | 0.0691 | 0.0707 | 0.0786 | 6.99 | 1.18 | 6.28 | 1.31 | 14.5 | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 142 | 0.03 | 0.0335 | 0.0281 | 8.97 | 1.33 | 6.28 | 1.31 | 14.5 | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 183 | 0.4742 | 0.4424 | 0.4162 | 6.54 | 1.45 | 6.28 | 1.31 | 14.5 | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 196 | 0.0229 | 0.0278 | 0.0307 | 14.5 | 2.12 | 6.28 | 1.31 | 14.5 | 0 |

| beta-Caryophyllene (% | 202.02 | GC,MS | 211 | 0.091 | 0.093 | 0.091 | 1.26 | 0.22 | 4.48 | 0.88 | 9.81 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| beta-Caryophyllene (% | 202.02 | GC,MS | 183 | 0.9287 | 0.9125 | 0.9206 | 0.88 | 0.22 | 4.48 | 0.88 | 9.81 | 0 |
| beta-Caryophyllene (% | 202.02 | GC,MS | 162 | 0.083 | 0.0871 | 0.0877 | 2.98 | 0.52 | 4.48 | 0.88 | 9.81 | 0 |
| beta-Caryophyllene (% | 202.02 | GC,MS | 194 | 0.0978 | 0.1026 | 0.1052 | 3.69 | 0.66 | 4.48 | 0.88 | 9.81 | 0 |
| beta-Caryophyllene (% | 202.02 | GC,MS | 108 | 0.124 | 0.1205 | 0.1329 | 5.08 | 0.93 | 4.48 | 0.88 | 9.81 | 0 |
| beta-Caryophyllene (% | 202.02 | GC,MS | 196 | 0.067 | 0.0745 | 0.078 | 7.68 | 1.3 | 4.48 | 0.88 | 9.81 | 0 |
| beta-Caryophyllene (% | 202.02 | GC,MS | 142 | 0.1134 | 0.1361 | 0.1179 | 9.81 | 1.79 | 4.48 | 0.88 | 9.81 | 0 |

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

(-)-Caryophyllene oxide (%AR)

RSDr	HorRat(r)

Population of Lab RSDr Rob Mean	min	max

Flag

alpha-Humulene (%AR)

Geraniol (%AR)

| (-)-Caryophyllene oxide | 203.02 | GC,MS | 142 | 0.0077 | 0.0082 | 0.0078 | 3.35 | 0.41 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| (-)-Caryophyllene oxide | 203.02 | GC,MS | 194 | 0.013 | 0.0131 | 0.0142 | 4.96 | 0.65 | 0 |
| (-)-Caryophyllene oxide | 203.02 | GC,MS | 211 | 0.025 | 0.022 | 0.023 | 6.55 | 0.93 | 0 |
| (-)-Caryophyllene oxide | 203.02 | GC,MS | 196 | 0.0106 | 0.0126 | 0.0123 | 9.11 | 1.17 | 0 |
| (-)-Caryophyllene oxide | 203.02 | GC,MS | 183 | 0.0164 | 0.0167 | 0.0195 | 9.75 | 1.33 | 0 |

| Geraniol (%AR) | 204.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Geraniol (%AR) | 204.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 | 1 |
| Geraniol (%AR) | 204.02 | GC,MS | 142 | <0.0001 | <0.0001 | <0.0001 | 1 |
| Geraniol (%AR) | 204.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | 1 |
| Geraniol (%AR) | 204.02 | GC,MS | 211 | <0.0050 | <0.0050 | <0.0050 | 1 |
| Geraniol (%AR) | 204.02 | GC,MS | 183 | 0 | 0 | 0 | 1 |

| alpha-Humulene (%AR) | 205.02 | GC,MS | 211 | 0.042 | 0.043 | 0.042 | 1.36 | 0.21 | 4.64 | 1.36 | 9.74 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 183 | 0.3019 | 0.2924 | 0.2947 | 1.67 | 0.35 | 4.64 | 1.36 | 9.74 | 0 |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 162 | 0.0274 | 0.0285 | 0.0288 | 2.61 | 0.38 | 4.64 | 1.36 | 9.74 | 0 |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 194 | 0.0367 | 0.0392 | 0.0392 | 3.76 | 0.58 | 4.64 | 1.36 | 9.74 | 0 |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 196 | 0.024 | 0.0272 | 0.0285 | 8.72 | 1.27 | 4.64 | 1.36 | 9.74 | 0 |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 142 | 0.0348 | 0.0418 | 0.0364 | 9.74 | 1.49 | 4.64 | 1.36 | 9.74 | 0 |

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

d-Limonene (%AR)

RSDr	HorRat(r)

Population of Lab RSDr Rob Mean	min	max

Flag

Linalool (%AR)

| d-Limonene (%AR) | 206.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| d-Limonene (%AR) | 206.02 | GC,MS | 211 | <0.0050 | <0.0050 | <0.0050 |  |  | 1 |
| d-Limonene (%AR) | 206.02 | GC,MS | 183 | 0.0161 | 0.0165 | 0.0167 | 1.86 | 0.25 | 0 |
| d-Limonene (%AR) | 206.02 | GC,MS | 194 | 0.0015 | 0.0015 | 0.0016 | 3.77 | 0.36 | 0 |
| d-Limonene (%AR) | 206.02 | GC,MS | 196 | 0.0018 | 0.0016 | 0.0016 | 6.93 | 0.67 | 0 |
| d-Limonene (%AR) | 206.02 | GC,MS | 142 | 0.0016 | 0.0016 | 0.0018 | 6.93 | 0.67 | 0 |

| Linalool (%AR) | 207.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Linalool (%AR) | 207.02 | GC,MS | 211 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| Linalool (%AR) | 207.02 | GC,MS | 142 | 0.0021 | 0.002 | 0.0021 | 2.79 | 0.28 | 0 |
| Linalool (%AR) | 207.02 | GC,MS | 194 | 0.0019 | 0.002 | 0.0021 | 5 | 0.49 | 0 |
| Linalool (%AR) | 207.02 | GC,MS | 183 | 0.0129 | 0.0132 | 0.0144 | 5.88 | 0.77 | 0 |
| Linalool (%AR) | 207.02 | GC,MS | 196 | 0.0017 | 0.0019 | 0.002 | 8.18 | 0.8 | 0 |

|  |  |  | beta-Myrcene (%AR) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 211 | 0.005 | 0.005 | 0.005 | 0 | 0 | 0 |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 196 | 0.006 | 0.0062 | 0.0061 | 1.64 | 0.19 | 0 |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 194 | 0.0064 | 0.0067 | 0.0066 | 2.33 | 0.27 | 0 |

Lab

Nopinene (%AR)

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) Rob Mean min max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 183 | 0.0593 | 0.0616 | 0.0625 | 2.7 | 0.44 | 0 |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 142 | 0.0047 | 0.0046 | 0.0054 | 8.9 | 1 | 0 |

Lab

Population of Lab RSDr

alpha-Pinene (%AR)

Ocimene (%AR)

| Nopinene (%AR) | 209.02 | GC,MS | 211 | <0.0050 <0.0050 <0.0050 |  |  | 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nopinene (%AR) | 209.02 | GC,MS | 183 | 0.0051 | 0.0049 | 0.005 | 2 | 0.23 | 0 |

| Ocimene (%AR) | 210.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Ocimene (%AR) | 210.02 | GC,MS | 183 | 0 | 0 | 0 |  |  | 1 |
| Ocimene (%AR) | 210.04 | Other | 142 | <0.0001 | <0.0001 | <0.0001 |  |  | 1 |
| Ocimene (%AR) | 210.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 |  |  | 1 |
| Ocimene (%AR) | 210.02 | GC,MS | 211 | <0.0050 | <0.0050 | <0.0050 |  |  | 1 |
| Ocimene (%AR) | 210.02 | GC,MS | 194 | 0.0005 | 0.0005 | 0.0005 | 0 | 0 | 0 |

| alpha-Pinene (%AR) | 211.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 211 | <0.0050 | <0.0050 | <0.0050 |  |  | 1 |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 196 | <0.0013 | 0.0015 | 0.0014 |  |  | 1 |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 194 | 0.0015 | 0.0015 | 0.0015 | 0 | 0 | 0 |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 183 | 0.0151 | 0.0153 | 0.0157 | 1.99 | 0.27 | 0 |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 142 | 0.0017 | 0.0016 | 0.002 | 11.8 | 1.14 | 0 |

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

(-)-beta-Pinene (%AR)

RSDr	HorRat(r)

Population of Lab RSDr Rob Mean	min	max

Flag

gamma-Terpinene (%AR)

alpha-Terpinene (%AR)

| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 |  |  | 1 |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 194 | 0.0006 | 0.0006 | 0.0006 | 0 | 0 | 0 |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 142 | 0.0006 | 0.0006 | 0.0007 | 9.12 | 0.76 | 0 |

| alpha-Terpinene (%AR) | 213.02 | GC,MS | 211 | <0.0050 | <0.0050 | <0.0050 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 |  |  | 1 |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 |  |  | 1 |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 142 | <0.0001 | <0.0001 | 0.0001 |  |  | 1 |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 183 | 0.0024 | 0.0024 | 0.0026 | 4.68 | 0.48 | 0 |

| gamma-Terpinene (%A | 214.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| gamma-Terpinene (%A | 214.02 | GC,MS | 211 | <0.0050 | <0.0050 | <0.0050 |  |  | 1 |
| gamma-Terpinene (%A | 214.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 |  |  | 1 |
| gamma-Terpinene (%A | 214.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 |  |  | 1 |
| gamma-Terpinene (%A | 214.02 | GC,MS | 142 | 0.0001 | 0.0001 | 0.0001 | 0 | 0 | 0 |
| gamma-Terpinene (%A | 214.02 | GC,MS | 183 | 0.0019 | 0.002 | 0.002 | 2.94 | 0.29 | 0 |

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

alpha-Terpinolene (%AR)

RSDr	HorRat(r)

Population of Lab RSDr Rob Mean	min	max

Flag

Moisture (%AR)

| alpha-Terpinolene (%A | 215.02 | GC,MS | 211 | <0.0050 | <0.0050 | <0.0050 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 |  |  | 1 |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 |  |  | 1 |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 142 | 0.0001 | 0.0001 | 0.0001 | 0 | 0 | 0 |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 183 | 0.0015 | 0.0016 | 0.0016 | 3.69 | 0.35 | 0 |

| Moisture (%AR) | 500.70 | Commerc. Anal. | 172 | 10.263 | 10.263 | 10.263 | 0 | 0 | 3.7 | 0 | 14.3 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 204 | 5.57 | 5.57 | 5.57 | 0 | 0 | 3.7 | 0 | 14.3 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 199 | 9.0385 | 9.0217 | 9.0407 | 0.12 | 0.04 | 3.7 | 0 | 14.3 | 0 |
| Moisture (%AR) | 500.99 | Other | 165 | 8.4984 | 8.426 | 8.4514 | 0.43 | 0.15 | 3.7 | 0 | 14.3 | 0 |
| Moisture (%AR) | 500.99 | Other | 103 | 6.63 | 6.61 | 6.54 | 0.72 | 0.24 | 3.7 | 0 | 14.3 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 192 | 10.2679 | 10.2921 | 10.4276 | 0.83 | 0.3 | 3.7 | 0 | 14.3 | 0 |
| Moisture (%AR) | 500.99 | Other | 105 | 9.94 | 9.74 | 9.77 | 1.1 | 0.39 | 3.7 | 0 | 14.3 | 0 |
| Moisture (%AR) | 500.72 | NIR | 187 | 6.05 | 6.07 | 6.22 | 1.52 | 0.5 | 3.7 | 0 | 14.3 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 183 | 8.2 | 7.97 | 7.92 | 1.86 | 0.64 | 3.7 | 0 | 14.3 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 116 | 6.08 | 5.85 | 5.91 | 2.01 | 0.66 | 3.7 | 0 | 14.3 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 106 | 7.39 | 7.74 | 7.71 | 2.55 | 0.87 | 3.7 | 0 | 14.3 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 212 | 6.01 | 6.36 | 6.11 | 2.93 | 0.96 | 3.7 | 0 | 14.3 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Moisture (%AR) | 500.99 | Other | 142 | 5.42 | 5.74 | 5.71 | 3.14 | 1.02 | 3.7 | 0 | 14.3 | 0 |
| Moisture (%AR) | 500.99 | Other | 120 | 7.2891 | 7.7758 | 7.4614 | 3.29 | 1.11 | 3.7 | 0 | 14.3 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 182 | 7.61 | 8.15 | 8.04 | 3.6 | 1.23 | 3.7 | 0 | 14.3 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 211 | 9.72 | 9.13 | 8.93 | 4.44 | 1.55 | 3.7 | 0 | 14.3 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 196 | 6.99 | 7.13 | 7.65 | 4.79 | 1.62 | 3.7 | 0 | 14.3 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 208 | 10.18 | 10.41 | 9.26 | 6.12 | 2.16 | 3.7 | 0 | 14.3 | 0 |
| Moisture (%AR) | 500.99 | Other | 132 | 4.82 | 4.97 | 5.55 | 7.54 | 2.41 | 3.7 | 0 | 14.3 | 0 |
| Moisture (%AR) | 500.72 | NIR | 138 | 8.8 | 7.9 | 7.6 | 7.71 | 2.64 | 3.7 | 0 | 14.3 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 209 | 8.7473 | 8.6057 | 7.5594 | 7.81 | 2.69 | 3.7 | 0 | 14.3 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 109 | 5.76 | 6.17 | 5.21 | 8.43 | 2.74 | 3.7 | 0 | 14.3 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 153 | 4.27 | 3.65 | 4.42 | 9.92 | 3.07 | 3.7 | 0 | 14.3 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 210 | 8.72 | 10.79 | 8.32 | 14.3 | 5 | 3.7 | 0 | 14.3 | 0 |
|  |  |  |  | Δ9-THC (%DW) |  |  |  |  |  |  |  |  |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 189 | 0.2723 | 0.2718 | 0.2732 | 0.26 | 0.05 | 3.31 | 0.26 | 8.93 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 204 | 0.2301 | 0.2284 | 0.2295 | 0.38 | 0.08 | 3.31 | 0.26 | 8.93 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 211 | 0.246 | 0.244 | 0.241 | 1.03 | 0.21 | 3.31 | 0.26 | 8.93 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 142 | 0.2301 | 0.224 | 0.2226 | 1.77 | 0.35 | 3.31 | 0.26 | 8.93 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 196 | 0.2539 | 0.2558 | 0.2632 | 1.91 | 0.39 | 3.31 | 0.26 | 8.93 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 118 | 0.212 | 0.209 | 0.203 | 2.2 | 0.44 | 3.31 | 0.26 | 8.93 | 0 |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, | 210 | 0.257 | 0.253 | 0.246 | 2.21 | 0.45 | 3.31 | 0.26 | 8.93 | 0 |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, | 218 | 0.23 | 0.224 | 0.22 | 2.24 | 0.45 | 3.31 | 0.26 | 8.93 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, | 199 | 0.2391 | 0.2358 | 0.2468 | 2.35 | 0.47 | 3.31 | 0.26 | 8.93 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 215 | 0.2353 | 0.2357 | 0.2259 | 2.39 | 0.48 | 3.31 | 0.26 | 8.93 | 0 |
| Δ9-THC (%DW) | 501.60 | NIR | 187 | 0.21 | 0.21 | 0.2 | 2.79 | 0.55 | 3.31 | 0.26 | 8.93 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 109 | 0.237 | 0.235 | 0.225 | 2.77 | 0.56 | 3.31 | 0.26 | 8.93 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 165 | 0.2433 | 0.2324 | 0.2475 | 3.23 | 0.65 | 3.31 | 0.26 | 8.93 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 116 | 0.2592 | 0.2695 | 0.2527 | 3.25 | 0.67 | 3.31 | 0.26 | 8.93 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 172 | 0.2964 | 0.302 | 0.2808 | 3.75 | 0.78 | 3.31 | 0.26 | 8.93 | 0 |
| Δ9-THC (%DW) | 501.30 | Other LC-MS | 102 | 0.2286 | 0.2109 | 0.2256 | 4.27 | 0.85 | 3.31 | 0.26 | 8.93 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 209 | 0.2418 | 0.2611 | 0.2599 | 4.25 | 0.87 | 3.31 | 0.26 | 8.93 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 154 | 0.2025 | 0.2114 | 0.1909 | 5.1 | 1 | 3.31 | 0.26 | 8.93 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 208 | 0.2204 | 0.2444 | 0.2497 | 6.56 | 1.32 | 3.31 | 0.26 | 8.93 | 0 |
| Δ9-THC (%DW) | 501.30 | Other LC-MS | 120 | 0.288 | 0.2508 | 0.2634 | 7.08 | 1.45 | 3.31 | 0.26 | 8.93 | 0 |
| Δ9-THC (%DW) | 501.30 | Other LC-MS | 105 | 0.211 | 0.248 | 0.239 | 8.29 | 1.67 | 3.31 | 0.26 | 8.93 | 0 |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, | 103 | 0.2196 | 0.2184 | 0.2547 | 8.93 | 1.8 | 3.31 | 0.26 | 8.93 | 0 |
|  |  |  |  | Δ9-THCA (%DW) |  |  |  |  |  |  |  |  |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 209 | 0.1244 | 0.1247 | 0.1243 | 0.17 | 0.03 | 3.07 | 0.17 | 13.1 | 0 |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, | 103 | 0.1264 | 0.1274 | 0.1284 | 0.78 | 0.14 | 3.07 | 0.17 | 13.1 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 189 | 0.1296 | 0.1315 | 0.1313 | 0.8 | 0.15 | 3.07 | 0.17 | 13.1 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 116 | 0.1156 | 0.1174 | 0.1172 | 0.84 | 0.15 | 3.07 | 0.17 | 13.1 | 0 |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, | 199 | 0.1215 | 0.1198 | 0.1198 | 0.82 | 0.15 | 3.07 | 0.17 | 13.1 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 204 | 0.1287 | 0.1257 | 0.1278 | 1.21 | 0.22 | 3.07 | 0.17 | 13.1 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, | 218 | 0.093 | 0.089 | 0.091 | 2.2 | 0.38 | 3.07 | 0.17 | 13.1 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 211 | 0.109 | 0.107 | 0.104 | 2.36 | 0.42 | 3.07 | 0.17 | 13.1 | 0 |
| Δ9-THCA (%DW) | 502.30 | Other LC-MS | 120 | 0.1228 | 0.1183 | 0.117 | 2.55 | 0.46 | 3.07 | 0.17 | 13.1 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 196 | 0.1131 | 0.1156 | 0.1088 | 3.06 | 0.55 | 3.07 | 0.17 | 13.1 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 109 | 0.128 | 0.129 | 0.122 | 3 | 0.55 | 3.07 | 0.17 | 13.1 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 118 | 0.11 | 0.11 | 0.104 | 3.21 | 0.58 | 3.07 | 0.17 | 13.1 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 142 | 0.1071 | 0.1076 | 0.1008 | 3.6 | 0.64 | 3.07 | 0.17 | 13.1 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 208 | 0.109 | 0.1172 | 0.114 | 3.64 | 0.66 | 3.07 | 0.17 | 13.1 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 172 | 0.1516 | 0.1605 | 0.1493 | 3.85 | 0.73 | 3.07 | 0.17 | 13.1 | 0 |
| Δ9-THCA (%DW) | 502.60 | NIR | 187 | 0.14 | 0.15 | 0.14 | 4.03 | 0.75 | 3.07 | 0.17 | 13.1 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 154 | 0.1056 | 0.1076 | 0.096 | 6.02 | 1.07 | 3.07 | 0.17 | 13.1 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 165 | 0.1351 | 0.1253 | 0.1418 | 6.19 | 1.15 | 3.07 | 0.17 | 13.1 | 0 |
| Δ9-THCA (%DW) | 502.30 | Other LC-MS | 105 | 0.0893 | 0.096 | 0.105 | 8.14 | 1.44 | 3.07 | 0.17 | 13.1 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 215 | 0.105 | 0.1361 | 0.1176 | 13.1 | 2.38 | 3.07 | 0.17 | 13.1 | 0 |
|  |  |  |  | CBD (%DW) |  |  |  |  |  |  |  |  |
| CBD (%DW) | 503.10 | Other LC-UV | 189 | 3.5362 | 3.5409 | 3.54 | 0.07 | 0.02 | 2.31 | 0.07 | 5.33 | 0 |
| CBD (%DW) | 503.02 | AOAC 2018.11, | 210 | 2.922 | 2.922 | 2.938 | 0.32 | 0.09 | 2.31 | 0.07 | 5.33 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 142 | 3.696 | 3.6593 | 3.6256 | 0.96 | 0.29 | 2.31 | 0.07 | 5.33 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 204 | 3.6347 | 3.5507 | 3.6057 | 1.19 | 0.36 | 2.31 | 0.07 | 5.33 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 118 | 3.12 | 3.118 | 3.044 | 1.4 | 0.42 | 2.31 | 0.07 | 5.33 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 196 | 3.9707 | 4.033 | 4.0801 | 1.36 | 0.42 | 2.31 | 0.07 | 5.33 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%DW) | 503.60 | NIR | 187 | 3.17 | 3.25 | 3.17 | 1.44 | 0.43 | 2.31 | 0.07 | 5.33 | 0 |
| CBD (%DW) | 503.02 | AOAC 2018.11, | 218 | 3.19 | 3.098 | 3.208 | 1.86 | 0.56 | 2.31 | 0.07 | 5.33 | 0 |
| CBD (%DW) | 503.02 | AOAC 2018.11, | 103 | 3.5482 | 3.525 | 3.654 | 1.92 | 0.58 | 2.31 | 0.07 | 5.33 | 0 |
| CBD (%DW) | 503.02 | AOAC 2018.11, | 199 | 3.8962 | 3.8295 | 3.9897 | 2.06 | 0.63 | 2.31 | 0.07 | 5.33 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 109 | 4.03 | 3.912 | 3.87 | 2.11 | 0.65 | 2.31 | 0.07 | 5.33 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 172 | 4.5622 | 4.6636 | 4.4084 | 2.83 | 0.89 | 2.31 | 0.07 | 5.33 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 211 | 3.873 | 3.86 | 3.662 | 3.11 | 0.95 | 2.31 | 0.07 | 5.33 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 116 | 3.8117 | 3.9263 | 3.6733 | 3.33 | 1.02 | 2.31 | 0.07 | 5.33 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 208 | 3.7207 | 3.9694 | 3.7817 | 3.39 | 1.04 | 2.31 | 0.07 | 5.33 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 215 | 3.4212 | 3.3249 | 3.5632 | 3.49 | 1.05 | 2.31 | 0.07 | 5.33 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 209 | 2.1295 | 2.2121 | 2.2996 | 3.84 | 1.08 | 2.31 | 0.07 | 5.33 | 0 |
| CBD (%DW) | 503.30 | Other LC-MS | 105 | 3.32 | 3.67 | 3.55 | 5.06 | 1.53 | 2.31 | 0.07 | 5.33 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 154 | 3.0163 | 3.2805 | 2.9775 | 5.33 | 1.58 | 2.31 | 0.07 | 5.33 | 0 |
|  |  |  |  | CBDA (%DW) |  |  |  |  |  |  |  |  |
| CBDA (%DW) | 504.10 | Other LC-UV | 189 | 6.1169 | 6.1215 | 6.1199 | 0.04 | 0.01 | 3.49 | 0.04 | 8.23 | 0 |
| CBDA (%DW) | 504.02 | AOAC 2018.11, | 210 | 7.167 | 7.168 | 7.204 | 0.29 | 0.1 | 3.49 | 0.04 | 8.23 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 204 | 6.8937 | 6.8022 | 6.8926 | 0.76 | 0.26 | 3.49 | 0.04 | 8.23 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 196 | 7.2749 | 7.405 | 7.3537 | 0.89 | 0.3 | 3.49 | 0.04 | 8.23 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 109 | 7.757 | 7.634 | 7.496 | 1.71 | 0.58 | 3.49 | 0.04 | 8.23 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 142 | 6.7807 | 6.6969 | 6.5146 | 2.04 | 0.68 | 3.49 | 0.04 | 8.23 | 0 |
| CBDA (%DW) | 504.02 | AOAC 2018.11, | 218 | 6.283 | 6.005 | 6.244 | 2.44 | 0.8 | 3.49 | 0.04 | 8.23 | 0 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDA (%DW) | 504.10 | Other LC-UV | 208 | 7.1519 | 7.5015 | 7.3859 | 2.42 | 0.82 | 3.49 | 0.04 | 8.23 | 0 |
| CBDA (%DW) | 504.02 | AOAC 2018.11, | 199 | 6.9463 | 7.2061 | 7.4011 | 3.18 | 1.07 | 3.49 | 0.04 | 8.23 | 0 |
| CBDA (%DW) | 504.60 | NIR | 187 | 5.86 | 6.33 | 6.14 | 3.87 | 1.27 | 3.49 | 0.04 | 8.23 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 118 | 5.856 | 6.048 | 5.56 | 4.22 | 1.38 | 3.49 | 0.04 | 8.23 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 172 | 8.3778 | 8.6676 | 8.0012 | 4 | 1.38 | 3.49 | 0.04 | 8.23 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 211 | 6.908 | 6.89 | 6.384 | 4.42 | 1.47 | 3.49 | 0.04 | 8.23 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 116 | 6.2401 | 6.5424 | 5.9352 | 4.87 | 1.6 | 3.49 | 0.04 | 8.23 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 154 | 5.7063 | 6.1486 | 5.4586 | 6.06 | 1.97 | 3.49 | 0.04 | 8.23 | 0 |
| CBDA (%DW) | 504.30 | Other LC-MS | 105 | 6.8 | 6.3 | 7.25 | 7.01 | 2.34 | 3.49 | 0.04 | 8.23 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 215 | 6.624 | 7.3284 | 7.6227 | 7.14 | 2.4 | 3.49 | 0.04 | 8.23 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 209 | 4.448 | 5.2441 | 4.938 | 8.23 | 2.62 | 3.49 | 0.04 | 8.23 | 0 |

Lab

Population of Lab RSDr

CBN (%DW)

| CBN (%DW) | 505.10 | Other LC-UV | 116 | <0.0400 | <0.0400 | <0.0400 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%DW) | 505.02 | AOAC 2018.11, | 210 | <0.0140 | <0.0140 | <0.0140 |  |  |  |  |  | 1 |
| CBN (%DW) | 505.10 | Other LC-UV | 196 | <0.0205 | <0.0205 | <0.0205 |  |  |  |  |  | 1 |
| CBN (%DW) | 505.10 | Other LC-UV | 208 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| CBN (%DW) | 505.02 | AOAC 2018.11, | 199 | <0.0250 | <0.0250 | <0.0250 |  |  |  |  |  | 1 |
| CBN (%DW) | 505.10 | Other LC-UV | 211 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| CBN (%DW) | 505.30 | GC-FID | 116 | <0.0400 | <0.0400 | <0.0400 |  |  |  |  |  | 1 |
| CBN (%DW) | 505.20 | Other LC-MS | 105 | <0.0144 | <0.0144 | <0.0144 |  |  |  |  |  | 1 |
| CBN (%DW) | 505.10 | Other LC-UV | 189 | 0.0185 | 0.0184 | 0.0186 | 0.54 | 0.07 | 6.6 | 0.54 | 15.5 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%DW) | 505.10 | Other LC-UV | 204 | 0.0168 | 0.0166 | 0.0167 | 0.6 | 0.08 | 6.6 | 0.54 | 15.5 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 142 | 0.0143 | 0.0141 | 0.0137 | 2.18 | 0.29 | 6.6 | 0.54 | 15.5 | 0 |
| CBN (%DW) | 505.02 | AOAC 2018.11, | 218 | 0.018 | 0.017 | 0.018 | 3.27 | 0.45 | 6.6 | 0.54 | 15.5 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 154 | 0.0193 | 0.0192 | 0.017 | 7.03 | 0.97 | 6.6 | 0.54 | 15.5 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 109 | 0.016 | 0.015 | 0.018 | 9.35 | 1.26 | 6.6 | 0.54 | 15.5 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 209 | 0.017 | 0.0129 | 0.0141 | 14.4 | 1.91 | 6.6 | 0.54 | 15.5 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 192 | 0.0456 | 0.045 | 0.0586 | 15.5 | 2.47 | 6.6 | 0.54 | 15.5 | 0 |
|  |  |  |  | Total Δ9-THC (%DW) |  |  |  |  |  |  |  |  |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 189 | 0.386 | 0.3871 | 0.3884 | 0.31 | 0.07 | 3.36 | 0.31 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 204 | 0.343 | 0.3387 | 0.3416 | 0.64 | 0.14 | 3.36 | 0.31 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 196 | 0.3531 | 0.3571 | 0.3586 | 0.8 | 0.17 | 3.36 | 0.31 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 116 | 0.3304 | 0.3253 | 0.3237 | 1.07 | 0.23 | 3.36 | 0.31 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 211 | 0.341 | 0.337 | 0.332 | 1.34 | 0.29 | 3.36 | 0.31 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, | 199 | 0.3457 | 0.3409 | 0.3519 | 1.59 | 0.34 | 3.36 | 0.31 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, | 218 | 0.311 | 0.302 | 0.3 | 1.93 | 0.4 | 3.36 | 0.31 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 142 | 0.324 | 0.3183 | 0.311 | 2.05 | 0.43 | 3.36 | 0.31 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 116 | 0.3606 | 0.3725 | 0.3556 | 2.39 | 0.51 | 3.36 | 0.31 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 118 | 0.308 | 0.305 | 0.294 | 2.44 | 0.51 | 3.36 | 0.31 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 192 | 0.3656 | 0.3611 | 0.3806 | 2.77 | 0.6 | 3.36 | 0.31 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 109 | 0.349 | 0.348 | 0.331 | 2.95 | 0.63 | 3.36 | 0.31 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.60 | NIR | 187 | 0.33 | 0.34 | 0.32 | 3.03 | 0.64 | 3.36 | 0.31 | 10.8 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 209 | 0.3509 | 0.3704 | 0.369 | 2.99 | 0.64 | 3.36 | 0.31 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 182 | 0.3206 | 0.3145 | 0.3013 | 3.16 | 0.67 | 3.36 | 0.31 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 172 | 0.4293 | 0.4427 | 0.4118 | 3.62 | 0.8 | 3.36 | 0.31 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 165 | 0.3617 | 0.3422 | 0.3718 | 4.2 | 0.9 | 3.36 | 0.31 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 215 | 0.3274 | 0.355 | 0.329 | 4.6 | 0.98 | 3.36 | 0.31 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 154 | 0.2951 | 0.3057 | 0.2751 | 5.32 | 1.11 | 3.36 | 0.31 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 208 | 0.316 | 0.3471 | 0.3497 | 5.55 | 1.18 | 3.36 | 0.31 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 153 | 0.2989 | 0.2949 | 0.3273 | 5.75 | 1.21 | 3.36 | 0.31 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.30 | Other LC-MS | 120 | 0.3957 | 0.3545 | 0.366 | 5.71 | 1.23 | 3.36 | 0.31 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, | 103 | 0.3304 | 0.3302 | 0.3672 | 6.22 | 1.33 | 3.36 | 0.31 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.30 | Other LC-MS | 105 | 0.289 | 0.332 | 0.331 | 7.73 | 1.63 | 3.36 | 0.31 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 138 | 0.3529 | 0.2941 | 0.3603 | 10.8 | 2.3 | 3.36 | 0.31 | 10.8 | 0 |
|  |  |  |  | Total CBD (%DW) |  |  |  |  |  |  |  |  |
| Total CBD (%DW) | 507.10 | Other LC-UV | 189 | 8.9008 | 8.9094 | 8.9072 | 0.05 | 0.02 | 3.1 | 0.05 | 7.44 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 204 | 9.6804 | 9.5162 | 9.6506 | 0.91 | 0.32 | 3.1 | 0.05 | 7.44 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 192 | 8.8952 | 8.9271 | 9.0608 | 0.98 | 0.34 | 3.1 | 0.05 | 7.44 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 196 | 10.3507 | 10.5272 | 10.5293 | 0.98 | 0.35 | 3.1 | 0.05 | 7.44 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 142 | 9.6427 | 9.5325 | 9.3389 | 1.62 | 0.57 | 3.1 | 0.05 | 7.44 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 109 | 10.833 | 10.607 | 10.444 | 1.84 | 0.66 | 3.1 | 0.05 | 7.44 | 0 |
| Total CBD (%DW) | 507.02 | AOAC 2018.11, | 218 | 8.701 | 8.364 | 8.684 | 2.21 | 0.77 | 3.1 | 0.05 | 7.44 | 0 |
| Total CBD (%DW) | 507.02 | AOAC 2018.11, | 199 | 9.9881 | 10.1492 | 10.4805 | 2.46 | 0.87 | 3.1 | 0.05 | 7.44 | 0 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%DW) | 507.10 | Other LC-UV | 208 | 9.9929 | 10.5482 | 10.2591 | 2.71 | 0.96 | 3.1 | 0.05 | 7.44 | 0 |
| Total CBD (%DW) | 507.60 | NIR | 187 | 8.31 | 8.81 | 8.55 | 2.92 | 1.01 | 3.1 | 0.05 | 7.44 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 118 | 8.256 | 8.422 | 7.921 | 3.11 | 1.07 | 3.1 | 0.05 | 7.44 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 172 | 11.9096 | 12.2651 | 11.4255 | 3.55 | 1.29 | 3.1 | 0.05 | 7.44 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 211 | 9.931 | 9.903 | 9.261 | 3.91 | 1.38 | 3.1 | 0.05 | 7.44 | 0 |
| Total CBD (%DW) | 507.30 | Other LC-MS | 105 | 9.28 | 9.2 | 9.91 | 4.11 | 1.44 | 3.1 | 0.05 | 7.44 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 116 | 9.2843 | 9.664 | 8.8785 | 4.24 | 1.48 | 3.1 | 0.05 | 7.44 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 215 | 9.2304 | 9.7519 | 10.2484 | 5.22 | 1.84 | 3.1 | 0.05 | 7.44 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 154 | 8.0208 | 8.6729 | 7.7647 | 5.74 | 1.97 | 3.1 | 0.05 | 7.44 | 0 |
| Total CBD (%DW) | 507.40 | GC-FID | 116 | 8.9932 | 8.33 | 9.4183 | 6.15 | 2.14 | 3.1 | 0.05 | 7.44 | 0 |
| Total CBD (%DW) | 507.40 | GC-FID | 138 | 9.0585 | 10.4835 | 10.0824 | 7.44 | 2.63 | 3.1 | 0.05 | 7.44 | 0 |

Lab

Population of Lab RSDr

CBDV (%DW)

| CBDV (%DW) | 508.99 | Other | 211 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDV (%DW) | 508.99 | Other | 109 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| CBDV (%DW) | 508.99 | Other | 142 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| CBDV (%DW) | 508.99 | Other | 189 | 0.0215 | 0.0213 | 0.0217 | 0.93 | 0.13 | 0 |
| CBDV (%DW) | 508.02 | AOAC 2018.11, | 210 | 0.027 | 0.025 | 0.026 | 3.85 | 0.56 | 0 |
| CBDV (%DW) | 508.99 | Other | 154 | 0.0411 | 0.043 | 0.0399 | 3.78 | 0.59 | 0 |
| CBDV (%DW) | 508.99 | Other | 196 | 0.0259 | 0.0265 | 0.0289 | 5.86 | 0.85 | 0 |
| CBDV (%DW) | 508.99 | Other | 208 | 0.082 | 0.104 | 0.0919 | 11.9 | 2.09 | 0 |

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

CBDVA (%DW)

RSDr	HorRat(r)

Population of Lab RSDr Rob Mean	min	max

Flag

CBGA (%DW)

CBG (%DW)

| CBDVA (%DW) | 509.99 | Other | 196 | 0.0697 | 0.0718 | 0.0706 | 1.49 | 0.25 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDVA (%DW) | 509.02 | AOAC 2018.11, | 210 | 0.072 | 0.07 | 0.07 | 1.63 | 0.28 | 0 |
| CBDVA (%DW) | 509.99 | Other | 154 | 0.0646 | 0.0674 | 0.0601 | 5.75 | 0.95 | 0 |
| CBDVA (%DW) | 509.99 | Other | 208 | 0.1223 | 0.2053 | 0.1246 | 31.4 | 5.92 | 0 |

| CBG (%DW) | 510.99 | Other | 189 | 0.1633 | 0.1629 | 0.1632 | 0.13 | 0.02 | 5.82 | 0.13 | 40.8 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%DW) | 510.99 | Other | 142 | 0.1339 | 0.1313 | 0.1302 | 1.44 | 0.27 | 5.82 | 0.13 | 40.8 | 0 |
| CBG (%DW) | 510.02 | AOAC 2018.11, | 210 | 0.125 | 0.127 | 0.13 | 1.98 | 0.36 | 5.82 | 0.13 | 40.8 | 0 |
| CBG (%DW) | 510.99 | Other | 196 | 0.122 | 0.1245 | 0.1274 | 2.17 | 0.4 | 5.82 | 0.13 | 40.8 | 0 |
| CBG (%DW) | 510.99 | Other | 208 | 0.1415 | 0.1438 | 0.1377 | 2.18 | 0.41 | 5.82 | 0.13 | 40.8 | 0 |
| CBG (%DW) | 510.99 | Other | 204 | 0.1587 | 0.152 | 0.1592 | 2.57 | 0.49 | 5.82 | 0.13 | 40.8 | 0 |
| CBG (%DW) | 510.99 | Other | 109 | 0.13 | 0.124 | 0.135 | 4.25 | 0.78 | 5.82 | 0.13 | 40.8 | 0 |
| CBG (%DW) | 510.99 | Other | 105 | 0.125 | 0.104 | 0.136 | 13.4 | 2.44 | 5.82 | 0.13 | 40.8 | 0 |
| CBG (%DW) | 510.99 | Other | 211 | 0.139 | 0.139 | 0.108 | 13.9 | 2.56 | 5.82 | 0.13 | 40.8 | 0 |
| CBG (%DW) | 510.99 | Other | 154 | 0.0658 | 0.1592 | 0.1443 | 40.8 | 7.46 | 5.82 | 0.13 | 40.8 | 0 |

| CBGA (%DW) | 511.99 | Other | 189 | 0.3206 | 0.3195 | 0.3203 | 0.18 | 0.04 | 3.08 | 0.18 | 14.6 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%DW) | 511.99 | Other | 196 | 0.3193 | 0.3257 | 0.3283 | 1.43 | 0.3 | 3.08 | 0.18 | 14.6 | 0 |
| CBGA (%DW) | 511.02 | AOAC 2018.11, | 210 | 0.28 | 0.271 | 0.277 | 1.66 | 0.34 | 3.08 | 0.18 | 14.6 | 0 |

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%DW) | 511.99 | Other | 204 | 0.321 | 0.3139 | 0.3243 | 1.66 | 0.35 | 3.08 | 0.18 | 14.6 | 0 |
| CBGA (%DW) | 511.99 | Other | 142 | 0.3249 | 0.3257 | 0.3113 | 2.52 | 0.53 | 3.08 | 0.18 | 14.6 | 0 |
| CBGA (%DW) | 511.99 | Other | 208 | 0.3219 | 0.3354 | 0.3409 | 2.94 | 0.62 | 3.08 | 0.18 | 14.6 | 0 |
| CBGA (%DW) | 511.99 | Other | 109 | 0.387 | 0.365 | 0.357 | 4.2 | 0.91 | 3.08 | 0.18 | 14.6 | 0 |
| CBGA (%DW) | 511.99 | Other | 154 | 0.239 | 0.2563 | 0.2266 | 6.2 | 1.25 | 3.08 | 0.18 | 14.6 | 0 |
| CBGA (%DW) | 511.99 | Other | 211 | 0.312 | 0.314 | 0.24 | 14.6 | 3.04 | 3.08 | 0.18 | 14.6 | 0 |

Lab

Population of Lab RSDr

Δ8-THC (%DW)

THCV (%DW)

| THCV (%DW) | 512.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| THCV (%DW) | 512.99 | Other | 208 | <0.0500 | <0.0500 | <0.0500 | 1 |
| THCV (%DW) | 512.99 | Other | 211 | <0.0100 | <0.0100 | <0.0100 | 1 |
| THCV (%DW) | 512.99 | Other | 189 | <0.0001 | <0.0001 | <0.0001 | 1 |
| THCV (%DW) | 512.02 | AOAC 2018.11, | 210 | <0.0170 | <0.0170 | <0.0170 | 1 |
| THCV (%DW) | 512.99 | Other | 142 | <0.0100 | <0.0100 | <0.0100 | 1 |
| THCV (%DW) | 512.99 | Other | 109 | <0.0100 | <0.0100 | <0.0100 | 1 |

| Δ8-THC (%DW) | 513.99 | Other | 109 | <0.0100 | <0.0100 | <0.0100 | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%DW) | 513.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 | 1 |
| Δ8-THC (%DW) | 513.99 | Other | 211 | <0.0100 | <0.0100 | <0.0100 | 1 |
| Δ8-THC (%DW) | 513.99 | Other | 208 | <0.0500 | <0.0500 | <0.0500 | 1 |
| Δ8-THC (%DW) | 513.02 | AOAC 2018.11, | 210 | <0.0140 | <0.0140 | <0.0140 | 1 |

Lab

CBC (%DW)

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) Rob Mean min max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%DW) | 513.02 | AOAC 2018.11, | 118 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| Δ8-THC (%DW) | 513.99 | Other | 142 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| Δ8-THC (%DW) | 513.99 | Other | 189 | 0.0105 | 0.0098 | 0.0092 | 6.62 | 0.83 | 0 |

Lab

Population of Lab RSDr

| CBC (%DW) | 514.99 | Other | 189 | 0.2732 | 0.2741 | 0.2742 | 0.2 | 0.04 | 3.39 | 0.2 | 7.07 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%DW) | 514.99 | Other | 196 | 0.2972 | 0.299 | 0.2955 | 0.59 | 0.12 | 3.39 | 0.2 | 7.07 | 0 |
| CBC (%DW) | 514.99 | Other | 204 | 0.2815 | 0.2766 | 0.2781 | 0.9 | 0.19 | 3.39 | 0.2 | 7.07 | 0 |
| CBC (%DW) | 514.99 | Other | 142 | 0.2586 | 0.2528 | 0.2493 | 1.85 | 0.38 | 3.39 | 0.2 | 7.07 | 0 |
| CBC (%DW) | 514.02 | AOAC 2018.11, | 103 | 0.242 | 0.2441 | 0.2525 | 2.26 | 0.46 | 3.39 | 0.2 | 7.07 | 0 |
| CBC (%DW) | 514.99 | Other | 109 | 0.302 | 0.292 | 0.284 | 3.08 | 0.64 | 3.39 | 0.2 | 7.07 | 0 |
| CBC (%DW) | 514.99 | Other | 208 | 0.2467 | 0.2725 | 0.2531 | 5.22 | 1.07 | 3.39 | 0.2 | 7.07 | 0 |
| CBC (%DW) | 514.99 | Other | 154 | 0.2252 | 0.2494 | 0.2269 | 5.78 | 1.16 | 3.39 | 0.2 | 7.07 | 0 |
| CBC (%DW) | 514.99 | Other | 105 | 0.28 | 0.245 | 0.255 | 6.93 | 1.42 | 3.39 | 0.2 | 7.07 | 0 |
| CBC (%DW) | 514.99 | Other | 211 | 0.282 | 0.274 | 0.246 | 7.07 | 1.45 | 3.39 | 0.2 | 7.07 | 0 |

- Flag definitions: 1 (Lab RSDr not calculated because there were less than 3 numeric lab results)

- Statistical parameters of robust mean, minimum, and maximum for RSDr only shown only if there were 6 or more lab RSDr values.

- AR refers to concentration on as-received basis. DW refers to concentration on dry weight basis.c