Hemp Proficiency Testing

All Labs Precision HM22NOV-1 (hemp)

Sample: HM22NOV-1 (hemp)

All Labs Precision

Issue Date:

January 20, 2023

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

Δ9-THC (%AR)

RSDr	HorRat(r)

Population of Lab RSDr Rob Mean	min	max

Flag

| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 120 | 0.1978 | 0.1982 | 0.1987 | 0.23 | 0.04 | 3.09 | 0.23 | 10.1 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 189 | 0.2151 | 0.2138 | 0.2151 | 0.35 | 0.07 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 122 | 0.2105 | 0.2124 | 0.2115 | 0.45 | 0.09 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 192 | 0.1205 | 0.1221 | 0.121 | 0.68 | 0.12 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 199 | 0.202 | 0.2025 | 0.2045 | 0.65 | 0.13 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 103 | 0.159 | 0.161 | 0.161 | 0.72 | 0.14 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 204 | 0.193 | 0.19 | 0.1911 | 0.79 | 0.16 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 142 | 0.198 | 0.202 | 0.197 | 1.33 | 0.26 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 197 | 0.1759 | 0.1813 | 0.1782 | 1.52 | 0.29 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 161 | 0.1657 | 0.1679 | 0.1617 | 1.9 | 0.36 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 200 | 0.183 | 0.18 | 0.187 | 1.92 | 0.37 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 175 | 0.2179 | 0.2142 | 0.2095 | 1.97 | 0.39 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 113 | 0.1652 | 0.1648 | 0.1711 | 2.11 | 0.4 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 224 | 0.1866 | 0.1817 | 0.1785 | 2.24 | 0.43 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.99 | Other | 183 | 0.202 | 0.208 | 0.199 | 2.26 | 0.45 | 3.09 | 0.23 | 10.1 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 191 | 0.1922 | 0.1836 | 0.1854 | 2.42 | 0.47 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 164 | 0.1779 | 0.1764 | 0.1696 | 2.53 | 0.49 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 115 | 0.24 | 0.23 | 0.24 | 2.44 | 0.49 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 116 | 0.2622 | 0.2574 | 0.2713 | 2.68 | 0.55 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 144 | 0.205 | 0.2131 | 0.2167 | 2.83 | 0.56 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.99 | Other | 132 | 0.175 | 0.171 | 0.165 | 2.95 | 0.57 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 172 | 0.193 | 0.201 | 0.205 | 3.06 | 0.6 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 106 | 0.19 | 0.19 | 0.18 | 3.09 | 0.6 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 196 | 0.1687 | 0.1794 | 0.1756 | 3.11 | 0.6 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 194 | 0.2046 | 0.2101 | 0.1975 | 3.1 | 0.61 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 149 | 0.177 | 0.189 | 0.182 | 3.3 | 0.64 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 182 | 0.1572 | 0.1533 | 0.147 | 3.37 | 0.64 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 101 | 0.216 | 0.2022 | 0.2076 | 3.33 | 0.66 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 102 | 0.1856 | 0.1977 | 0.2 | 3.98 | 0.78 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 163 | 0.206 | 0.221 | 0.207 | 3.97 | 0.79 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 170 | 0.2124 | 0.2039 | 0.194 | 4.53 | 0.89 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 125 | 0.187 | 0.17 | 0.177 | 4.8 | 0.93 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.03 | AOAC 2018.11, M | 139 | 0.2489 | 0.2288 | 0.2503 | 4.96 | 1 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 148 | 0.117 | 0.114 | 0.104 | 6.1 | 1.1 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 155 | 0.19 | 0.2 | 0.216 | 6.49 | 1.28 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 225 | 0.1843 | 0.1956 | 0.1708 | 6.76 | 1.31 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 184 | 0.1989 | 0.2265 | 0.2199 | 6.7 | 1.33 | 3.09 | 0.23 | 10.1 | 0 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 152 | 0.2235 | 0.2154 | 0.1935 | 7.36 | 1.46 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.60 | NIR | 187 | 0.12 | 0.14 | 0.14 | 8.66 | 1.6 | 3.09 | 0.23 | 10.1 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 217 | 0.1803 | 0.2118 | 0.2187 | 10.1 | 1.98 | 3.09 | 0.23 | 10.1 | 0 |

Lab

Population of Lab RSDr

Δ9-THCA (%AR)

| Δ9-THCA (%AR) | 002.60 | NIR | 187 | 0.01 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.99 | Other | 132 | <0.1000 | <0.1000 | <0.1000 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 197 | <0.0694 | <0.0739 | <0.0711 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 184 | 0.0722 | 0.0725 | 0.0723 | 0.21 | 0.04 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 115 | 0.07 | 0.07 | 0.071 | 0.82 | 0.14 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 103 | 0.055 | 0.056 | 0.055 | 1.04 | 0.17 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 142 | 0.058 | 0.0583 | 0.0594 | 1.26 | 0.21 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 204 | 0.0643 | 0.0652 | 0.0665 | 1.69 | 0.28 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 144 | 0.0778 | 0.0773 | 0.0751 | 1.87 | 0.32 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 199 | 0.064 | 0.065 | 0.0625 | 1.97 | 0.33 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 189 | 0.0945 | 0.0979 | 0.0974 | 1.9 | 0.34 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 200 | 0.066 | 0.064 | 0.067 | 2.33 | 0.39 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 148 | 0.064 | 0.067 | 0.064 | 2.66 | 0.44 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 196 | 0.0507 | 0.0516 | 0.0539 | 3.17 | 0.51 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.99 | Other | 183 | 0.085 | 0.088 | 0.083 | 2.95 | 0.51 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 122 | 0.0686 | 0.0729 | 0.0707 | 3.04 | 0.51 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 194 | 0.075 | 0.0737 | 0.0703 | 3.32 | 0.56 | 4.9 | 0.21 | 16.2 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 155 | 0.076 | 0.075 | 0.071 | 3.58 | 0.61 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 164 | 0.0678 | 0.072 | 0.0727 | 3.74 | 0.63 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 113 | 0.0616 | 0.0666 | 0.0645 | 3.91 | 0.65 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 175 | 0.0628 | 0.0591 | 0.0642 | 4.25 | 0.7 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 191 | 0.077 | 0.0714 | 0.0719 | 4.22 | 0.71 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.03 | AOAC 2018.11, M | 139 | 0.0675 | 0.0626 | 0.0686 | 4.82 | 0.8 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 172 | 0.101 | 0.103 | 0.094 | 4.76 | 0.84 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 225 | 0.0497 | 0.0524 | 0.0466 | 5.86 | 0.93 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 101 | 0.0822 | 0.0736 | 0.0769 | 5.59 | 0.96 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 161 | 0.0524 | 0.0589 | 0.0581 | 6.28 | 1.02 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 125 | 0.0776 | 0.0736 | 0.0832 | 6.17 | 1.06 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 120 | 0.065 | 0.057 | 0.0611 | 6.55 | 1.08 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 182 | 0.067 | 0.0657 | 0.059 | 6.72 | 1.11 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 116 | 0.0689 | 0.0693 | 0.0783 | 7.37 | 1.24 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 102 | 0.0726 | 0.0834 | 0.0816 | 7.31 | 1.25 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 106 | 0.07 | 0.07 | 0.08 | 7.87 | 1.33 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 224 | 0.2738 | 0.243 | 0.2453 | 6.75 | 1.38 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 149 | 0.068 | 0.081 | 0.074 | 8.75 | 1.49 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 152 | 0.069 | 0.0811 | 0.0694 | 9.39 | 1.59 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 170 | 0.0611 | 0.0782 | 0.0763 | 13 | 2.2 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 192 | 0.0811 | 0.0676 | 0.0612 | 14.5 | 2.44 | 4.9 | 0.21 | 16.2 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 217 | 0.0726 | 0.0698 | 0.0904 | 14.4 | 2.46 | 4.9 | 0.21 | 16.2 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 163 | 0.062 | 0.045 | 0.058 | 16.2 | 2.62 | 4.9 | 0.21 | 16.2 | 0 |
|  |  |  |  | CBD (%AR) |  |  |  |  |  |  |  |  |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 199 | 1.8445 | 1.85 | 1.846 | 0.15 | 0.04 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 204 | 1.882 | 1.8692 | 1.8701 | 0.38 | 0.1 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 189 | 1.8179 | 1.8133 | 1.8273 | 0.39 | 0.11 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.03 | AOAC 2018.11, M | 139 | 1.8539 | 1.8779 | 1.874 | 0.69 | 0.19 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 122 | 1.9553 | 1.9849 | 1.9701 | 0.75 | 0.21 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 116 | 1.9432 | 1.975 | 1.9641 | 0.82 | 0.23 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 113 | 2.376 | 2.4 | 2.42 | 0.92 | 0.26 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 125 | 1.67 | 1.63 | 1.66 | 1.26 | 0.34 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 170 | 2.1422 | 2.1137 | 2.0862 | 1.32 | 0.37 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 106 | 1.94 | 1.98 | 1.93 | 1.36 | 0.38 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 197 | 1.9204 | 1.8867 | 1.9431 | 1.48 | 0.41 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 161 | 1.6398 | 1.6424 | 1.5972 | 1.56 | 0.42 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 200 | 1.75 | 1.74 | 1.806 | 2.02 | 0.55 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 149 | 1.794 | 1.802 | 1.864 | 2.11 | 0.58 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 191 | 1.8717 | 1.8038 | 1.8055 | 2.12 | 0.58 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 196 | 1.6726 | 1.7454 | 1.7321 | 2.26 | 0.61 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.99 | Other | 132 | 1.67 | 1.641 | 1.596 | 2.28 | 0.62 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 194 | 1.9511 | 1.9714 | 1.8809 | 2.45 | 0.68 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 163 | 1.955 | 2.054 | 2.038 | 2.64 | 0.73 | 2.63 | 0.15 | 10.3 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%AR) | 003.10 | Other LC-UV | 142 | 1.71 | 1.8 | 1.72 | 2.83 | 0.77 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 144 | 1.832 | 1.7109 | 1.775 | 3.42 | 0.93 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 164 | 1.6482 | 1.6534 | 1.7512 | 3.45 | 0.93 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 192 | 1.8225 | 1.8164 | 1.7058 | 3.69 | 1.01 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 184 | 1.698 | 1.768 | 1.841 | 4.04 | 1.1 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 172 | 1.837 | 2.01 | 1.89 | 4.63 | 1.28 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 103 | 1.796 | 1.645 | 1.662 | 4.86 | 1.32 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 102 | 1.8234 | 1.7835 | 1.9669 | 5.19 | 1.43 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 152 | 1.8106 | 2.0231 | 1.92 | 5.54 | 1.53 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 155 | 1.52 | 1.61 | 1.8 | 8.7 | 2.35 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.60 | NIR | 187 | 2.23 | 2.45 | 2.69 | 9.37 | 2.69 | 2.63 | 0.15 | 10.3 | 0 |
| CBD (%AR) | 003.99 | Other | 183 | 1.863 | 1.53 | 1.806 | 10.3 | 2.8 | 2.63 | 0.15 | 10.3 | 0 |
|  |  |  |  | CBDA (%AR) |  |  |  |  |  |  |  |  |
| CBDA (%AR) | 004.60 | NIR | 187 | 0.55 | 0.41 | <0.0100 |  |  |  |  |  | 1 |
| CBDA (%AR) | 004.30 | Other LC-MS | 122 | 1.0278 | 1.0256 | 1.0267 | 0.11 | 0.03 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 204 | 1.0447 | 1.048 | 1.0423 | 0.27 | 0.07 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 189 | 0.9598 | 0.956 | 0.9533 | 0.34 | 0.08 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 199 | 1.0315 | 1.0345 | 1.0245 | 0.5 | 0.13 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 161 | 0.8457 | 0.8328 | 0.8374 | 0.78 | 0.19 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 197 | 0.9639 | 0.9704 | 0.9886 | 1.31 | 0.33 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 116 | 1.0072 | 1.0291 | 1.0016 | 1.44 | 0.36 | 2.37 | 0.11 | 6.83 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDA (%AR) | 004.10 | Other LC-UV | 144 | 1.1298 | 1.0942 | 1.1041 | 1.66 | 0.42 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 172 | 1.032 | 1.045 | 1.067 | 1.69 | 0.43 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.03 | AOAC 2018.11, M | 139 | 1.162 | 1.1689 | 1.1994 | 1.69 | 0.43 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 142 | 0.885 | 0.912 | 0.883 | 1.81 | 0.45 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 103 | 0.957 | 0.92 | 0.954 | 2.18 | 0.54 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 106 | 1.03 | 1.07 | 1.03 | 2.21 | 0.56 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.99 | Other | 183 | 1.14 | 1.193 | 1.163 | 2.28 | 0.58 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 125 | 0.965 | 0.922 | 0.965 | 2.61 | 0.65 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 194 | 1.1457 | 1.1481 | 1.0948 | 2.67 | 0.68 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 196 | 0.869 | 0.911 | 0.9163 | 2.88 | 0.71 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.99 | Other | 132 | 0.822 | 0.789 | 0.774 | 3.09 | 0.75 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 192 | 0.6943 | 0.6901 | 0.655 | 3.17 | 0.75 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 163 | 0.996 | 0.937 | 0.985 | 3.23 | 0.8 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 191 | 1.1397 | 1.0762 | 1.0727 | 3.44 | 0.87 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 164 | 1.1349 | 1.1907 | 1.2122 | 3.38 | 0.87 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 149 | 0.967 | 1.018 | 1.044 | 3.88 | 0.97 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 184 | 1.8904 | 1.982 | 2.0366 | 3.75 | 1.04 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 200 | 0.904 | 0.88 | 0.958 | 4.37 | 1.08 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 152 | 0.9508 | 1.0284 | 1.0308 | 4.54 | 1.14 | 2.37 | 0.11 | 6.83 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 155 | 0.835 | 0.874 | 0.954 | 6.83 | 1.68 | 2.37 | 0.11 | 6.83 | 0 |

Population of Lab RSDr

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

RSDr	HorRat(r)

Population of Lab RSDr	 Rob Mean	min	max

Flag

CBN (%AR)

| CBN (%AR) | 005.10 | Other LC-UV | 197 | <0.0694 | <0.0739 | <0.0711 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%AR) | 005.99 | Other | 132 | <0.1000 | <0.1000 | <0.1000 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 189 | 0.0633 | 0.0631 | 0.063 | 0.24 | 0.04 | 3.57 | 0.21 | 12.7 | 0 |
| CBN (%AR) | 005.20 | Other LC-MS | 122 | 0.071 | 0.0713 | 0.0711 | 0.22 | 0.04 | 3.57 | 0.21 | 12.7 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 204 | 0.0562 | 0.0557 | 0.0565 | 0.72 | 0.12 | 3.57 | 0.21 | 12.7 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 116 | 0.0812 | 0.0803 | 0.0817 | 0.88 | 0.15 | 3.57 | 0.21 | 12.7 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 142 | 0.0507 | 0.0507 | 0.0499 | 0.92 | 0.15 | 3.57 | 0.21 | 12.7 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 194 | 0.0544 | 0.0548 | 0.0533 | 1.43 | 0.23 | 3.57 | 0.21 | 12.7 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 144 | 0.0488 | 0.0503 | 0.0497 | 1.52 | 0.24 | 3.57 | 0.21 | 12.7 | 0 |
| CBN (%AR) | 005.99 | Other | 183 | 0.064 | 0.064 | 0.062 | 1.82 | 0.3 | 3.57 | 0.21 | 12.7 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 200 | 0.056 | 0.056 | 0.058 | 2.04 | 0.33 | 3.57 | 0.21 | 12.7 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 149 | 0.055 | 0.053 | 0.053 | 2.15 | 0.35 | 3.57 | 0.21 | 12.7 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 192 | 0.0478 | 0.0474 | 0.0458 | 2.25 | 0.36 | 3.57 | 0.21 | 12.7 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 192 | 0.0344 | 0.0358 | 0.036 | 2.46 | 0.37 | 3.57 | 0.21 | 12.7 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 125 | 0.0615 | 0.0587 | 0.06 | 2.33 | 0.38 | 3.57 | 0.21 | 12.7 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 199 | 0.0575 | 0.056 | 0.059 | 2.61 | 0.43 | 3.57 | 0.21 | 12.7 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 196 | 0.0465 | 0.0482 | 0.0493 | 2.94 | 0.47 | 3.57 | 0.21 | 12.7 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 172 | 0.055 | 0.059 | 0.056 | 3.67 | 0.6 | 3.57 | 0.21 | 12.7 | 0 |
| CBN (%AR) | 005.03 | AOAC 2018.11, M | 139 | 0.0531 | 0.0496 | 0.0533 | 4 | 0.64 | 3.57 | 0.21 | 12.7 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 138 | 0.0627 | 0.0722 | 0.0672 | 7.05 | 1.18 | 3.57 | 0.21 | 12.7 | 0 |
| CBN (%AR) | 005.20 | Other LC-MS | 102 | 0.0524 | 0.0582 | 0.0612 | 7.81 | 1.27 | 3.57 | 0.21 | 12.7 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%AR) | 005.10 | Other LC-UV | 106 | 0.07 | 0.06 | 0.06 | 9.12 | 1.51 | 3.57 | 0.21 | 12.7 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 116 | 0.0701 | 0.0798 | 0.0672 | 9.12 | 1.54 | 3.57 | 0.21 | 12.7 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 163 | 0.073 | 0.085 | 0.07 | 10.4 | 1.78 | 3.57 | 0.21 | 12.7 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 107 | 0.0546 | 0.0628 | 0.0705 | 12.7 | 2.1 | 3.57 | 0.21 | 12.7 | 0 |
|  |  |  |  | Total Δ9-THC (%AR) |  |  |  |  |  |  |  |  |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 190 | 0.273 | 0.279 |  |  |  |  |  |  | 1 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 199 | 0.2581 | 0.2595 | 0.2593 | 0.29 | 0.06 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 189 | 0.2979 | 0.2996 | 0.3005 | 0.44 | 0.09 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 204 | 0.2494 | 0.2472 | 0.2494 | 0.51 | 0.1 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 103 | 0.2072 | 0.2101 | 0.2092 | 0.71 | 0.14 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 142 | 0.249 | 0.253 | 0.249 | 0.92 | 0.19 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 122 | 0.2707 | 0.2763 | 0.2735 | 1.02 | 0.21 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 120 | 0.2548 | 0.2482 | 0.2522 | 1.32 | 0.27 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 164 | 0.2374 | 0.2396 | 0.2334 | 1.33 | 0.27 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 197 | 0.1759 | 0.1813 | 0.1782 | 1.52 | 0.29 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 169 | 0.2517 | 0.2519 | 0.2457 | 1.41 | 0.29 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 175 | 0.2729 | 0.2661 | 0.2657 | 1.51 | 0.31 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 172 | 0.2816 | 0.2913 | 0.2874 | 1.7 | 0.35 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 144 | 0.2732 | 0.2809 | 0.2826 | 1.8 | 0.37 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 161 | 0.2117 | 0.2195 | 0.2127 | 1.98 | 0.39 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 115 | 0.3 | 0.29 | 0.3 | 1.95 | 0.41 | 3.06 | 0.29 | 12 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 200 | 0.241 | 0.236 | 0.246 | 2.07 | 0.42 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 191 | 0.26 | 0.25 | 0.25 | 2.28 | 0.46 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.99 | Other | 183 | 0.277 | 0.285 | 0.271 | 2.53 | 0.52 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 116 | 0.2664 | 0.2805 | 0.2734 | 2.58 | 0.53 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 163 | 0.276 | 0.288 | 0.275 | 2.59 | 0.54 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 153 | 0.2797 | 0.2678 | 0.2666 | 2.67 | 0.55 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 196 | 0.2132 | 0.2247 | 0.2228 | 2.8 | 0.56 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 121 | 0.3 | 0.316 | 0.311 | 2.65 | 0.56 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 194 | 0.2704 | 0.2747 | 0.2592 | 2.98 | 0.61 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 138 | 0.2389 | 0.2511 | 0.2555 | 3.46 | 0.7 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.99 | Other | 132 | 0.221 | 0.214 | 0.206 | 3.51 | 0.7 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 114 | 0.2145 | 0.2305 | 0.2243 | 3.62 | 0.72 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.60 | NIR | 138 | 0.485 | 0.515 | 0.49 | 3.24 | 0.73 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 116 | 0.3227 | 0.3181 | 0.34 | 3.53 | 0.75 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.60 | NIR | 187 | 0.13 | 0.14 | 0.14 | 4.22 | 0.79 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 101 | 0.2881 | 0.2667 | 0.275 | 3.9 | 0.81 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 148 | 0.173 | 0.172 | 0.16 | 4.3 | 0.82 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 182 | 0.2159 | 0.2108 | 0.1987 | 4.24 | 0.84 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 155 | 0.256 | 0.266 | 0.278 | 4.13 | 0.85 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 170 | 0.2767 | 0.2931 | 0.2709 | 4.11 | 0.85 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 125 | 0.255 | 0.235 | 0.25 | 4.22 | 0.86 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 192 | 0.1916 | 0.1814 | 0.1747 | 4.66 | 0.91 | 3.06 | 0.29 | 12 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 102 | 0.2493 | 0.2708 | 0.2716 | 4.79 | 0.98 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 107 | 0.2544 | 0.2765 | 0.2773 | 4.82 | 0.99 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 149 | 0.236 | 0.26 | 0.247 | 4.85 | 0.99 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 192 | 0.1373 | 0.1527 | 0.1451 | 5.31 | 1 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.03 | AOAC 2018.11, M | 139 | 0.3081 | 0.2837 | 0.3104 | 4.92 | 1.03 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 225 | 0.2278 | 0.2415 | 0.2117 | 6.57 | 1.32 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 152 | 0.2925 | 0.2964 | 0.2628 | 6.47 | 1.34 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 217 | 0.2439 | 0.2729 | 0.298 | 9.97 | 2.05 | 3.06 | 0.29 | 12 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 128 | 0.3548 | 0.2852 | 0.3535 | 12 | 2.55 | 3.06 | 0.29 | 12 | 0 |
|  |  |  |  | Total CBD (%AR) |  |  |  |  |  |  |  |  |
| Total CBD (%AR) | 007.10 | Other LC-UV | 189 | 2.6596 | 2.6518 | 2.6633 | 0.22 | 0.06 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 199 | 2.7491 | 2.7573 | 2.7445 | 0.24 | 0.07 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 204 | 2.7982 | 2.7883 | 2.7842 | 0.26 | 0.08 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 122 | 2.8567 | 2.8843 | 2.8705 | 0.48 | 0.14 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 192 | 1.6368 | 1.659 | 1.6653 | 0.9 | 0.24 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.03 | AOAC 2018.11, M | 139 | 2.8731 | 2.903 | 2.9259 | 0.91 | 0.27 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 116 | 2.8265 | 2.8775 | 2.8425 | 0.92 | 0.27 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 161 | 2.3815 | 2.3728 | 2.3316 | 1.13 | 0.32 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 197 | 2.7657 | 2.7377 | 2.8101 | 1.32 | 0.38 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 169 | 2.8274 | 2.8875 | 2.7979 | 1.61 | 0.47 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 125 | 2.51 | 2.43 | 2.5 | 1.76 | 0.5 | 2.77 | 0.22 | 8.45 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%AR) | 007.60 | NIR | 187 | 2.71 | 2.81 | 2.69 | 2.35 | 0.68 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 196 | 2.4348 | 2.5443 | 2.5357 | 2.43 | 0.7 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.99 | Other | 132 | 2.392 | 2.334 | 2.275 | 2.51 | 0.71 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 142 | 2.49 | 2.6 | 2.49 | 2.51 | 0.72 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 191 | 2.87 | 2.75 | 2.75 | 2.48 | 0.73 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 149 | 2.642 | 2.695 | 2.779 | 2.55 | 0.74 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 194 | 2.9559 | 2.9783 | 2.841 | 2.52 | 0.74 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 200 | 2.543 | 2.512 | 2.646 | 2.73 | 0.79 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 144 | 2.8231 | 2.6707 | 2.7435 | 2.78 | 0.81 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 164 | 2.6434 | 2.6977 | 2.8142 | 3.21 | 0.93 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 172 | 2.7421 | 2.9265 | 2.8258 | 3.26 | 0.96 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 192 | 2.4314 | 2.4215 | 2.2802 | 3.56 | 1.02 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 116 | 2.9876 | 3.1874 | 3.1682 | 3.54 | 1.05 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 103 | 2.6353 | 2.4518 | 2.4987 | 3.77 | 1.09 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 163 | 2.698 | 2.914 | 2.796 | 3.86 | 1.13 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 170 | 2.5371 | 2.7267 | 2.5677 | 3.9 | 1.13 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 152 | 2.7594 | 3.0514 | 2.9507 | 5.08 | 1.49 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.99 | Other | 183 | 2.863 | 2.577 | 2.826 | 5.65 | 1.65 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 138 | 2.961 | 3.2262 | 3.3319 | 6.02 | 1.79 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.60 | NIR | 138 | 1.61 | 1.47 | 1.74 | 8.4 | 2.26 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 155 | 2.25 | 2.37 | 2.64 | 8.25 | 2.36 | 2.77 | 0.22 | 8.45 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 107 | 2.4653 | 2.7864 | 2.911 | 8.45 | 2.46 | 2.77 | 0.22 | 8.45 | 0 |

Population of Lab RSDr

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

RSDr	HorRat(r)

Population of Lab RSDr	 Rob Mean	min	max

Flag

CBDV (%AR)

| CBDV (%AR) | 008.99 | Other | 194 | <0.0270 | <0.0270 | <0.0270 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDV (%AR) | 008.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 |  |  | 1 |
| CBDV (%AR) | 008.99 | Other | 132 | <0.1000 | <0.1000 | <0.1000 |  |  | 1 |
| CBDV (%AR) | 008.99 | Other | 197 | <0.0463 | <0.0493 | <0.0474 |  |  | 1 |
| CBDV (%AR) | 008.99 | Other | 200 | 0.031 | 0.031 | 0.032 | 1.84 | 0.27 | 0 |
| CBDV (%AR) | 008.99 | Other | 142 | 0.0294 | 0.0299 | 0.0288 | 1.88 | 0.28 | 0 |
| CBDV (%AR) | 008.99 | Other | 183 | 0.034 | 0.036 | 0.034 | 3.33 | 0.5 | 0 |
| CBDV (%AR) | 008.99 | Other | 189 | 0.029 | 0.0258 | 0.0279 | 5.9 | 0.86 | 0 |
| CBDV (%AR) | 008.99 | Other | 106 | 0.09 | 0.08 | 0.07 | 12.5 | 2.14 | 0 |

CBDVA (%AR)

| CBDVA (%AR) | 009.99 | Other | 197 | <0.0463 | <0.0493 | <0.0474 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDVA (%AR) | 009.99 | Other | 149 | 0.028 | <0.0250 | <0.0250 |  |  | 1 |
| CBDVA (%AR) | 009.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 |  |  | 1 |
| CBDVA (%AR) | 009.99 | Other | 142 | 0.0263 | 0.0254 | 0.0255 | 1.92 | 0.28 | 0 |
| CBDVA (%AR) | 009.99 | Other | 183 | 0.037 | 0.04 | 0.038 | 3.98 | 0.61 | 0 |

CBG (%AR)

| CBG (%AR) | 010.99 | Other | 204 | 0.8452 | 0.846 | 0.8498 | 0.29 | 0.07 | 2.27 | 0.29 | 7.1 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%AR) | 010.99 | Other | 189 | 0.8817 | 0.8791 | 0.8746 | 0.41 | 0.1 | 2.27 | 0.29 | 7.1 | 0 |
| CBG (%AR) | 010.99 | Other | 183 | 0.856 | 0.845 | 0.853 | 0.67 | 0.16 | 2.27 | 0.29 | 7.1 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%AR) | 010.99 | Other | 197 | 0.738 | 0.7251 | 0.7336 | 0.9 | 0.21 | 2.27 | 0.29 | 7.1 | 0 |
| CBG (%AR) | 010.99 | Other | 200 | 0.719 | 0.733 | 0.739 | 1.41 | 0.34 | 2.27 | 0.29 | 7.1 | 0 |
| CBG (%AR) | 010.99 | Other | 142 | 0.763 | 0.782 | 0.755 | 1.81 | 0.44 | 2.27 | 0.29 | 7.1 | 0 |
| CBG (%AR) | 010.99 | Other | 122 | 1.0097 | 0.9664 | 0.9881 | 2.19 | 0.55 | 2.27 | 0.29 | 7.1 | 0 |
| CBG (%AR) | 010.99 | Other | 194 | 0.8549 | 0.8732 | 0.834 | 2.3 | 0.56 | 2.27 | 0.29 | 7.1 | 0 |
| CBG (%AR) | 010.99 | Other | 132 | 0.757 | 0.736 | 0.719 | 2.58 | 0.62 | 2.27 | 0.29 | 7.1 | 0 |
| CBG (%AR) | 010.99 | Other | 196 | 0.7161 | 0.7505 | 0.7506 | 2.69 | 0.64 | 2.27 | 0.29 | 7.1 | 0 |
| CBG (%AR) | 010.99 | Other | 149 | 0.795 | 0.804 | 0.838 | 2.79 | 0.68 | 2.27 | 0.29 | 7.1 | 0 |
| CBG (%AR) | 010.99 | Other | 106 | 0.79 | 0.81 | 0.76 | 3.2 | 0.77 | 2.27 | 0.29 | 7.1 | 0 |
| CBG (%AR) | 010.02 | AOAC 2018.11, U | 125 | 0.736 | 0.716 | 0.774 | 3.97 | 0.95 | 2.27 | 0.29 | 7.1 | 0 |
| CBG (%AR) | 010.99 | Other | 172 | 0.83 | 0.856 | 0.787 | 4.23 | 1.03 | 2.27 | 0.29 | 7.1 | 0 |
| CBG (%AR) | 010.99 | Other | 155 | 0.69 | 0.722 | 0.792 | 7.1 | 1.7 | 2.27 | 0.29 | 7.1 | 0 |
|  |  |  |  | CBGA (%AR) |  |  |  |  |  |  |  |  |
| CBGA (%AR) | 011.99 | Other | 122 | 1.2842 | 1.2797 | 1.2819 | 0.18 | 0.05 | 2.26 | 0.18 | 6.45 | 0 |
| CBGA (%AR) | 011.99 | Other | 189 | 1.2146 | 1.2104 | 1.2113 | 0.18 | 0.05 | 2.26 | 0.18 | 6.45 | 0 |
| CBGA (%AR) | 011.99 | Other | 204 | 1.1929 | 1.2008 | 1.2001 | 0.36 | 0.09 | 2.26 | 0.18 | 6.45 | 0 |
| CBGA (%AR) | 011.99 | Other | 183 | 1.288 | 1.29 | 1.304 | 0.67 | 0.18 | 2.26 | 0.18 | 6.45 | 0 |
| CBGA (%AR) | 011.99 | Other | 149 | 1.216 | 1.256 | 1.252 | 1.77 | 0.46 | 2.26 | 0.18 | 6.45 | 0 |
| CBGA (%AR) | 011.99 | Other | 142 | 1.14 | 1.18 | 1.18 | 1.98 | 0.51 | 2.26 | 0.18 | 6.45 | 0 |
| CBGA (%AR) | 011.99 | Other | 106 | 1.2 | 1.21 | 1.16 | 2.22 | 0.57 | 2.26 | 0.18 | 6.45 | 0 |
| CBGA (%AR) | 011.99 | Other | 196 | 1.123 | 1.1592 | 1.1776 | 2.41 | 0.62 | 2.26 | 0.18 | 6.45 | 0 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%AR) | 011.99 | Other | 194 | 1.2878 | 1.2971 | 1.2378 | 2.5 | 0.65 | 2.26 | 0.18 | 6.45 | 0 |
| CBGA (%AR) | 011.99 | Other | 132 | 1.086 | 1.023 | 1.025 | 3.43 | 0.86 | 2.26 | 0.18 | 6.45 | 0 |
| CBGA (%AR) | 011.99 | Other | 172 | 1.283 | 1.303 | 1.209 | 3.91 | 1.02 | 2.26 | 0.18 | 6.45 | 0 |
| CBGA (%AR) | 011.99 | Other | 197 | 1.287 | 1.1793 | 1.273 | 4.7 | 1.22 | 2.26 | 0.18 | 6.45 | 0 |
| CBGA (%AR) | 011.02 | AOAC 2018.11, U | 125 | 1.17 | 1.06 | 1.2 | 6.45 | 1.65 | 2.26 | 0.18 | 6.45 | 0 |

Lab

Population of Lab RSDr

THCV (%AR)

| THCV (%AR) | 012.99 | Other | 197 | <0.0463 | <0.0493 | <0.0474 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| THCV (%AR) | 012.99 | Other | 106 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 132 | <0.1000 | <0.1000 | <0.1000 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 194 | <0.0270 | <0.0270 | <0.0270 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 149 | <0.0250 | <0.0250 | <0.0250 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 200 | 0.02 | 0.02 | 0.021 | 2.84 | 0.4 | 0 |
| THCV (%AR) | 012.99 | Other | 183 | 0.021 | 0.023 | 0.021 | 5.33 | 0.75 | 0 |
| THCV (%AR) | 012.99 | Other | 189 | 0.0574 | 0.0567 | 0.0343 | 26.6 | 4.24 | 0 |
| THCV (%AR) | 012.99 | Other | 142 | 0.0291 | 0.0262 | 0.298 | 133 | 24.1 | 0 |

Δ8-THC (%AR)

| Δ8-THC (%AR) | 013.99 | Other | 106 | <0.0100 | <0.0100 | <0.0100 | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%AR) | 013.99 | Other | 183 | 0 | 0 | 0 | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 200 | <0.0100 | <0.0100 | <0.0100 | 1 |

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) Rob Mean min max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%AR) | 013.99 | Other | 142 | <0.0022 | <0.0022 | <0.0022 |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 197 | <0.0463 | <0.0493 | <0.0474 |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 194 | <0.0270 | <0.0270 | <0.0270 |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 102 | <0.0197 | <0.0095 | <0.0199 |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 149 | <0.0250 | <0.0250 | <0.0250 |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 138 | 0.0492 | 0.054 | 0.0535 | 5.05 | 0.81 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 189 | 0.0085 | 0.0079 | 0.0072 | 8.27 | 1 | 0 |

Lab

Population of Lab RSDr

CBC (%AR)

| CBC (%AR) | 014.99 | Other | 189 | 0.2086 | 0.2087 | 0.2087 | 0.03 | 0.01 | 2.33 | 0.03 | 7.92 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%AR) | 014.99 | Other | 204 | 0.1991 | 0.1958 | 0.2003 | 1.17 | 0.23 | 2.33 | 0.03 | 7.92 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, U | 125 | 0.178 | 0.175 | 0.179 | 1.17 | 0.23 | 2.33 | 0.03 | 7.92 | 0 |
| CBC (%AR) | 014.99 | Other | 196 | 0.1595 | 0.1645 | 0.1629 | 1.57 | 0.3 | 2.33 | 0.03 | 7.92 | 0 |
| CBC (%AR) | 014.99 | Other | 142 | 0.176 | 0.178 | 0.172 | 1.74 | 0.34 | 2.33 | 0.03 | 7.92 | 0 |
| CBC (%AR) | 014.99 | Other | 200 | 0.2 | 0.204 | 0.208 | 1.96 | 0.39 | 2.33 | 0.03 | 7.92 | 0 |
| CBC (%AR) | 014.99 | Other | 197 | 0.175 | 0.1754 | 0.182 | 2.22 | 0.43 | 2.33 | 0.03 | 7.92 | 0 |
| CBC (%AR) | 014.99 | Other | 149 | 0.173 | 0.181 | 0.179 | 2.34 | 0.45 | 2.33 | 0.03 | 7.92 | 0 |
| CBC (%AR) | 014.99 | Other | 183 | 0.206 | 0.206 | 0.197 | 2.56 | 0.5 | 2.33 | 0.03 | 7.92 | 0 |
| CBC (%AR) | 014.99 | Other | 155 | 0.176 | 0.177 | 0.185 | 2.75 | 0.53 | 2.33 | 0.03 | 7.92 | 0 |
| CBC (%AR) | 014.99 | Other | 194 | 0.1931 | 0.1959 | 0.1856 | 2.78 | 0.54 | 2.33 | 0.03 | 7.92 | 0 |
| CBC (%AR) | 014.99 | Other | 132 | 0.163 | 0.16 | 0.154 | 2.88 | 0.55 | 2.33 | 0.03 | 7.92 | 0 |

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%AR) | 014.99 | Other | 106 | 0.19 | 0.19 | 0.18 | 3.09 | 0.6 | 2.33 | 0.03 | 7.92 | 0 |
| CBC (%AR) | 014.99 | Other | 172 | 0.193 | 0.209 | 0.194 | 4.51 | 0.89 | 2.33 | 0.03 | 7.92 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, U | 103 | 0.17 | 0.173 | 0.196 | 7.92 | 1.53 | 2.33 | 0.03 | 7.92 | 0 |

Lab

Population of Lab RSDr

Cd (ug/g AR)

| Cd (ug/g AR) | 101.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cd (ug/g AR) | 101.04 | ICP-MS | 224 | <0.2225 | <0.2061 | <0.2166 |  |  |  |  |  | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 102 | 0.1251 | 0.1274 | 0.1274 | 1.05 | 0.19 | 4.39 | 1.05 | 8.27 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 194 | 0.1322 | 0.1286 | 0.1273 | 1.96 | 0.36 | 4.39 | 1.05 | 8.27 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 183 | 0.097 | 0.095 | 0.092 | 2.66 | 0.47 | 4.39 | 1.05 | 8.27 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 197 | 0.136 | 0.1459 | 0.1357 | 4.17 | 0.78 | 4.39 | 1.05 | 8.27 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 196 | 0.118 | 0.107 | 0.111 | 4.97 | 0.9 | 4.39 | 1.05 | 8.27 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 106 | 0.165 | 0.15 | 0.142 | 7.66 | 1.45 | 4.39 | 1.05 | 8.27 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 142 | 0.135 | 0.136 | 0.117 | 8.27 | 1.52 | 4.39 | 1.05 | 8.27 | 0 |

As (ug/g AR)

| As (ug/g AR) | 102.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| As (ug/g AR) | 102.04 | ICP-MS | 197 | <0.0977 | <0.0956 | <0.0986 |  |  |  |  |  | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 196 | 0.101 | 0.0944 | 0.1 | 3.61 | 0.64 | 8.82 | 3.61 | 17.3 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 183 | 0.07 | 0.073 | 0.077 | 4.79 | 0.81 | 8.82 | 3.61 | 17.3 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 106 | 0.086 | 0.096 | 0.093 | 5.6 | 0.98 | 8.82 | 3.61 | 17.3 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 102 | 0.0797 | 0.0877 | 0.0975 | 10.1 | 1.76 | 8.82 | 3.61 | 17.3 | 0 |

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| As (ug/g AR) | 102.04 | ICP-MS | 224 | 0.1377 | 0.1176 | 0.1143 | 10.3 | 1.88 | 8.82 | 3.61 | 17.3 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 194 | 0.1197 | 0.1254 | 0.1012 | 11 | 1.99 | 8.82 | 3.61 | 17.3 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 142 | 0.115 | 0.091 | 0.083 | 17.3 | 3.05 | 8.82 | 3.61 | 17.3 | 0 |

Lab

Population of Lab RSDr

Hg (ug/g AR)

| Hg (ug/g AR) | 103.04 | ICP-MS | 224 | <0.0222 | <0.0206 | <0.0217 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Hg (ug/g AR) | 103.04 | ICP-MS | 197 | <0.0977 | <0.0956 | <0.0986 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 194 | <0.0200 | <0.0200 | <0.0200 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 142 | <0.0500 | <0.0500 | <0.0500 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 196 | 0.0083 | 0.0079 | 0.0085 | 3.71 | 0.45 | 0 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 102 | 0.0095 | 0.0091 | 0.0088 | 3.85 | 0.48 | 0 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 183 | 0.049 | 0.045 | 0.041 | 8.89 | 1.4 | 0 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 106 | 0.011 | 0.009 | 0.009 | 11.9 | 1.49 | 0 |

Pb (ug/g AR)

| Pb (ug/g AR) | 104.04 | ICP-MS | 162 | 0.7167 | 0.6857 | 0.6594 | 4.17 | 0.99 | 9.58 | 4.17 | 19.2 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Pb (ug/g AR) | 104.04 | ICP-MS | 196 | 0.502 | 0.544 | 0.548 | 4.8 | 1.09 | 9.58 | 4.17 | 19.2 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 194 | 0.5172 | 0.4644 | 0.5176 | 6.12 | 1.38 | 9.58 | 4.17 | 19.2 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 142 | 0.464 | 0.524 | 0.481 | 6.32 | 1.42 | 9.58 | 4.17 | 19.2 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 197 | 0.5526 | 0.6543 | 0.6269 | 8.61 | 2 | 9.58 | 4.17 | 19.2 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 183 | 0.569 | 0.558 | 0.667 | 10 | 2.33 | 9.58 | 4.17 | 19.2 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Pb (ug/g AR) | 104.04 | ICP-MS | 102 | 0.7783 | 0.6324 | 0.6436 | 11.9 | 2.81 | 9.58 | 4.17 | 19.2 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 106 | 0.875 | 0.641 | 0.707 | 16.3 | 3.9 | 9.58 | 4.17 | 19.2 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 224 | 0.8215 | 0.7051 | 0.5558 | 19.2 | 4.55 | 9.58 | 4.17 | 19.2 | 0 |
|  |  |  |  | Cu (ug/g AR) |  |  |  |  |  |  |  |  |
| Cu (ug/g AR) | 105.04 | ICP-MS | 196 | 8.4 | 7.81 | 7.92 | 3.9 | 1.34 |  |  |  | 0 |
| Cu (ug/g AR) | 105.04 | ICP-MS | 142 | 13.62 | 14.2 | 12.7 | 5.6 | 2.07 |  |  |  | 0 |
|  |  |  |  | Ni (ug/g AR) |  |  |  |  |  |  |  |  |
| Ni (ug/g AR) | 106.04 | ICP-MS | 102 | 2.9943 | 2.8617 | 2.961 | 2.35 | 0.69 |  |  |  | 0 |
| Ni (ug/g AR) | 106.04 | ICP-MS | 196 | 1.34 | 1.26 | 1.3 | 3.08 | 0.8 |  |  |  | 0 |
| Ni (ug/g AR) | 106.04 | ICP-MS | 142 | 1.965 | 1.3 | 0.965 | 36.1 | 9.52 |  |  |  | 0 |
|  |  |  |  | Cr (ug/g AR) |  |  |  |  |  |  |  |  |
| Cr (ug/g AR) | 107.04 | ICP-MS | 196 | 0.875 | 0.877 | 0.812 | 4.32 | 1.06 |  |  |  | 0 |
| Cr (ug/g AR) | 107.04 | ICP-MS | 197 | 0.7828 | 0.8679 | 0.7859 | 5.94 | 1.44 |  |  |  | 0 |
| Cr (ug/g AR) | 107.04 | ICP-MS | 142 | 0.7 | 0.643 | 0.621 | 6.23 | 1.46 |  |  |  | 0 |
|  |  |  |  | Se (ug/g AR) |  |  |  |  |  |  |  |  |
| Se (ug/g AR) | 108.04 | ICP-MS | 102 | 0.1157 | 0.1133 | 0.1107 | 2.21 | 0.4 |  |  |  | 0 |
| Se (ug/g AR) | 108.04 | ICP-MS | 196 | 0.138 | 0.129 | 0.135 | 3.42 | 0.63 |  |  |  | 0 |
| Se (ug/g AR) | 108.04 | ICP-MS | 197 | 0.1269 | 0.1376 | 0.1184 | 7.54 | 1.39 |  |  |  | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Se (ug/g AR) | 108.04 | ICP-MS | 142 | 0.108 | 0.085 | 0.078 | 17.4 | 3.04 |  |  |  | 0 |
|  |  |  |  | Ag (ug/g AR) |  |  |  |  |  |  |  |  |
| Ag (ug/g AR) | 109.04 | ICP-MS | 142 | <0.0500 <0.0500 |  | <0.0500 |  |  |  |  |  | 1 |
| Ag (ug/g AR) | 109.04 | ICP-MS | 102 | 0.0049 | 0.0052 | 0.0045 | 7.22 | 0.81 |  |  |  | 0 |
|  |  |  |  | Ba (ug/g AR) |  |  |  |  |  |  |  |  |
| Ba (ug/g AR) | 110.04 | ICP-MS | 142 | 52.56 | 53.3 | 49.5 | 3.89 | 1.76 |  |  |  | 0 |
|  |  |  |  | Zn (ug/g AR) |  |  |  |  |  |  |  |  |
| Zn (ug/g AR) | 111.04 | ICP-MS | 196 | 59.6 | 55.9 | 56.8 | 3.36 | 1.55 |  |  |  | 0 |
|  |  |  |  | Sb (ug/g AR) |  |  |  |  |  |  |  |  |
| Sb (ug/g AR) | 112.04 | ICP-MS | 196 | <0.0241 <0.0241 |  | <0.0241 |  |  |  |  |  | 1 |
| Sb (ug/g AR) | 112.04 | ICP-MS | 102 | 0.0098 | 0.0131 | 0.0127 | 15.2 | 1.96 |  |  |  | 0 |
|  |  |  |  | (-)-alpha-Bisabolol (%AR) |  |  |  |  |  |  |  |  |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 194 | 0.0325 | 0.0325 | 0.0328 | 0.53 | 0.08 |  |  |  | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 162 | 0.0226 | 0.0239 | 0.0238 | 3.09 | 0.44 |  |  |  | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 183 | 0.056 | 0.055 | 0.061 | 5.61 | 0.92 |  |  |  | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 196 | 0.0155 | 0.0174 | 0.0153 | 7.21 | 0.97 |  |  |  | 0 |
| (-)-alpha-Bisabolol (%A | 201.01 | GC,FID | 142 | 0.0617 | 0.0547 | 0.0594 | 6.09 | 1 |  |  |  | 0 |

Population of Lab RSDr

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

RSDr	HorRat(r)

Population of Lab RSDr	 Rob Mean	min	max

Flag

beta-Caryophyllene (%AR)

| beta-Caryophyllene (% | 202.02 | GC,MS | 194 | 0.0141 | 0.013 | 0.0138 | 4.17 | 0.55 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| beta-Caryophyllene (% | 202.02 | GC,MS | 162 | 0.0171 | 0.0191 | 0.0177 | 5.71 | 0.78 | 0 |
| beta-Caryophyllene (% | 202.01 | GC,FID | 142 | 0.0229 | 0.0202 | 0.0224 | 6.58 | 0.93 | 0 |
| beta-Caryophyllene (% | 202.02 | GC,MS | 196 | 0.0136 | 0.0135 | 0.0117 | 8.27 | 1.08 | 0 |
| beta-Caryophyllene (% | 202.02 | GC,MS | 183 | 0.016 | 0.02 | 0.019 | 11.4 | 1.56 | 0 |

(-)-Caryophyllene oxide (%AR)

| (-)-Caryophyllene oxide | 203.02 | GC,MS | 194 | 0.0108 | 0.0099 | 0.0104 | 4.35 | 0.55 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| (-)-Caryophyllene oxide | 203.02 | GC,MS | 196 | 0.0049 | 0.0057 | 0.0059 | 9.62 | 1.1 | 0 |
| (-)-Caryophyllene oxide | 203.01 | GC,FID | 142 | 0.0134 | 0.0112 | 0.0133 | 9.83 | 1.28 | 0 |
| (-)-Caryophyllene oxide | 203.02 | GC,MS | 183 | 0.007 | 0.011 | 0.009 | 22.2 | 2.75 | 0 |

Geraniol (%AR)

| Geraniol (%AR) | 204.01 | GC,FID | 142 | <0.0010 | <0.0010 | <0.0010 | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Geraniol (%AR) | 204.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | 1 |
| Geraniol (%AR) | 204.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | 1 |
| Geraniol (%AR) | 204.02 | GC,MS | 183 | <0.0100 | <0.0100 | <0.0100 | 1 |
| Geraniol (%AR) | 204.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 | 1 |

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

RSDr	HorRat(r)

Population of Lab RSDr	 Rob Mean	min	max

Flag

alpha-Humulene (%AR)

| alpha-Humulene (%AR) | 205.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 194 | 0.0072 | 0.0067 | 0.007 | 3.61 | 0.43 | 0 |
| alpha-Humulene (%AR) | 205.01 | GC,FID | 142 | 0.0134 | 0.0119 | 0.0132 | 6.35 | 0.83 | 0 |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 196 | 0.0062 | 0.0062 | 0.0054 | 7.78 | 0.9 | 0 |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 183 | 0.006 | 0.006 | 0.007 | 9.12 | 1.07 | 0 |

d-Limonene (%AR)

| d-Limonene (%AR) | 206.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| d-Limonene (%AR) | 206.01 | GC,FID | 142 | <0.0010 | <0.0010 | <0.0010 | 1 |
| d-Limonene (%AR) | 206.02 | GC,MS | 183 | <0.0100 | <0.0100 | <0.0100 | 1 |
| d-Limonene (%AR) | 206.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | 1 |
| d-Limonene (%AR) | 206.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | 1 |

Linalool (%AR)

| Linalool (%AR) | 207.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Linalool (%AR) | 207.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 | 1 |
| Linalool (%AR) | 207.01 | GC,FID | 142 | <0.0010 | <0.0010 | <0.0010 | 1 |
| Linalool (%AR) | 207.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | 1 |
| Linalool (%AR) | 207.02 | GC,MS | 183 | <0.0100 | <0.0100 | <0.0100 | 1 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | beta-Myrcene (%AR) |  |  |  |  |  |  |  |  |
| beta-Myrcene (%AR) | 208.01 | GC,FID | 142 | <0.0010 | <0.0010 | <0.0010 |  |  |  |  |  | 1 |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 |  |  |  |  |  | 1 |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 183 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 194 | 0.0026 | 0.0026 | 0.0025 | 2.25 | 0.23 |  |  |  | 0 |
|  |  |  |  | Nopinene (%AR) |  |  |  |  |  |  |  |  |
| Nopinene (%AR) | 209.01 | GC,FID | 142 | <0.0010 <0.0010 |  | <0.0010 |  |  |  |  |  | 1 |
|  |  |  |  | Ocimene (%AR) |  |  |  |  |  |  |  |  |
| Ocimene (%AR) | 210.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| Ocimene (%AR) | 210.01 | GC,FID | 142 | <0.0010 | <0.0010 | <0.0010 |  |  |  |  |  | 1 |
| Ocimene (%AR) | 210.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 |  |  |  |  |  | 1 |
| Ocimene (%AR) | 210.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 |  |  |  |  |  | 1 |
| Ocimene (%AR) | 210.02 | GC,MS | 183 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
|  |  |  |  | alpha-Pinene (%AR) |  |  |  |  |  |  |  |  |
| alpha-Pinene (%AR) | 211.01 | GC,FID | 142 | <0.0010 | <0.0010 | <0.0010 |  |  |  |  |  | 1 |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 |  |  |  |  |  | 1 |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 |  |  |  |  |  | 1 |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 183 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |

Population of Lab RSDr

|  |  |  | Lab |  |  |  | Lab |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Analyte | Code | Method | Num | Result1 Result2 Result3 |  |  | RSDr | HorRat(r) |  | Rob Mean | min | max | Flag |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 162 | <0.0100 <0.0100 |  | <0.0100 |  | 1 |  |  |  |  |  |
|  |  |  |  | (-)-beta-Pinene (%AR) |  |  |  |  |  |  |  |  |  |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 183 | <0.0100 | <0.0100 | <0.0100 |  | 1 |  |  |  |  |  |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 |  | 1 |  |  |  |  |  |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 |  | 1 |  |  |  |  |  |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 |  | 1 |  |  |  |  |  |
|  |  |  |  | alpha-Terpinene (%AR) |  |  |  |  |  |  |  |  |  |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 |  | 1 |  |  |  |  |  |
| alpha-Terpinene (%AR) | 213.01 | GC,FID | 142 | <0.0010 | <0.0010 | <0.0010 |  | 1 |  |  |  |  |  |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 |  | 1 |  |  |  |  |  |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 |  | 1 |  |  |  |  |  |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 183 | <0.0100 | <0.0100 | <0.0100 |  | 1 |  |  |  |  |  |
|  |  |  |  | gamma-Terpinene (%AR) |  |  |  |  |  |  |  |  |  |
| gamma-Terpinene (%A | 214.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 |  | 1 |  |  |  |  |  |
| gamma-Terpinene (%A | 214.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 |  | 1 |  |  |  |  |  |
| gamma-Terpinene (%A | 214.02 | GC,MS | 183 | <0.0100 | <0.0100 | <0.0100 |  | 1 |  |  |  |  |  |
| gamma-Terpinene (%A | 214.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 |  | 1 |  |  |  |  |  |

Population of Lab RSDr

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

RSDr	HorRat(r)

Population of Lab RSDr	 Rob Mean	min	max

Flag

alpha-Terpinolene (%AR)

| alpha-Terpinolene (%A | 215.01 | GC,FID | 142 | <0.0010 | <0.0010 | <0.0010 | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | 1 |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 | 1 |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 183 | <0.0100 | <0.0100 | <0.0100 | 1 |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | 1 |

Moisture (%AR)

| Moisture (%AR) | 500.99 | Other | 142 | 5 | 5 | 5 | 0 | 0 | 3.79 | 0 | 14.4 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 189 | 4.894 | 4.894 | 4.894 | 0 | 0 | 3.79 | 0 | 14.4 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 204 | 6.12 | 6.12 | 6.12 | 0 | 0 | 3.79 | 0 | 14.4 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 210 | 2.887 | 2.8849 | 2.902 | 0.32 | 0.09 | 3.79 | 0 | 14.4 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 116 | 5.21 | 5.25 | 5.2 | 0.51 | 0.16 | 3.79 | 0 | 14.4 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 183 | 7.2 | 7.1 | 7.15 | 0.7 | 0.24 | 3.79 | 0 | 14.4 | 0 |
| Moisture (%AR) | 500.99 | Other | 105 | 8.42 | 8.31 | 8.32 | 0.73 | 0.25 | 3.79 | 0 | 14.4 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 199 | 8.8079 | 8.6485 | 8.7914 | 1 | 0.35 | 3.79 | 0 | 14.4 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 172 | 6.566 | 6.412 | 6.518 | 1.21 | 0.4 | 3.79 | 0 | 14.4 | 0 |
| Moisture (%AR) | 500.72 | NIR | 138 | 7.45 | 7.25 | 7.15 | 2.1 | 0.71 | 3.79 | 0 | 14.4 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 194 | 7.77 | 7.4 | 7.73 | 2.66 | 0.9 | 3.79 | 0 | 14.4 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 192 | 8.949 | 8.9837 | 9.4486 | 3.06 | 1.07 | 3.79 | 0 | 14.4 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 106 | 6.39 | 6.16 | 5.87 | 4.24 | 1.4 | 3.79 | 0 | 14.4 | 0 |
| Moisture (%AR) | 500.99 | Other | 102 | 6.6 | 6 | 6.5 | 5.05 | 1.67 | 3.79 | 0 | 14.4 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 196 | 7.13 | 6.43 | 6.47 | 5.89 | 1.96 | 3.79 | 0 | 14.4 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 153 | 3.83 | 3.45 | 3.41 | 6.51 | 1.97 | 3.79 | 0 | 14.4 | 0 |
| Moisture (%AR) | 500.99 | Other | 103 | 4.94 | 4.59 | 5.22 | 6.42 | 2.04 | 3.79 | 0 | 14.4 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 162 | 8.57 | 9.8 | 8.91 | 6.98 | 2.44 | 3.79 | 0 | 14.4 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 182 | 6.125 | 6.745 | 7.16 | 7.8 | 2.6 | 3.79 | 0 | 14.4 | 0 |
| Moisture (%AR) | 500.99 | Other | 120 | 7.3324 | 6.2316 | 6.3837 | 8.97 | 2.99 | 3.79 | 0 | 14.4 | 0 |
| Moisture (%AR) | 500.72 | NIR | 187 | 7.1 | 7.4 | 8.5 | 9.61 | 3.27 | 3.79 | 0 | 14.4 | 0 |
| Moisture (%AR) | 500.99 | Other | 224 | 9.0184 | 6.7916 | 8.5429 | 14.4 | 4.96 | 3.79 | 0 | 14.4 | 0 |
|  |  |  |  | Δ9-THC (%DW) |  |  |  |  |  |  |  |  |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 189 | 0.2261 | 0.2248 | 0.2262 | 0.35 | 0.07 | 3.75 | 0.35 | 16.5 | 0 |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, U | 120 | 0.2135 | 0.2114 | 0.2122 | 0.5 | 0.1 | 3.75 | 0.35 | 16.5 | 0 |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, U | 199 | 0.2215 | 0.2217 | 0.2242 | 0.68 | 0.14 | 3.75 | 0.35 | 16.5 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 204 | 0.2044 | 0.2024 | 0.2012 | 0.8 | 0.16 | 3.75 | 0.35 | 16.5 | 0 |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, U | 103 | 0.1673 | 0.1687 | 0.1611 | 2.44 | 0.47 | 3.75 | 0.35 | 16.5 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 116 | 0.2767 | 0.2715 | 0.2863 | 2.7 | 0.56 | 3.75 | 0.35 | 16.5 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 172 | 0.2065 | 0.215 | 0.2193 | 3.05 | 0.61 | 3.75 | 0.35 | 16.5 | 0 |
| Δ9-THC (%DW) | 501.30 | Other LC-MS | 102 | 0.1988 | 0.2103 | 0.2138 | 3.78 | 0.75 | 3.75 | 0.35 | 16.5 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 162 | 0.18 | 0.1682 | 0.1842 | 4.67 | 0.9 | 3.75 | 0.35 | 16.5 | 0 |
| Δ9-THC (%DW) | 501.30 | Other LC-MS | 105 | 0.198 | 0.216 | 0.197 | 5.25 | 1.04 | 3.75 | 0.35 | 16.5 | 0 |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, U | 218 | 0.176 | 0.198 | 0.172 | 7.69 | 1.49 | 3.75 | 0.35 | 16.5 | 0 |
| Δ9-THC (%DW) | 501.60 | NIR | 187 | 0.13 | 0.15 | 0.15 | 8.06 | 1.51 | 3.75 | 0.35 | 16.5 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%DW) | 501.01 | AOAC 2018.10 | 210 | 0.021 | 0.018 | 0.025 | 16.5 | 2.32 | 3.75 | 0.35 | 16.5 | 0 |
|  |  |  |  | Δ9-THCA (%DW) |  |  |  |  |  |  |  |  |
| Δ9-THCA (%DW) | 502.60 | NIR | 187 | 0.01 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 204 | 0.0681 | 0.069 | 0.0704 | 1.68 | 0.28 | 4.33 | 1.68 | 7.91 | 0 |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, U | 199 | 0.0702 | 0.0712 | 0.0685 | 1.95 | 0.33 | 4.33 | 1.68 | 7.91 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 189 | 0.0993 | 0.1029 | 0.1024 | 1.92 | 0.34 | 4.33 | 1.68 | 7.91 | 0 |
| Δ9-THCA (%DW) | 502.01 | AOAC 2018.10 | 210 | 0.073 | 0.075 | 0.076 | 2.05 | 0.35 | 4.33 | 1.68 | 7.91 | 0 |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, U | 103 | 0.0579 | 0.0587 | 0.055 | 3.4 | 0.56 | 4.33 | 1.68 | 7.91 | 0 |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, U | 218 | 0.096 | 0.101 | 0.095 | 3.3 | 0.58 | 4.33 | 1.68 | 7.91 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 162 | 0.0669 | 0.0666 | 0.0628 | 3.49 | 0.58 | 4.33 | 1.68 | 7.91 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 172 | 0.108 | 0.1102 | 0.1006 | 4.73 | 0.85 | 4.33 | 1.68 | 7.91 | 0 |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, U | 120 | 0.0701 | 0.0608 | 0.0652 | 7.12 | 1.18 | 4.33 | 1.68 | 7.91 | 0 |
| Δ9-THCA (%DW) | 502.30 | Other LC-MS | 102 | 0.0778 | 0.0887 | 0.0873 | 7.01 | 1.21 | 4.33 | 1.68 | 7.91 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 116 | 0.0727 | 0.0731 | 0.0826 | 7.36 | 1.25 | 4.33 | 1.68 | 7.91 | 0 |
| Δ9-THCA (%DW) | 502.30 | Other LC-MS | 105 | 0.0598 | 0.0594 | 0.0518 | 7.91 | 1.29 | 4.33 | 1.68 | 7.91 | 0 |
|  |  |  |  | CBD (%DW) |  |  |  |  |  |  |  |  |
| CBD (%DW) | 503.02 | AOAC 2018.11, U | 199 | 2.0227 | 2.0251 | 2.0239 | 0.06 | 0.02 | 2.59 | 0.06 | 10.1 | 0 |
| CBD (%DW) | 503.01 | AOAC 2018.10 | 210 | 1.951 | 1.956 | 1.957 | 0.16 | 0.05 | 2.59 | 0.06 | 10.1 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 204 | 1.993 | 1.9795 | 1.9804 | 0.38 | 0.11 | 2.59 | 0.06 | 10.1 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 189 | 1.9114 | 1.9067 | 1.9213 | 0.39 | 0.11 | 2.59 | 0.06 | 10.1 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%DW) | 503.10 | Other LC-UV | 162 | 2.1321 | 2.1457 | 2.1599 | 0.65 | 0.18 | 2.59 | 0.06 | 10.1 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 116 | 2.0502 | 2.0838 | 2.0723 | 0.82 | 0.23 | 2.59 | 0.06 | 10.1 | 0 |
| CBD (%DW) | 503.02 | AOAC 2018.11, U | 218 | 1.97 | 1.99 | 1.953 | 0.94 | 0.26 | 2.59 | 0.06 | 10.1 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 172 | 1.9651 | 2.1502 | 2.0218 | 4.64 | 1.29 | 2.59 | 0.06 | 10.1 | 0 |
| CBD (%DW) | 503.02 | AOAC 2018.11, U | 103 | 1.8893 | 1.7241 | 1.7535 | 4.93 | 1.35 | 2.59 | 0.06 | 10.1 | 0 |
| CBD (%DW) | 503.30 | Other LC-MS | 102 | 1.953 | 1.8969 | 2.103 | 5.37 | 1.49 | 2.59 | 0.06 | 10.1 | 0 |
| CBD (%DW) | 503.30 | Other LC-MS | 105 | 1.87 | 1.97 | 1.76 | 5.63 | 1.55 | 2.59 | 0.06 | 10.1 | 0 |
| CBD (%DW) | 503.60 | NIR | 187 | 2.4 | 2.65 | 2.94 | 10.1 | 2.95 | 2.59 | 0.06 | 10.1 | 0 |
|  |  |  |  | CBDA (%DW) |  |  |  |  |  |  |  |  |
| CBDA (%DW) | 504.60 | NIR | 187 | 0.59 | 0.44 | <0.0100 |  |  |  |  |  | 1 |
| CBDA (%DW) | 504.10 | Other LC-UV | 204 | 1.1063 | 1.1098 | 1.1038 | 0.27 | 0.07 | 1.33 | 0.27 | 4.4 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 189 | 1.0092 | 1.0052 | 1.0024 | 0.34 | 0.09 | 1.33 | 0.27 | 4.4 | 0 |
| CBDA (%DW) | 504.02 | AOAC 2018.11, U | 199 | 1.1311 | 1.1324 | 1.1232 | 0.44 | 0.11 | 1.33 | 0.27 | 4.4 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 162 | 1.1633 | 1.1494 | 1.1572 | 0.6 | 0.15 | 1.33 | 0.27 | 4.4 | 0 |
| CBDA (%DW) | 504.01 | AOAC 2018.10 | 210 | 1.223 | 1.209 | 1.22 | 0.61 | 0.16 | 1.33 | 0.27 | 4.4 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 116 | 1.0627 | 1.0857 | 1.0568 | 1.43 | 0.36 | 1.33 | 0.27 | 4.4 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 172 | 1.104 | 1.1179 | 1.1414 | 1.69 | 0.43 | 1.33 | 0.27 | 4.4 | 0 |
| CBDA (%DW) | 504.02 | AOAC 2018.11, U | 218 | 0.898 | 0.93 | 0.94 | 2.38 | 0.59 | 1.33 | 0.27 | 4.4 | 0 |
| CBDA (%DW) | 504.02 | AOAC 2018.11, U | 103 | 1.0068 | 0.9643 | 1.007 | 2.48 | 0.62 | 1.33 | 0.27 | 4.4 | 0 |
| CBDA (%DW) | 504.30 | Other LC-MS | 105 | 0.897 | 0.883 | 0.825 | 4.4 | 1.08 | 1.33 | 0.27 | 4.4 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | CBN (%DW) |  |  |  |  |  |  |  |  |
| CBN (%DW) | 505.10 | Other LC-UV | 189 | 0.0665 | 0.0664 | 0.0663 | 0.15 | 0.03 | 2.92 | 0.15 | 9.11 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 204 | 0.0595 | 0.059 | 0.0598 | 0.68 | 0.11 | 2.92 | 0.15 | 9.11 | 0 |
| CBN (%DW) | 505.30 | GC-FID | 116 | 0.0857 | 0.0847 | 0.0862 | 0.89 | 0.15 | 2.92 | 0.15 | 9.11 | 0 |
| CBN (%DW) | 505.02 | AOAC 2018.11, U | 218 | 0.054 | 0.055 | 0.055 | 1.06 | 0.17 | 2.92 | 0.15 | 9.11 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 192 | 0.0378 | 0.0393 | 0.0397 | 2.57 | 0.4 | 2.92 | 0.15 | 9.11 | 0 |
| CBN (%DW) | 505.02 | AOAC 2018.11, U | 199 | 0.0631 | 0.0613 | 0.0647 | 2.7 | 0.45 | 2.92 | 0.15 | 9.11 | 0 |
| CBN (%DW) | 505.01 | AOAC 2018.10 | 210 | 0.053 | 0.051 | 0.054 | 2.9 | 0.47 | 2.92 | 0.15 | 9.11 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 162 | 0.0605 | 0.0619 | 0.0648 | 3.51 | 0.58 | 2.92 | 0.15 | 9.11 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 172 | 0.0588 | 0.0631 | 0.0599 | 3.69 | 0.61 | 2.92 | 0.15 | 9.11 | 0 |
| CBN (%DW) | 505.20 | Other LC-MS | 105 | 0.0537 | 0.0546 | 0.0506 | 3.96 | 0.64 | 2.92 | 0.15 | 9.11 | 0 |
| CBN (%DW) | 505.20 | Other LC-MS | 102 | 0.0562 | 0.0619 | 0.0654 | 7.59 | 1.25 | 2.92 | 0.15 | 9.11 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 116 | 0.074 | 0.0842 | 0.0709 | 9.11 | 1.55 | 2.92 | 0.15 | 9.11 | 0 |
|  |  |  |  | Total Δ9-THC (%DW) |  |  |  |  |  |  |  |  |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, U | 199 | 0.2831 | 0.2841 | 0.2843 | 0.23 | 0.05 | 3.21 | 0.23 | 6.09 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 189 | 0.3132 | 0.315 | 0.316 | 0.45 | 0.09 | 3.21 | 0.23 | 6.09 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 204 | 0.2641 | 0.2618 | 0.2641 | 0.5 | 0.1 | 3.21 | 0.23 | 6.09 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 172 | 0.3012 | 0.3116 | 0.3075 | 1.71 | 0.36 | 3.21 | 0.23 | 6.09 | 0 |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, U | 120 | 0.275 | 0.2647 | 0.2694 | 1.91 | 0.39 | 3.21 | 0.23 | 6.09 | 0 |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, U | 103 | 0.218 | 0.2202 | 0.2094 | 2.64 | 0.53 | 3.21 | 0.23 | 6.09 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 116 | 0.281 | 0.2959 | 0.2885 | 2.58 | 0.54 | 3.21 | 0.23 | 6.09 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 153 | 0.2908 | 0.2774 | 0.276 | 2.9 | 0.6 | 3.21 | 0.23 | 6.09 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 162 | 0.2387 | 0.2266 | 0.2392 | 3.04 | 0.61 | 3.21 | 0.23 | 6.09 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 138 | 0.2577 | 0.2708 | 0.2726 | 3.05 | 0.63 | 3.21 | 0.23 | 6.09 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 116 | 0.3405 | 0.3357 | 0.3587 | 3.52 | 0.75 | 3.21 | 0.23 | 6.09 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 192 | 0.2105 | 0.1993 | 0.1929 | 4.43 | 0.87 | 3.21 | 0.23 | 6.09 | 0 |
| Total Δ9-THC (%DW) | 506.01 | AOAC 2018.10 | 210 | 0.085 | 0.084 | 0.092 | 5.01 | 0.87 | 3.21 | 0.23 | 6.09 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 182 | 0.23 | 0.2108 | 0.214 | 4.71 | 0.94 | 3.21 | 0.23 | 6.09 | 0 |
| Total Δ9-THC (%DW) | 506.30 | Other LC-MS | 102 | 0.267 | 0.2881 | 0.2904 | 4.58 | 0.95 | 3.21 | 0.23 | 6.09 | 0 |
| Total Δ9-THC (%DW) | 506.60 | NIR | 187 | 0.139 | 0.151 | 0.153 | 5.13 | 0.96 | 3.21 | 0.23 | 6.09 | 0 |
| Total Δ9-THC (%DW) | 506.30 | Other LC-MS | 105 | 0.25 | 0.268 | 0.242 | 5.26 | 1.07 | 3.21 | 0.23 | 6.09 | 0 |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, U | 218 | 0.26 | 0.286 | 0.256 | 6.09 | 1.25 | 3.21 | 0.23 | 6.09 | 0 |
|  |  |  |  | Total CBD (%DW) |  |  |  |  |  |  |  |  |
| Total CBD (%DW) | 507.02 | AOAC 2018.11, U | 199 | 3.0147 | 3.0183 | 3.009 | 0.16 | 0.05 | 2.1 | 0.16 | 6.02 | 0 |
| Total CBD (%DW) | 507.01 | AOAC 2018.10 | 210 | 3.024 | 3.016 | 3.027 | 0.19 | 0.06 | 2.1 | 0.16 | 6.02 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 189 | 2.7965 | 2.7882 | 2.8004 | 0.22 | 0.07 | 2.1 | 0.16 | 6.02 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 204 | 2.9633 | 2.9528 | 2.9484 | 0.26 | 0.08 | 2.1 | 0.16 | 6.02 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 162 | 3.1523 | 3.1537 | 3.1747 | 0.4 | 0.12 | 2.1 | 0.16 | 6.02 | 0 |
| Total CBD (%DW) | 507.02 | AOAC 2018.11, U | 218 | 2.758 | 2.806 | 2.777 | 0.87 | 0.25 | 2.1 | 0.16 | 6.02 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 116 | 2.9822 | 3.036 | 2.9991 | 0.92 | 0.27 | 2.1 | 0.16 | 6.02 | 0 |
| Total CBD (%DW) | 507.60 | NIR | 187 | 2.92 | 3.03 | 2.94 | 1.98 | 0.58 | 2.1 | 0.16 | 6.02 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 192 | 2.6703 | 2.6606 | 2.5181 | 3.26 | 0.94 | 2.1 | 0.16 | 6.02 | 0 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%DW) | 507.10 | Other LC-UV | 172 | 2.9333 | 3.1306 | 3.0228 | 3.26 | 0.97 | 2.1 | 0.16 | 6.02 | 0 |
| Total CBD (%DW) | 507.40 | GC-FID | 116 | 3.1522 | 3.363 | 3.3427 | 3.54 | 1.06 | 2.1 | 0.16 | 6.02 | 0 |
| Total CBD (%DW) | 507.02 | AOAC 2018.11, U | 103 | 2.7722 | 2.5698 | 2.6363 | 3.88 | 1.13 | 2.1 | 0.16 | 6.02 | 0 |
| Total CBD (%DW) | 507.30 | Other LC-MS | 105 | 2.66 | 2.74 | 2.48 | 5.07 | 1.47 | 2.1 | 0.16 | 6.02 | 0 |
| Total CBD (%DW) | 507.40 | GC-FID | 138 | 3.1936 | 3.4796 | 3.5936 | 6.02 | 1.81 | 2.1 | 0.16 | 6.02 | 0 |

Lab

Population of Lab RSDr

CBDV (%DW)

| CBDV (%DW) | 508.99 | Other | 162 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDV (%DW) | 508.02 | AOAC 2018.11, U | 210 | 0.031 | 0.029 | 0.03 | 3.33 | 0.49 | 0 |
| CBDV (%DW) | 508.99 | Other | 189 | 0.0305 | 0.0272 | 0.0293 | 5.76 | 0.85 | 0 |

CBDVA (%DW)

| CBDVA (%DW) | 509.01 | AOAC 2018.10 | 210 | 0.024 | 0.027 | 0.024 | 6.93 | 1 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

CBG (%DW)

| CBG (%DW) | 510.99 | Other | 204 | 0.8951 | 0.8959 | 0.8999 | 0.29 | 0.07 | 1.78 | 0.29 | 4.23 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%DW) | 510.99 | Other | 189 | 0.9271 | 0.9243 | 0.9196 | 0.41 | 0.1 | 1.78 | 0.29 | 4.23 | 0 |
| CBG (%DW) | 510.01 | AOAC 2018.10 | 210 | 0.88 | 0.873 | 0.886 | 0.74 | 0.18 | 1.78 | 0.29 | 4.23 | 0 |
| CBG (%DW) | 510.99 | Other | 105 | 0.808 | 0.796 | 0.784 | 1.51 | 0.37 | 1.78 | 0.29 | 4.23 | 0 |
| CBG (%DW) | 510.99 | Other | 162 | 0.8818 | 0.9174 | 0.8558 | 3.49 | 0.86 | 1.78 | 0.29 | 4.23 | 0 |
| CBG (%DW) | 510.99 | Other | 172 | 0.8879 | 0.9157 | 0.8419 | 4.23 | 1.04 | 1.78 | 0.29 | 4.23 | 0 |

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

RSDr	HorRat(r)

Population of Lab RSDr	 Rob Mean	min	max

Flag

CBGA (%DW)

| CBGA (%DW) | 511.99 | Other | 189 | 1.2771 | 1.2727 | 1.2736 | 0.18 | 0.05 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%DW) | 511.99 | Other | 204 | 1.2633 | 1.2716 | 1.2709 | 0.36 | 0.09 | 0 |
| CBGA (%DW) | 511.01 | AOAC 2018.10 | 210 | 1.404 | 1.393 | 1.39 | 0.53 | 0.14 | 0 |
| CBGA (%DW) | 511.99 | Other | 162 | 1.4879 | 1.5101 | 1.4391 | 2.46 | 0.65 | 0 |
| CBGA (%DW) | 511.99 | Other | 172 | 1.3725 | 1.3939 | 1.2933 | 3.92 | 1.03 | 0 |

THCV (%DW)

| THCV (%DW) | 512.02 | AOAC 2018.11, U | 210 | 0.019 | 0.016 | <0.0130 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| THCV (%DW) | 512.99 | Other | 189 | 0.0603 | 0.0596 | 0.0361 | 26.5 | 4.26 | 0 |

Δ8-THC (%DW)

| Δ8-THC (%DW) | 513.99 | Other | 102 | <0.0211 | <0.0252 | <0.0213 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%DW) | 513.99 | Other | 162 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| Δ8-THC (%DW) | 513.99 | Other | 189 | 0.0089 | 0.0083 | 0.0076 | 7.87 | 0.96 | 0 |
| Δ8-THC (%DW) | 513.01 | AOAC 2018.10 | 210 | 0.021 | 0.018 | 0.025 | 16.5 | 2.32 | 0 |

CBC (%DW)

| CBC (%DW) | 514.99 | Other | 189 | 0.2194 | 0.2195 | 0.2194 | 0.03 | 0.01 | 4.15 | 0.03 | 13.4 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%DW) | 514.01 | AOAC 2018.10 | 210 | 0.181 | 0.18 | 0.18 | 0.32 | 0.06 | 4.15 | 0.03 | 13.4 | 0 |
| CBC (%DW) | 514.99 | Other | 204 | 0.2108 | 0.2073 | 0.2121 | 1.18 | 0.23 | 4.15 | 0.03 | 13.4 | 0 |
| CBC (%DW) | 514.99 | Other | 172 | 0.2065 | 0.2236 | 0.2075 | 4.52 | 0.9 | 4.15 | 0.03 | 13.4 | 0 |

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%DW) | 514.99 | Other | 105 | 0.194 | 0.214 | 0.193 | 5.91 | 1.16 | 4.15 | 0.03 | 13.4 | 0 |
| CBC (%DW) | 514.02 | AOAC 2018.11, U | 103 | 0.1788 | 0.182 | 0.2005 | 6.26 | 1.22 | 4.15 | 0.03 | 13.4 | 0 |
| CBC (%DW) | 514.99 | Other | 162 | 0.1904 | 0.1526 | 0.1976 | 13.4 | 2.6 | 4.15 | 0.03 | 13.4 | 0 |

Lab

Population of Lab RSDr

- Flag definitions: 1 (Lab RSDr not calculated because there were less than 3 numeric lab results)

- Statistical parameters of robust mean, minimum, and maximum for RSDr only shown only if there were 6 or more lab RSDr values.

- AR refers to concentration on as-received basis. DW refers to concentration on dry weight basis.c