# Homogeneity and Stability

Sample ID . . . . . . . . . . . . . . . . . . . . . . . .

Samples shipped . . . . . . . . . . . . . . . . . . .

Results due date . . . . . . . . . . . . . . . . . . . .

HM22NOV-1

8-Nov-22

20-Dec-22

## Homogeneity

10 sample packets randomly chosen for analysis.

| Sample # | Replicate 1 | Replicate 2 |
| --- | --- | --- |
| 1 | 0.2292 | 0.2227 |
| 2 | 0.2355 | 0.2395 |
| 3 | 0.2482 | 0.2615 |
| 4 | 0.2527 | 0.2570 |
| 5 | 0.2442 | 0.2386 |
| 6 | 0.2470 | 0.2520 |
| 7 | 0.2427 | 0.2498 |
| 8 | 0.2579 | 0.2360 |
| 9 | 0.2477 | 0.2587 |
| 10 | 0.2584 | 0.2442 |

Duplicate test portions analyzed for analyte and method shown in table caption.

| Total THC (%AR) via GC-FID (method 006.40) Analysis Date: 11/16/22 | Overall average: | 0.2462 | %RSD | HorRat-r |
| --- | --- | --- | --- | --- |
|  | SD of sample avgs: | 0.009098 | 3.70 |  |
|  | repeatability SD: | 0.007624 | 3.10 | 0.6 |
|  | between-sample SD: | 0.00733 | 2.98 |  |
|  | reproducibilty SD: | 0.01058 | 4.30 |  |
|  | SD used for Analyte z scores: | 0.0388 |  |  |
|  | Check value: | 0.01164 |  |  |
|  | Total THC (%AR) via GC-FID (method 006.40) Is between-sample SD less than check value? YES Homogeneity test passed |  |  |  |
| Total CBD (%AR) via GC-FID (method 007.40) Analysis Date: 11/16/22 | Overall average: | 3.1049 | %RSD | HorRat-r |
|  | SD of sample avgs: | 0.1356 | 4.37 |  |
|  | repeatability SD: | 0.06103 | 1.97 | 0.6 |
|  | between-sample SD: | 0.1286 | 4.14 |  |
|  | reproducibilty SD: | 0.1423 | 4.58 |  |
|  | SD used for Analyte z scores: | 0.2332 |  |  |
|  | Check value: | 0.06996 |  |  |
|  | Total CBD (%AR) via GC-FID (method 007.40) Is between-sample SD less than check value? NO Homogeneity test did not pass |  |  |  |

| Sample # | Replicate 1 | Replicate 2 |
| --- | --- | --- |
| 1 | 2.7657 | 2.8584 |
| 2 | 3.1194 | 3.1344 |
| 3 | 3.0250 | 2.9753 |
| 4 | 3.1731 | 3.0205 |
| 5 | 3.0595 | 3.0385 |
| 6 | 3.2339 | 3.2032 |
| 7 | 3.1295 | 3.0555 |
| 8 | 3.1935 | 3.3643 |
| 9 | 3.1525 | 3.0930 |
| 10 | 3.2421 | 3.2607 |

Stability

Sample tested on different days from date sample shipped to results due date. Sample stored at room temperature during period of testing.

| Total THC (%AR) via GC-FID (method 006.40) |  |  |
| --- | --- | --- |
| Date of Analysis | Days | Conc. |
| 11/7/22 | 0 | 0.2264 |
| 11/16/22 | 9 | 0.2381 |
| 11/17/22 | 10 | 0.2488 |
| 12/2/22 | 25 | 0.2635 |
| 12/12/22 | 35 | 0.2661 |
| 12/16/22 | 39 | 0.2133 |
| 1/12/23 | 66 | 0.2147 |

Days = x variable, Conc = y variable SUMMARY OUTPUT

Regression Statistics

Multiple R   0.3046953 R Square	0.0928392 Adjusted R Squ -0.0885929 Standard Error  0.022702

Observations	7

ANOVA

df	SS	MS	F	Significance F

Regression	1  0.00026372  0.0002637  0.5117  0.506406366

y = -0.0003x + 0.2464 R² = 0.0928

0.30

0.25

0.20

0.15

0.10

0.05

0.00

0

20

40

60

80

Residual	5	0.0025769  0.0005154

Total	6 0.00284062

Coefficients tandard Erro	t Stat	P-value	Lower 95%	Upper 95%

Intercept	0.2464023  0.01376822  17.896454	1E-05  0.211009989  0.281794662

Coefficient for x variable is not statistically significantly different from 0. No signficant change in result during testing period.

Stability test passed.

X Variable 1	-0.000293  0.00040963 -0.7153336  0.5064  -0.001346015  0.000759968

%RSD = 9.1	HorRat-R = 1.8

| Total CBD (%AR) via GC-FID (method 006.40) |  |  |
| --- | --- | --- |
| Date of Analysis | Days | Conc. |
| 11/7/22 | 0 | 2.9610 |
| 11/16/22 | 9 | 3.2262 |
| 11/17/22 | 10 | 3.3319 |
| 12/2/22 | 25 | 3.1862 |
| 12/12/22 | 35 | 3.6524 |
| 12/16/22 | 39 | 3.4914 |
| 1/12/23 | 66 | 3.1779 |

Days = x variable, Conc = y variable SUMMARY OUTPUT

Regression Statistics

Multiple R   0.3396481 R Square	0.1153608 Adjusted R Squ -0.061567 Standard Error  0.233932

Observations	7

| ANOVA |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | df |  | SS | MS | F | Significance F |  |
| Regression |  | 1 | 0.03568137 | 0.0356814 | 0.652 | 0.456070592 |  |
| Residual |  | 5 | 0.27362096 | 0.0547242 |  |  |  |
| Total |  | 6 | 0.30930233 |  |  |  |  |

y = 0.0034x + 3.2 R² = 0.1154

3.8

3.6

3.4

3.2

3.0

2.8

2.6

2.4

2.2

2.0

0

20

40

60

80

Coefficients tandard Erro	t Stat	P-value	Lower 95%	Upper 95%

Intercept	3.1999792  0.14187432  22.555028	3E-06  2.835279659	3.564678741

Coefficient for x variable is not statistically significantly different from 0. No signficant change in result during testing period.

Stability test passed.

X Variable 1	0.0034084  0.00422104  0.8074787  0.4561  -0.007442129	0.014258929

%RSD = 6.9	HorRat-R = 2.1