Hemp Proficiency Testing

All Labs Precision HM22NOV-2 (hemp)

Sample: HM22NOV-2 (hemp)

All Labs Precision

Issue Date:

January 20, 2023

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

Δ9-THC (%AR)

RSDr	HorRat(r)

Population of Lab RSDr Rob Mean	min	max

Flag

| Δ9-THC (%AR) | 001.99 | Other | 183 | 0.213 | 0.211 | 0.213 | 0.54 | 0.11 | 3.29 | 0.54 | 97 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 120 | 0.2098 | 0.2078 | 0.2076 | 0.58 | 0.12 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 196 | 0.1963 | 0.1995 | 0.1958 | 1.02 | 0.2 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 125 | 0.177 | 0.18 | 0.181 | 1.16 | 0.22 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 122 | 0.2051 | 0.2007 | 0.202 | 1.12 | 0.22 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 197 | 0.1931 | 0.1911 | 0.196 | 1.27 | 0.25 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 116 | 0.2325 | 0.2309 | 0.227 | 1.23 | 0.25 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 200 | 0.18 | 0.185 | 0.182 | 1.38 | 0.27 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 175 | 0.2143 | 0.2154 | 0.21 | 1.34 | 0.27 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 199 | 0.213 | 0.212 | 0.2075 | 1.39 | 0.28 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 113 | 0.1722 | 0.178 | 0.1768 | 1.74 | 0.34 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 204 | 0.1914 | 0.1866 | 0.1932 | 1.79 | 0.35 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 194 | 0.2055 | 0.2008 | 0.1979 | 1.9 | 0.38 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 224 | 0.1872 | 0.1793 | 0.1836 | 2.16 | 0.42 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 142 | 0.198 | 0.194 | 0.203 | 2.27 | 0.45 | 3.29 | 0.54 | 97 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 170 | 0.2013 | 0.2089 | 0.2103 | 2.34 | 0.46 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 182 | 0.1575 | 0.1554 | 0.1629 | 2.44 | 0.46 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 161 | 0.1737 | 0.166 | 0.1671 | 2.47 | 0.47 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 144 | 0.1988 | 0.2051 | 0.2091 | 2.54 | 0.5 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 155 | 0.185 | 0.196 | 0.189 | 2.93 | 0.57 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 189 | 0.2164 | 0.2109 | 0.2041 | 2.93 | 0.58 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 106 | 0.2 | 0.19 | 0.19 | 2.99 | 0.58 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 103 | 0.171 | 0.173 | 0.163 | 3.13 | 0.6 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 184 | 0.2423 | 0.2326 | 0.2464 | 2.95 | 0.6 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 225 | 0.1741 | 0.186 | 0.1783 | 3.36 | 0.65 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.03 | AOAC 2018.11, M | 139 | 0.2436 | 0.2289 | 0.2268 | 3.93 | 0.79 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 101 | 0.2082 | 0.1918 | 0.199 | 4.12 | 0.81 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 148 | 0.106 | 0.116 | 0.112 | 4.52 | 0.82 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 115 | 0.25 | 0.23 | 0.24 | 4.17 | 0.84 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 102 | 0.2025 | 0.2028 | 0.1876 | 4.4 | 0.86 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 149 | 0.178 | 0.191 | 0.176 | 4.48 | 0.87 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 192 | 0.1165 | 0.1288 | 0.1185 | 5.44 | 0.99 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.99 | Other | 132 | 0.199 | 0.185 | 0.178 | 5.71 | 1.11 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 191 | 0.2058 | 0.1846 | 0.1882 | 5.88 | 1.15 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.60 | NIR | 187 | 0.23 | 0.24 | 0.21 | 6.74 | 1.35 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 164 | 0.2151 | 0.2481 | 0.216 | 8.3 | 1.67 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 152 | 0.2404 | 0.2375 | 0.2026 | 9.27 | 1.86 | 3.29 | 0.54 | 97 | 0 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 217 | 0.1731 | 0.209 | 0.1598 | 14.1 | 2.73 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 163 | 0.288 | 0.238 | 0.217 | 14.7 | 2.99 | 3.29 | 0.54 | 97 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 172 | 0.91 | 0.188 | 0.19 | 97 | 21.4 | 3.29 | 0.54 | 97 | 0 |

Lab

Population of Lab RSDr

Δ9-THCA (%AR)

| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 196 | <0.0205 | <0.0205 | <0.0205 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 199 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 161 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 170 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.60 | NIR | 187 | <0.0100 | 0.02 | 0.01 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.99 | Other | 132 | <0.1000 | <0.1000 | <0.1000 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 149 | <0.0250 | <0.0250 | 0.026 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 197 | <0.0495 | <0.0495 | <0.0500 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 194 | <0.0270 | <0.0270 | <0.0270 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 191 | <0.0050 | <0.0050 | <0.0050 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 102 | <0.0494 | 0.0331 | <0.0487 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 200 | 0.024 | 0.024 | 0.024 | 0 | 0 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 120 | 0.0177 | 0.0174 | 0.0177 | 0.98 | 0.13 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 122 | 0.0258 | 0.0265 | 0.026 | 1.38 | 0.2 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 192 | 0.0792 | 0.0808 | 0.0817 | 1.57 | 0.27 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 175 | 0.0181 | 0.0174 | 0.018 | 2.12 | 0.29 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 189 | 0.0434 | 0.0434 | 0.0448 | 1.84 | 0.29 | 7.37 | 0 | 25.1 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 144 | 0.0253 | 0.0255 | 0.0263 | 2.06 | 0.3 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 115 | 0.023 | 0.023 | 0.022 | 2.55 | 0.36 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 103 | 0.019 | 0.019 | 0.02 | 2.99 | 0.41 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 225 | 0.019 | 0.0206 | 0.0201 | 4.11 | 0.57 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 125 | 0.0249 | 0.0264 | 0.0274 | 4.8 | 0.7 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 152 | 0.037 | 0.0381 | 0.0343 | 5.36 | 0.82 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 172 | 0.045 | 0.041 | 0.045 | 5.29 | 0.83 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 142 | 0.0183 | 0.0174 | 0.0197 | 6.28 | 0.86 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 101 | 0.0273 | 0.0242 | 0.0261 | 6.04 | 0.88 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 116 | 0.026 | 0.0292 | 0.0253 | 7.75 | 1.13 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 164 | 0.023 | 0.0269 | 0.0234 | 8.78 | 1.26 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.03 | AOAC 2018.11, M | 139 | 0.0241 | 0.0242 | 0.0203 | 9.72 | 1.38 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 217 | 0.0273 | 0.0266 | 0.032 | 10.3 | 1.51 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 224 | 0.285 | 0.2463 | 0.2555 | 7.71 | 1.58 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 184 | 0.0258 | 0.0321 | 0.0309 | 11.3 | 1.67 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 163 | 0.02 | 0.017 | 0.022 | 12.8 | 1.78 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 155 | 0.033 | 0.026 | 0.032 | 12.5 | 1.85 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 182 | 0.0326 | 0.0402 | 0.0325 | 12.6 | 1.91 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 113 | 0.0292 | 0.0222 | 0.0222 | 16.5 | 2.37 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.99 | Other | 183 | 0.043 | 0.028 | 0.035 | 21.2 | 3.22 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 148 | 0.014 | 0.023 | 0.022 | 25.1 | 3.49 | 7.37 | 0 | 25.1 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 106 | 0.02 | 0.02 | 0.03 | 24.7 | 3.53 | 7.37 | 0 | 25.1 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | CBD (%AR) |  |  |  |  |  |  |  |  |
| CBD (%AR) | 003.30 | Other LC-MS | 122 | 5.0295 | 5.0291 | 5.025 | 0.05 | 0.02 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 149 | 4.495 | 4.506 | 4.487 | 0.21 | 0.07 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 199 | 4.811 | 4.796 | 4.821 | 0.26 | 0.08 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 197 | 5.1792 | 5.1413 | 5.184 | 0.45 | 0.15 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 200 | 4.498 | 4.495 | 4.542 | 0.58 | 0.18 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 189 | 4.5717 | 4.5217 | 4.5249 | 0.62 | 0.19 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 161 | 4.3917 | 4.4308 | 4.3753 | 0.65 | 0.2 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 196 | 4.7536 | 4.8173 | 4.7896 | 0.67 | 0.21 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 170 | 5.0786 | 5.0136 | 5.0451 | 0.64 | 0.21 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 172 | 4.688 | 4.634 | 4.63 | 0.7 | 0.22 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.99 | Other | 132 | 4.258 | 4.241 | 4.2 | 0.7 | 0.22 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 142 | 4.64 | 4.61 | 4.71 | 1.1 | 0.35 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.03 | AOAC 2018.11, M | 139 | 4.4588 | 4.3582 | 4.3869 | 1.18 | 0.37 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 192 | 3.2066 | 3.3183 | 3.2478 | 1.73 | 0.52 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 152 | 4.4578 | 4.6102 | 4.5488 | 1.69 | 0.53 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 204 | 4.8451 | 4.7251 | 4.8961 | 1.82 | 0.58 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 194 | 4.9607 | 4.9544 | 4.7681 | 2.24 | 0.71 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 184 | 4.858 | 4.697 | 4.915 | 2.34 | 0.74 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 103 | 4.443 | 4.282 | 4.483 | 2.42 | 0.76 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 116 | 5.1825 | 4.9388 | 5.1238 | 2.5 | 0.8 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 106 | 5.04 | 5 | 4.8 | 2.6 | 0.83 | 2 | 0.05 | 16.3 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%AR) | 003.10 | Other LC-UV | 113 | 4.42 | 4.623 | 4.64 | 2.68 | 0.84 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 163 | 5.211 | 5.491 | 5.423 | 2.72 | 0.88 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 125 | 4.5 | 4.26 | 4.27 | 3.13 | 0.98 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.60 | NIR | 187 | 6.68 | 6.78 | 6.4 | 2.98 | 0.99 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 164 | 4.5144 | 4.733 | 4.4006 | 3.71 | 1.17 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 155 | 4.16 | 4.45 | 4.5 | 4.2 | 1.31 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 144 | 5.4874 | 5.2401 | 5.6911 | 4.13 | 1.33 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 102 | 4.8363 | 4.4241 | 4.6149 | 4.46 | 1.41 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 191 | 5.2364 | 4.6999 | 4.6689 | 6.55 | 2.08 | 2 | 0.05 | 16.3 | 0 |
| CBD (%AR) | 003.99 | Other | 183 | 5.616 | 4.085 | 5.364 | 16.3 | 5.22 | 2 | 0.05 | 16.3 | 0 |
|  |  |  |  | CBDA (%AR) |  |  |  |  |  |  |  |  |
| CBDA (%AR) | 004.30 | Other LC-MS | 122 | 3.065 | 3.0525 | 3.0656 | 0.24 | 0.07 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 116 | 3.0287 | 3.0556 | 3.0426 | 0.44 | 0.13 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 200 | 2.766 | 2.752 | 2.778 | 0.47 | 0.14 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 170 | 3.9737 | 3.9688 | 3.935 | 0.53 | 0.16 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 196 | 2.9019 | 2.9322 | 2.8989 | 0.63 | 0.19 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 189 | 2.8567 | 2.8231 | 2.8224 | 0.69 | 0.2 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 152 | 2.3809 | 2.3921 | 2.3427 | 1.09 | 0.31 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 199 | 3.097 | 3.163 | 3.146 | 1.09 | 0.33 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 197 | 3.0554 | 3.1382 | 3.098 | 1.34 | 0.4 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 204 | 3.1591 | 3.0925 | 3.1724 | 1.36 | 0.41 | 2.22 | 0.24 | 17.2 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDA (%AR) | 004.10 | Other LC-UV | 149 | 2.953 | 2.98 | 2.898 | 1.42 | 0.42 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 163 | 3.077 | 3.005 | 3.109 | 1.74 | 0.52 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 192 | 1.9806 | 2.0449 | 1.9754 | 1.94 | 0.54 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 194 | 3.3131 | 3.2871 | 3.192 | 1.95 | 0.58 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.99 | Other | 132 | 2.557 | 2.483 | 2.459 | 2.04 | 0.59 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 161 | 2.7823 | 2.6649 | 2.7113 | 2.17 | 0.63 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 142 | 2.87 | 2.79 | 2.93 | 2.45 | 0.72 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 125 | 2.81 | 2.94 | 2.94 | 2.59 | 0.76 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 106 | 3.09 | 3.15 | 2.99 | 2.63 | 0.78 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 155 | 2.57 | 2.72 | 2.67 | 2.88 | 0.84 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 103 | 2.951 | 2.794 | 2.822 | 2.93 | 0.86 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 144 | 3.2784 | 3.2611 | 3.4428 | 3.01 | 0.9 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 172 | 3.106 | 2.926 | 3.096 | 3.32 | 0.98 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.03 | AOAC 2018.11, M | 139 | 3.3271 | 3.1058 | 3.228 | 3.44 | 1.03 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 184 | 3.671 | 3.51 | 3.787 | 3.8 | 1.16 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 191 | 3.7046 | 3.2471 | 3.2535 | 7.71 | 2.32 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 164 | 3.3874 | 3.9094 | 3.2443 | 9.96 | 3.01 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.60 | NIR | 187 | 2.53 | 2.79 | 2.14 | 13.2 | 3.78 | 2.22 | 0.24 | 17.2 | 0 |
| CBDA (%AR) | 004.99 | Other | 183 | 3.781 | 2.695 | 3.581 | 17.2 | 5.18 | 2.22 | 0.24 | 17.2 | 0 |
|  |  |  |  | CBN (%AR) |  |  |  |  |  |  |  |  |
| CBN (%AR) | 005.20 | Other LC-MS | 102 | <0.0494 | 0.0448 | <0.0487 |  |  |  |  |  | 1 |

Population of Lab RSDr

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 199 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.99 | Other | 132 | <0.1000 | <0.1000 | <0.1000 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 197 | <0.0743 | <0.0743 | <0.0750 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 106 | 0.05 | 0.05 | 0.05 | 0 | 0 | 6.02 | 0 | 26.3 | 0 |
| CBN (%AR) | 005.20 | Other LC-MS | 122 | 0.055 | 0.0543 | 0.0547 | 0.64 | 0.1 | 6.02 | 0 | 26.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 196 | 0.0386 | 0.0394 | 0.0388 | 1.07 | 0.16 | 6.02 | 0 | 26.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 189 | 0.0458 | 0.0447 | 0.0451 | 1.23 | 0.19 | 6.02 | 0 | 26.3 | 0 |
| CBN (%AR) | 005.99 | Other | 183 | 0.047 | 0.047 | 0.046 | 1.24 | 0.2 | 6.02 | 0 | 26.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 172 | 0.042 | 0.041 | 0.042 | 1.39 | 0.22 | 6.02 | 0 | 26.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 200 | 0.04 | 0.041 | 0.04 | 1.43 | 0.22 | 6.02 | 0 | 26.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 204 | 0.0441 | 0.0428 | 0.0439 | 1.61 | 0.25 | 6.02 | 0 | 26.3 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 125 | 0.0533 | 0.0533 | 0.0548 | 1.61 | 0.26 | 6.02 | 0 | 26.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 194 | 0.0372 | 0.0368 | 0.0357 | 2.12 | 0.32 | 6.02 | 0 | 26.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 149 | 0.038 | 0.041 | 0.039 | 3.88 | 0.6 | 6.02 | 0 | 26.3 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 116 | 0.0428 | 0.0403 | 0.0439 | 4.36 | 0.68 | 6.02 | 0 | 26.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 142 | 0.035 | 0.0348 | 0.0382 | 5.3 | 0.81 | 6.02 | 0 | 26.3 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 192 | 0.0498 | 0.0445 | 0.0436 | 7.29 | 1.15 | 6.02 | 0 | 26.3 | 0 |
| CBN (%AR) | 005.03 | AOAC 2018.11, M | 139 | 0.0388 | 0.0332 | 0.0337 | 8.8 | 1.33 | 6.02 | 0 | 26.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 144 | 0.0369 | 0.0307 | 0.0374 | 10.7 | 1.62 | 6.02 | 0 | 26.3 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 138 | 0.0448 | 0.0557 | 0.0492 | 11 | 1.76 | 6.02 | 0 | 26.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 116 | 0.0588 | 0.0686 | 0.0506 | 15.2 | 2.49 | 6.02 | 0 | 26.3 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 107 | 0.0344 | 0.0457 | 0.0485 | 17.4 | 2.72 | 6.02 | 0 | 26.3 | 0 |

Lab

Lab

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%AR) | 005.30 | GC-FID | 163 | 0.052 | 0.058 | 0.04 | 18.3 | 2.93 | 6.02 | 0 | 26.3 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 192 | 0.0188 | 0.0324 | 0.0278 | 26.3 | 3.82 | 6.02 | 0 | 26.3 | 0 |
|  |  |  |  | Total Δ9-THC (%AR) |  |  |  |  |  |  |  |  |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 190 | 0.273 | 0.28 |  |  |  |  |  |  | 1 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 120 | 0.2253 | 0.223 | 0.2231 | 0.58 | 0.12 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 116 | 0.2344 | 0.2315 | 0.2343 | 0.7 | 0.14 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 122 | 0.2278 | 0.224 | 0.2262 | 0.84 | 0.17 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 169 | 0.1941 | 0.1962 | 0.198 | 1 | 0.2 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 196 | 0.1963 | 0.1995 | 0.1958 | 1.02 | 0.2 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 155 | 0.214 | 0.218 | 0.218 | 1.07 | 0.21 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 175 | 0.2301 | 0.2307 | 0.2258 | 1.17 | 0.23 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 199 | 0.2301 | 0.233 | 0.2277 | 1.15 | 0.23 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 200 | 0.201 | 0.206 | 0.203 | 1.24 | 0.24 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 197 | 0.1931 | 0.1911 | 0.196 | 1.27 | 0.25 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 125 | 0.199 | 0.203 | 0.205 | 1.51 | 0.3 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 182 | 0.1861 | 0.1907 | 0.1914 | 1.52 | 0.3 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 116 | 0.2553 | 0.2565 | 0.2492 | 1.54 | 0.31 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 172 | 0.2305 | 0.224 | 0.2295 | 1.54 | 0.31 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 204 | 0.1914 | 0.1866 | 0.1932 | 1.79 | 0.35 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 194 | 0.2055 | 0.2008 | 0.1979 | 1.9 | 0.38 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 192 | 0.1283 | 0.1231 | 0.1271 | 2.16 | 0.4 | 3.76 | 0.58 | 11.9 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 189 | 0.2545 | 0.249 | 0.2434 | 2.23 | 0.45 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 161 | 0.1737 | 0.166 | 0.1671 | 2.47 | 0.47 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 144 | 0.221 | 0.2275 | 0.2322 | 2.48 | 0.5 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 142 | 0.214 | 0.209 | 0.22 | 2.57 | 0.51 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 102 | 0.2274 | 0.2318 | 0.2191 | 2.85 | 0.57 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.99 | Other | 183 | 0.251 | 0.236 | 0.244 | 3.08 | 0.62 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 225 | 0.1908 | 0.204 | 0.196 | 3.38 | 0.66 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 114 | 0.2079 | 0.2071 | 0.195 | 3.55 | 0.7 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 138 | 0.2082 | 0.2235 | 0.2161 | 3.54 | 0.71 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 192 | 0.1859 | 0.1996 | 0.1901 | 3.66 | 0.72 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 163 | 0.25 | 0.262 | 0.244 | 3.64 | 0.74 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.03 | AOAC 2018.11, M | 139 | 0.2647 | 0.2502 | 0.2446 | 4.1 | 0.84 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 101 | 0.2322 | 0.213 | 0.2219 | 4.32 | 0.86 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 115 | 0.27 | 0.25 | 0.25 | 4.5 | 0.92 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 103 | 0.1877 | 0.1995 | 0.1805 | 5.07 | 0.99 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 128 | 0.2845 | 0.2629 | 0.288 | 4.88 | 1.01 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 149 | 0.178 | 0.191 | 0.198 | 5.37 | 1.05 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 170 | 0.2328 | 0.2595 | 0.2426 | 5.51 | 1.12 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.99 | Other | 132 | 0.219 | 0.203 | 0.195 | 5.94 | 1.17 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.60 | NIR | 138 | 0.38 | 0.395 | 0.425 | 5.73 | 1.25 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 148 | 0.118 | 0.136 | 0.131 | 7.24 | 1.33 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 191 | 0.21 | 0.18 | 0.19 | 7.9 | 1.55 | 3.76 | 0.58 | 11.9 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 164 | 0.2353 | 0.2717 | 0.2365 | 8.34 | 1.7 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.60 | NIR | 187 | 0.23 | 0.26 | 0.22 | 8.8 | 1.78 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 152 | 0.2774 | 0.276 | 0.2326 | 9.72 | 1.99 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 153 | 0.2852 | 0.247 | 0.2326 | 10.7 | 2.18 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 217 | 0.197 | 0.2324 | 0.1878 | 11.4 | 2.26 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 107 | 0.2018 | 0.2475 | 0.2522 | 11.9 | 2.4 | 3.76 | 0.58 | 11.9 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 121 | 0.234 | 0.276 | 0.297 | 11.9 | 2.45 | 3.76 | 0.58 | 11.9 | 0 |
|  |  |  |  | Total CBD (%AR) |  |  |  |  |  |  |  |  |
| Total CBD (%AR) | 007.30 | Other LC-MS | 122 | 7.7175 | 7.7061 | 7.7135 | 0.08 | 0.03 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 197 | 7.8588 | 7.8935 | 7.9009 | 0.28 | 0.1 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 199 | 7.5271 | 7.5345 | 7.58 | 0.38 | 0.13 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 200 | 6.924 | 6.909 | 6.978 | 0.52 | 0.18 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.60 | NIR | 138 | 7.625 | 7.675 | 7.59 | 0.56 | 0.19 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 196 | 7.2985 | 7.3888 | 7.3319 | 0.62 | 0.21 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 161 | 6.8318 | 6.7679 | 6.7531 | 0.62 | 0.21 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 149 | 7.085 | 7.12 | 7.028 | 0.66 | 0.22 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 189 | 7.0771 | 6.9976 | 7.0002 | 0.64 | 0.22 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 125 | 6.96 | 6.84 | 6.86 | 0.93 | 0.31 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.99 | Other | 132 | 6.502 | 6.421 | 6.358 | 1.12 | 0.37 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 169 | 6.8352 | 6.9787 | 6.8116 | 1.32 | 0.44 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 116 | 8.4107 | 8.1832 | 8.3016 | 1.37 | 0.47 | 2.16 | 0.07 | 16.7 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%AR) | 007.10 | Other LC-UV | 172 | 7.412 | 7.2001 | 7.3452 | 1.48 | 0.5 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 116 | 7.8386 | 7.6186 | 7.7922 | 1.5 | 0.51 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 142 | 7.16 | 7.06 | 7.28 | 1.54 | 0.52 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 192 | 4.9436 | 5.1117 | 4.9802 | 1.76 | 0.56 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 204 | 7.6156 | 7.4372 | 7.6783 | 1.65 | 0.56 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 152 | 6.8499 | 7.0967 | 6.8915 | 1.9 | 0.64 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.03 | AOAC 2018.11, M | 139 | 7.3766 | 7.082 | 7.2179 | 2.04 | 0.69 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 194 | 7.8663 | 7.8372 | 7.5675 | 2.12 | 0.72 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 103 | 7.031 | 6.7323 | 6.9579 | 2.25 | 0.75 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 192 | 4.5204 | 4.3021 | 4.3898 | 2.49 | 0.78 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.60 | NIR | 187 | 6.68 | 6.78 | 6.4 | 2.98 | 0.99 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 155 | 6.41 | 6.83 | 6.84 | 3.67 | 1.22 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 144 | 8.3632 | 8.1077 | 8.7111 | 3.61 | 1.24 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 163 | 7.521 | 8.083 | 7.904 | 3.66 | 1.25 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 164 | 7.4851 | 8.1615 | 7.2459 | 6.22 | 2.12 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 138 | 7.702 | 8.5002 | 8.7512 | 6.59 | 2.27 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 191 | 8.49 | 7.55 | 7.52 | 7.02 | 2.4 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 170 | 6.0589 | 7.3848 | 6.8405 | 9.86 | 3.29 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 107 | 5.294 | 6.3365 | 6.3886 | 10.3 | 3.37 | 2.16 | 0.07 | 16.7 | 0 |
| Total CBD (%AR) | 007.99 | Other | 183 | 8.932 | 6.448 | 8.504 | 16.7 | 5.71 | 2.16 | 0.07 | 16.7 | 0 |

Population of Lab RSDr

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

RSDr	HorRat(r)

Population of Lab RSDr	 Rob Mean	min	max

Flag

CBDV (%AR)

| CBDV (%AR) | 008.99 | Other | 132 | <0.1000 | <0.1000 | <0.1000 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDV (%AR) | 008.99 | Other | 197 | <0.0495 | <0.0495 | <0.0500 |  |  |  |  |  | 1 |
| CBDV (%AR) | 008.99 | Other | 194 | <0.0270 | <0.0270 | <0.0270 |  |  |  |  |  | 1 |
| CBDV (%AR) | 008.99 | Other | 200 | 0.041 | 0.041 | 0.041 | 0 | 0 | 8.02 | 0 | 128 | 0 |
| CBDV (%AR) | 008.99 | Other | 196 | 0.0318 | 0.0323 | 0.032 | 0.79 | 0.12 | 8.02 | 0 | 128 | 0 |
| CBDV (%AR) | 008.99 | Other | 183 | 0.046 | 0.047 | 0.046 | 1.25 | 0.2 | 8.02 | 0 | 128 | 0 |
| CBDV (%AR) | 008.99 | Other | 189 | 0.0447 | 0.0518 | 0.0449 | 8.58 | 1.36 | 8.02 | 0 | 128 | 0 |
| CBDV (%AR) | 008.99 | Other | 106 | 0.05 | 0.04 | 0.04 | 13.3 | 2.09 | 8.02 | 0 | 128 | 0 |
| CBDV (%AR) | 008.99 | Other | 142 | 0.0459 | 0.451 | 0.048 | 128 | 24.92 | 8.02 | 0 | 128 | 0 |

CBDVA (%AR)

| CBDVA (%AR) | 009.99 | Other | 197 | <0.0495 | <0.0495 | <0.0500 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDVA (%AR) | 009.99 | Other | 196 | 0.0255 | 0.0262 | 0.0258 | 1.36 | 0.2 | 0 |
| CBDVA (%AR) | 009.99 | Other | 183 | 0.046 | 0.044 | 0.045 | 2.22 | 0.35 | 0 |
| CBDVA (%AR) | 009.99 | Other | 142 | 0.0314 | 0.031 | 0.0331 | 3.5 | 0.52 | 0 |
| CBDVA (%AR) | 009.99 | Other | 149 | 0.043 | 0.034 | 0.034 | 14 | 2.15 | 0 |

CBG (%AR)

| CBG (%AR) | 010.99 | Other | 189 | 0.3616 | 0.3587 | 0.3562 | 0.75 | 0.16 | 2.59 | 0.75 | 5.42 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%AR) | 010.99 | Other | 200 | 0.295 | 0.289 | 0.292 | 1.03 | 0.21 | 2.59 | 0.75 | 5.42 | 0 |
| CBG (%AR) | 010.99 | Other | 142 | 0.295 | 0.297 | 0.303 | 1.4 | 0.29 | 2.59 | 0.75 | 5.42 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%AR) | 010.99 | Other | 196 | 0.2874 | 0.2917 | 0.2961 | 1.49 | 0.31 | 2.59 | 0.75 | 5.42 | 0 |
| CBG (%AR) | 010.99 | Other | 122 | 0.3993 | 0.3883 | 0.3895 | 1.54 | 0.33 | 2.59 | 0.75 | 5.42 | 0 |
| CBG (%AR) | 010.99 | Other | 183 | 0.356 | 0.363 | 0.348 | 2.11 | 0.45 | 2.59 | 0.75 | 5.42 | 0 |
| CBG (%AR) | 010.99 | Other | 132 | 0.337 | 0.328 | 0.323 | 2.15 | 0.46 | 2.59 | 0.75 | 5.42 | 0 |
| CBG (%AR) | 010.02 | AOAC 2018.11, U | 125 | 0.295 | 0.289 | 0.303 | 2.38 | 0.5 | 2.59 | 0.75 | 5.42 | 0 |
| CBG (%AR) | 010.99 | Other | 106 | 0.5 | 0.5 | 0.48 | 2.34 | 0.53 | 2.59 | 0.75 | 5.42 | 0 |
| CBG (%AR) | 010.99 | Other | 172 | 0.29 | 0.305 | 0.301 | 2.6 | 0.54 | 2.59 | 0.75 | 5.42 | 0 |
| CBG (%AR) | 010.99 | Other | 155 | 0.332 | 0.322 | 0.344 | 3.31 | 0.7 | 2.59 | 0.75 | 5.42 | 0 |
| CBG (%AR) | 010.99 | Other | 194 | 0.3418 | 0.3264 | 0.3192 | 3.51 | 0.74 | 2.59 | 0.75 | 5.42 | 0 |
| CBG (%AR) | 010.99 | Other | 204 | 0.3357 | 0.3337 | 0.3614 | 4.5 | 0.96 | 2.59 | 0.75 | 5.42 | 0 |
| CBG (%AR) | 010.99 | Other | 197 | 0.2307 | 0.2569 | 0.244 | 5.37 | 1.09 | 2.59 | 0.75 | 5.42 | 0 |
| CBG (%AR) | 010.99 | Other | 149 | 0.267 | 0.297 | 0.288 | 5.42 | 1.12 | 2.59 | 0.75 | 5.42 | 0 |
|  |  |  |  | CBGA (%AR) |  |  |  |  |  |  |  |  |
| CBGA (%AR) | 011.99 | Other | 122 | 0.4123 | 0.4123 | 0.4121 | 0.03 | 0.01 | 2.42 | 0.03 | 4.42 | 0 |
| CBGA (%AR) | 011.99 | Other | 183 | 0.44 | 0.435 | 0.437 | 0.57 | 0.13 | 2.42 | 0.03 | 4.42 | 0 |
| CBGA (%AR) | 011.99 | Other | 189 | 0.3835 | 0.3774 | 0.3787 | 0.85 | 0.18 | 2.42 | 0.03 | 4.42 | 0 |
| CBGA (%AR) | 011.99 | Other | 194 | 0.407 | 0.4078 | 0.3939 | 1.94 | 0.42 | 2.42 | 0.03 | 4.42 | 0 |
| CBGA (%AR) | 011.99 | Other | 196 | 0.3663 | 0.3658 | 0.3792 | 2.05 | 0.44 | 2.42 | 0.03 | 4.42 | 0 |
| CBGA (%AR) | 011.99 | Other | 172 | 0.376 | 0.382 | 0.365 | 2.3 | 0.5 | 2.42 | 0.03 | 4.42 | 0 |
| CBGA (%AR) | 011.99 | Other | 204 | 0.3982 | 0.3944 | 0.4127 | 2.4 | 0.53 | 2.42 | 0.03 | 4.42 | 0 |
| CBGA (%AR) | 011.99 | Other | 149 | 0.395 | 0.41 | 0.389 | 2.72 | 0.59 | 2.42 | 0.03 | 4.42 | 0 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%AR) | 011.99 | Other | 132 | 0.341 | 0.327 | 0.32 | 3.25 | 0.69 | 2.42 | 0.03 | 4.42 | 0 |
| CBGA (%AR) | 011.99 | Other | 106 | 0.35 | 0.35 | 0.33 | 3.36 | 0.72 | 2.42 | 0.03 | 4.42 | 0 |
| CBGA (%AR) | 011.99 | Other | 142 | 0.368 | 0.368 | 0.39 | 3.38 | 0.73 | 2.42 | 0.03 | 4.42 | 0 |
| CBGA (%AR) | 011.02 | AOAC 2018.11, U | 125 | 0.359 | 0.385 | 0.385 | 3.99 | 0.86 | 2.42 | 0.03 | 4.42 | 0 |
| CBGA (%AR) | 011.99 | Other | 197 | 0.3871 | 0.4228 | 0.408 | 4.42 | 0.97 | 2.42 | 0.03 | 4.42 | 0 |

Lab

Population of Lab RSDr

THCV (%AR)

| THCV (%AR) | 012.99 | Other | 106 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| THCV (%AR) | 012.99 | Other | 142 | <0.0022 | <0.0022 | <0.0022 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 132 | <0.1000 | <0.1000 | <0.1000 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 197 | <0.0495 | <0.0495 | <0.0500 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 183 | 0 | 0 | 0 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 200 | <0.0050 | <0.0050 | <0.0050 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 149 | <0.0250 | <0.0250 | <0.0250 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 194 | <0.0270 | <0.0270 | <0.0270 |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 189 | 0.0459 | 0.0438 | 0.042 | 4.45 | 0.7 | 0 |

Δ8-THC (%AR)

| Δ8-THC (%AR) | 013.99 | Other | 106 | <0.0100 | <0.0100 | <0.0100 | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%AR) | 013.99 | Other | 200 | <0.0100 | <0.0100 | <0.0100 | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 183 | 0 | 0 | 0 | 1 |

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) Rob Mean min max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%AR) | 013.99 | Other | 149 | <0.0250 | <0.0250 | <0.0250 |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 142 | <0.0022 | <0.0022 | <0.0022 |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 197 | <0.0495 | <0.0495 | <0.0500 |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 102 | <0.0198 | <0.0238 | <0.0195 |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 194 | <0.0270 | <0.0270 | <0.0270 |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 138 | 0.0479 | 0.0516 | 0.0503 | 3.76 | 0.6 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 189 | 0.0127 | 0.01 | 0.009 | 18.1 | 2.29 | 0 |

Lab

Population of Lab RSDr

CBC (%AR)

| CBC (%AR) | 014.99 | Other | 196 | 0.2185 | 0.2202 | 0.2172 | 0.69 | 0.14 | 1.6 | 0.69 | 3.5 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%AR) | 014.99 | Other | 155 | 0.225 | 0.228 | 0.225 | 0.77 | 0.15 | 1.6 | 0.69 | 3.5 | 0 |
| CBC (%AR) | 014.99 | Other | 149 | 0.224 | 0.222 | 0.22 | 0.9 | 0.18 | 1.6 | 0.69 | 3.5 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, U | 125 | 0.23 | 0.227 | 0.231 | 0.91 | 0.18 | 1.6 | 0.69 | 3.5 | 0 |
| CBC (%AR) | 014.99 | Other | 189 | 0.2797 | 0.2742 | 0.2782 | 1.03 | 0.21 | 1.6 | 0.69 | 3.5 | 0 |
| CBC (%AR) | 014.99 | Other | 172 | 0.243 | 0.238 | 0.239 | 1.1 | 0.22 | 1.6 | 0.69 | 3.5 | 0 |
| CBC (%AR) | 014.99 | Other | 194 | 0.2421 | 0.2394 | 0.2358 | 1.32 | 0.27 | 1.6 | 0.69 | 3.5 | 0 |
| CBC (%AR) | 014.99 | Other | 204 | 0.2319 | 0.225 | 0.2284 | 1.51 | 0.3 | 1.6 | 0.69 | 3.5 | 0 |
| CBC (%AR) | 014.99 | Other | 183 | 0.265 | 0.258 | 0.257 | 1.68 | 0.34 | 1.6 | 0.69 | 3.5 | 0 |
| CBC (%AR) | 014.99 | Other | 200 | 0.26 | 0.269 | 0.265 | 1.7 | 0.35 | 1.6 | 0.69 | 3.5 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, U | 103 | 0.215 | 0.216 | 0.222 | 1.74 | 0.35 | 1.6 | 0.69 | 3.5 | 0 |
| CBC (%AR) | 014.99 | Other | 106 | 0.24 | 0.24 | 0.23 | 2.44 | 0.49 | 1.6 | 0.69 | 3.5 | 0 |

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%AR) | 014.99 | Other | 142 | 0.219 | 0.218 | 0.228 | 2.48 | 0.5 | 1.6 | 0.69 | 3.5 | 0 |
| CBC (%AR) | 014.99 | Other | 197 | 0.2307 | 0.2188 | 0.233 | 3.35 | 0.67 | 1.6 | 0.69 | 3.5 | 0 |
| CBC (%AR) | 014.99 | Other | 132 | 0.219 | 0.208 | 0.205 | 3.5 | 0.69 | 1.6 | 0.69 | 3.5 | 0 |

Lab

Population of Lab RSDr

Cd (ug/g AR)

| Cd (ug/g AR) | 101.04 | ICP-MS | 224 | <0.2100 | <0.2291 | <0.2274 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cd (ug/g AR) | 101.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 |  |  |  |  |  | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 194 | 0.1552 | 0.1578 | 0.154 | 1.25 | 0.24 | 2.45 | 1.25 | 8.61 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 197 | 0.1636 | 0.1614 | 0.1571 | 2.06 | 0.39 | 2.45 | 1.25 | 8.61 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 196 | 0.139 | 0.135 | 0.133 | 2.25 | 0.42 | 2.45 | 1.25 | 8.61 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 183 | 0.111 | 0.112 | 0.116 | 2.34 | 0.42 | 2.45 | 1.25 | 8.61 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 102 | 0.1517 | 0.1591 | 0.1546 | 2.4 | 0.46 | 2.45 | 1.25 | 8.61 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 106 | 0.161 | 0.169 | 0.16 | 3.02 | 0.58 | 2.45 | 1.25 | 8.61 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 142 | 0.149 | 0.131 | 0.155 | 8.61 | 1.62 | 2.45 | 1.25 | 8.61 | 0 |

As (ug/g AR)

| As (ug/g AR) | 102.04 | ICP-MS | 197 | <0.0952 | <0.0954 | <0.0980 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| As (ug/g AR) | 102.04 | ICP-MS | 142 | <0.0500 | 0.063 | 0.053 |  |  | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 106 | <0.0800 | <0.0800 | <0.0800 |  |  | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 224 | <0.1050 | <0.1145 | <0.1137 |  |  | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 |  |  | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 194 | 0.0649 | 0.0632 | 0.0635 | 1.42 | 0.24 | 0 |

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) Rob Mean min max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| As (ug/g AR) | 102.04 | ICP-MS | 196 | 0.0533 | 0.0515 | 0.0543 | 2.68 | 0.43 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 183 | 0.053 | 0.051 | 0.05 | 2.98 | 0.48 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 102 | 0.0597 | 0.0451 | 0.0434 | 18.1 | 2.89 | 0 |

Lab

Population of Lab RSDr

Hg (ug/g AR)

| Hg (ug/g AR) | 103.04 | ICP-MS | 197 | <0.0952 | <0.0954 | <0.0980 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Hg (ug/g AR) | 103.04 | ICP-MS | 142 | <0.0500 | <0.0500 | <0.0500 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 224 | <0.0210 | <0.0229 | <0.0227 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 196 | <0.0072 | <0.0072 | <0.0072 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 194 | <0.0200 | <0.0200 | <0.0200 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 106 | 0.008 | 0.008 | 0.008 | 0 | 0 | 0 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 102 | 0.0087 | 0.0083 | 0.0087 | 2.7 | 0.33 | 0 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 183 | 0.038 | 0.035 | 0.034 | 5.84 | 0.89 | 0 |

Pb (ug/g AR)

| Pb (ug/g AR) | 104.04 | ICP-MS | 196 | 0.336 | 0.341 | 0.316 | 4 | 0.85 | 11.94 | 4 | 17.8 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Pb (ug/g AR) | 104.04 | ICP-MS | 194 | 0.2431 | 0.2523 | 0.2236 | 6.11 | 1.24 | 11.94 | 4 | 17.8 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 197 | 0.2991 | 0.3363 | 0.277 | 9.85 | 2.07 | 11.94 | 4 | 17.8 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 106 | 0.302 | 0.339 | 0.276 | 10.4 | 2.17 | 11.94 | 4 | 17.8 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 102 | 0.3891 | 0.3621 | 0.2957 | 13.8 | 2.95 | 11.94 | 4 | 17.8 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 224 | 0.4149 | 0.3124 | 0.3823 | 14.2 | 3.06 | 11.94 | 4 | 17.8 | 0 |

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Pb (ug/g AR) | 104.04 | ICP-MS | 162 | 0.4156 | 0.3112 | 0.3647 | 14.3 | 3.09 | 11.94 | 4 | 17.8 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 142 | 0.264 | 0.187 | 0.228 | 17 | 3.41 | 11.94 | 4 | 17.8 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 183 | 0.324 | 0.256 | 0.231 | 17.8 | 3.67 | 11.94 | 4 | 17.8 | 0 |

Lab

Population of Lab RSDr

Cu (ug/g AR)

| Cu (ug/g AR) | 105.04 | ICP-MS | 196 | 8.52 | 8.41 | 8.38 | 0.87 | 0.3 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cu (ug/g AR) | 105.04 | ICP-MS | 142 | 14 | 14.2 | 14.2 | 0.82 | 0.3 | 0 |

Ni (ug/g AR)

| Ni (ug/g AR) | 106.04 | ICP-MS | 196 | 1.06 | 1.03 | 0.999 | 2.96 | 0.75 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Ni (ug/g AR) | 106.04 | ICP-MS | 102 | 2.696 | 2.6176 | 2.485 | 4.1 | 1.19 | 0 |
| Ni (ug/g AR) | 106.04 | ICP-MS | 142 | 0.673 | 0.573 | 0.579 | 9.22 | 2.14 | 0 |

Cr (ug/g AR)

| Cr (ug/g AR) | 107.04 | ICP-MS | 196 | 0.482 | 0.48 | 0.464 | 2.08 | 0.47 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cr (ug/g AR) | 107.04 | ICP-MS | 197 | 0.3598 | 0.3707 | 0.3249 | 6.8 | 1.46 | 0 |
| Cr (ug/g AR) | 107.04 | ICP-MS | 142 | 0.537 | 0.357 | 0.455 | 20 | 4.45 | 0 |

Se (ug/g AR)

| Se (ug/g AR) | 108.04 | ICP-MS | 197 | <1.0000 | <0.0954 | <0.0980 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Se (ug/g AR) | 108.04 | ICP-MS | 142 | <0.0500 | <0.0500 | <0.0500 |  |  | 1 |
| Se (ug/g AR) | 108.04 | ICP-MS | 196 | 0.124 | 0.126 | 0.119 | 2.93 | 0.54 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Se (ug/g AR) | 108.04 | ICP-MS | 102 | 0.1186 | 0.1231 | 0.114 | 3.84 | 0.7 |  |  |  | 0 |
|  |  |  |  | Ag (ug/g AR) |  |  |  |  |  |  |  |  |
| Ag (ug/g AR) | 109.04 | ICP-MS | 142 | <0.0500 <0.0500 |  | <0.0500 |  |  |  |  |  | 1 |
| Ag (ug/g AR) | 109.04 | ICP-MS | 102 | 0.0041 | 0.0044 | 0.0039 | 6.09 | 0.67 |  |  |  | 0 |
|  |  |  |  | Ba (ug/g AR) |  |  |  |  |  |  |  |  |
| Ba (ug/g AR) | 110.04 | ICP-MS | 142 | 50.5 | 52.2 | 49.1 | 3.07 | 1.39 |  |  |  | 0 |
|  |  |  |  | Zn (ug/g AR) |  |  |  |  |  |  |  |  |
| Zn (ug/g AR) | 111.04 | ICP-MS | 196 | 52.2 | 51.9 | 52.6 | 0.67 | 0.3 |  |  |  | 0 |
|  |  |  |  | Sb (ug/g AR) |  |  |  |  |  |  |  |  |
| Sb (ug/g AR) | 112.04 | ICP-MS | 196 | <0.0241 <0.0241 |  | <0.0241 |  |  |  |  |  | 1 |
| Sb (ug/g AR) | 112.04 | ICP-MS | 102 | 0.0086 | 0.0093 | 0.0078 | 8.76 | 1.08 |  |  |  | 0 |
|  |  |  |  | (-)-alpha-Bisabolol (%AR) |  |  |  |  |  |  |  |  |
| (-)-alpha-Bisabolol (%A | 201.01 | GC,FID | 142 | 0.106 | 0.104 | 0.106 | 1.1 | 0.2 |  |  |  | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 162 | 0.0317 | 0.0328 | 0.0303 | 3.97 | 0.59 |  |  |  | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 194 | 0.0592 | 0.0595 | 0.0554 | 3.94 | 0.64 |  |  |  | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 183 | 0.095 | 0.105 | 0.096 | 5.58 | 0.99 |  |  |  | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 196 | 0.0205 | 0.0208 | 0.0177 | 8.69 | 1.21 |  |  |  | 0 |

Population of Lab RSDr

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

RSDr	HorRat(r)

Population of Lab RSDr	 Rob Mean	min	max

Flag

beta-Caryophyllene (%AR)

| beta-Caryophyllene (% | 202.01 | GC,FID | 142 | 0.0719 | 0.0708 | 0.072 | 0.93 | 0.16 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| beta-Caryophyllene (% | 202.02 | GC,MS | 194 | 0.047 | 0.0483 | 0.0446 | 4.03 | 0.64 | 0 |
| beta-Caryophyllene (% | 202.02 | GC,MS | 183 | 0.056 | 0.061 | 0.059 | 4.29 | 0.7 | 0 |
| beta-Caryophyllene (% | 202.02 | GC,MS | 196 | 0.0384 | 0.0383 | 0.0336 | 7.46 | 1.14 | 0 |
| beta-Caryophyllene (% | 202.02 | GC,MS | 162 | 0.0472 | 0.0546 | 0.0486 | 7.84 | 1.25 | 0 |

(-)-Caryophyllene oxide (%AR)

| (-)-Caryophyllene oxide | 203.01 | GC,FID | 142 | 0.0169 | 0.0163 | 0.0164 | 1.94 | 0.26 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| (-)-Caryophyllene oxide | 203.02 | GC,MS | 194 | 0.0124 | 0.0121 | 0.0119 | 2.07 | 0.27 | 0 |
| (-)-Caryophyllene oxide | 203.02 | GC,MS | 196 | 0.0062 | 0.0061 | 0.0055 | 6.38 | 0.74 | 0 |
| (-)-Caryophyllene oxide | 203.02 | GC,MS | 183 | 0.008 | 0.013 | 0.012 | 24.1 | 3.06 | 0 |

Geraniol (%AR)

| Geraniol (%AR) | 204.02 | GC,MS | 183 | <0.0100 | <0.0100 | <0.0100 | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Geraniol (%AR) | 204.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | 1 |
| Geraniol (%AR) | 204.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 | 1 |
| Geraniol (%AR) | 204.01 | GC,FID | 142 | <0.0010 | <0.0010 | <0.0010 | 1 |
| Geraniol (%AR) | 204.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | 1 |

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

RSDr	HorRat(r)

Population of Lab RSDr	 Rob Mean	min	max

Flag

alpha-Humulene (%AR)

| alpha-Humulene (%AR) | 205.01 | GC,FID | 142 | 0.0377 | 0.0372 | 0.038 | 1.07 | 0.16 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 194 | 0.0194 | 0.0194 | 0.0178 | 4.9 | 0.68 | 0 |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 196 | 0.0155 | 0.0156 | 0.0135 | 7.97 | 1.06 | 0 |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 162 | 0.0188 | 0.0213 | 0.0184 | 8.06 | 1.12 | 0 |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 183 | 0.021 | 0.025 | 0.023 | 8.7 | 1.24 | 0 |

d-Limonene (%AR)

| d-Limonene (%AR) | 206.02 | GC,MS | 183 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| d-Limonene (%AR) | 206.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| d-Limonene (%AR) | 206.01 | GC,FID | 142 | <0.0010 | <0.0010 | <0.0010 |  |  | 1 |
| d-Limonene (%AR) | 206.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 |  |  | 1 |
| d-Limonene (%AR) | 206.02 | GC,MS | 194 | 0.0023 | 0.0024 | 0.0023 | 2.47 | 0.25 | 0 |

Linalool (%AR)

| Linalool (%AR) | 207.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Linalool (%AR) | 207.02 | GC,MS | 183 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| Linalool (%AR) | 207.01 | GC,FID | 142 | <0.0010 | <0.0010 | <0.0010 |  |  | 1 |
| Linalool (%AR) | 207.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 |  |  | 1 |
| Linalool (%AR) | 207.02 | GC,MS | 194 | 0.0056 | 0.0056 | 0.0055 | 1.04 | 0.12 | 0 |

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

RSDr	HorRat(r)

Population of Lab RSDr	 Rob Mean	min	max

Flag

beta-Myrcene (%AR)

| beta-Myrcene (%AR) | 208.01 | GC,FID | 142 | <0.0010 | <0.0010 | <0.0010 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 183 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 194 | 0.0035 | 0.0035 | 0.0033 | 3.36 | 0.36 | 0 |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 196 | 0.0015 | 0.0014 | 0.0013 | 7.14 | 0.67 | 0 |

Nopinene (%AR)

Nopinene (%AR)

209.01

GC,FID

142

<0.0010	<0.0010	<0.0010	1

Ocimene (%AR)

| Ocimene (%AR) | 210.01 | GC,FID | 142 | <0.0010 | <0.0010 | <0.0010 | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Ocimene (%AR) | 210.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | 1 |
| Ocimene (%AR) | 210.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | 1 |
| Ocimene (%AR) | 210.02 | GC,MS | 183 | <0.0100 | <0.0100 | <0.0100 | 1 |
| Ocimene (%AR) | 210.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 | 1 |

alpha-Pinene (%AR)

| alpha-Pinene (%AR) | 211.02 | GC,MS | 183 | <0.0100 | <0.0100 | <0.0100 | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | 1 |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | 1 |
| alpha-Pinene (%AR) | 211.01 | GC,FID | 142 | <0.0010 | <0.0010 | <0.0010 | 1 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 194 | 0.0015 | 0.0016 | 0.0015 | 3.77 | 0.36 |  |  |  | 0 |
|  |  |  |  | (-)-beta-Pinene (%AR) |  |  |  |  |  |  |  |  |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 |  |  |  |  |  | 1 |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 183 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 |  |  |  |  |  | 1 |
|  |  |  |  | alpha-Terpinene (%AR) |  |  |  |  |  |  |  |  |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 183 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 |  |  |  |  |  | 1 |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 |  |  |  |  |  | 1 |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| alpha-Terpinene (%AR) | 213.01 | GC,FID | 142 | <0.0010 | <0.0010 | <0.0010 |  |  |  |  |  | 1 |
|  |  |  |  | gamma-Terpinene (%AR) |  |  |  |  |  |  |  |  |
| gamma-Terpinene (%A | 214.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 |  |  |  |  |  | 1 |
| gamma-Terpinene (%A | 214.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| gamma-Terpinene (%A | 214.02 | GC,MS | 183 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  |  | 1 |
| gamma-Terpinene (%A | 214.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 |  |  |  |  |  | 1 |

Population of Lab RSDr

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

RSDr	HorRat(r)

Population of Lab RSDr	 Rob Mean	min	max

Flag

alpha-Terpinolene (%AR)

| alpha-Terpinolene (%A | 215.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | 1 |
| alpha-Terpinolene (%A | 215.01 | GC,FID | 142 | <0.0010 | <0.0010 | <0.0010 | 1 |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 | 1 |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 183 | <0.0100 | <0.0100 | <0.0100 | 1 |

Moisture (%AR)

| Moisture (%AR) | 500.70 | Commerc. Anal. | 189 | 4.712 | 4.712 | 4.712 | 0 | 0 | 3.41 | 0 | 18.4 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 204 | 6.05 | 6.05 | 6.05 | 0 | 0 | 3.41 | 0 | 18.4 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 183 | 7.2 | 7.1 | 7.15 | 0.7 | 0.24 | 3.41 | 0 | 18.4 | 0 |
| Moisture (%AR) | 500.99 | Other | 105 | 8.63 | 8.5 | 8.55 | 0.77 | 0.26 | 3.41 | 0 | 18.4 | 0 |
| Moisture (%AR) | 500.99 | Other | 142 | 5 | 5 | 5.1 | 1.15 | 0.37 | 3.41 | 0 | 18.4 | 0 |
| Moisture (%AR) | 500.72 | NIR | 138 | 6.9 | 6.75 | 6.8 | 1.12 | 0.37 | 3.41 | 0 | 18.4 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 106 | 6.13 | 6.19 | 6.33 | 1.65 | 0.54 | 3.41 | 0 | 18.4 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 172 | 6.986 | 7.001 | 6.778 | 1.8 | 0.6 | 3.41 | 0 | 18.4 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 116 | 5.48 | 5.34 | 5.26 | 2.08 | 0.67 | 3.41 | 0 | 18.4 | 0 |
| Moisture (%AR) | 500.99 | Other | 120 | 6.6188 | 6.4026 | 6.6553 | 2.08 | 0.69 | 3.41 | 0 | 18.4 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 199 | 9.7294 | 9.405 | 9.8214 | 2.27 | 0.8 | 3.41 | 0 | 18.4 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 210 | 7.13 | 7.46 | 7.09 | 2.81 | 0.95 | 3.41 | 0 | 18.4 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 196 | 6.62 | 7.05 | 6.82 | 3.15 | 1.05 | 3.41 | 0 | 18.4 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 153 | 3.67 | 3.46 | 3.37 | 4.4 | 1.33 | 3.41 | 0 | 18.4 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 194 | 7.26 | 6.73 | 6.82 | 4.09 | 1.37 | 3.41 | 0 | 18.4 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 192 | 9.0646 | 9.7532 | 9.7308 | 4.11 | 1.44 | 3.41 | 0 | 18.4 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 182 | 7.345 | 6.95 | 6.69 | 4.71 | 1.58 | 3.41 | 0 | 18.4 | 0 |
| Moisture (%AR) | 500.99 | Other | 102 | 7 | 6.2 | 6.6 | 6.06 | 2.02 | 3.41 | 0 | 18.4 | 0 |
| Moisture (%AR) | 500.99 | Other | 103 | 4.94 | 4.95 | 5.73 | 8.71 | 2.79 | 3.41 | 0 | 18.4 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 162 | 9.9 | 7.84 | 7.92 | 13.6 | 4.72 | 3.41 | 0 | 18.4 | 0 |
| Moisture (%AR) | 500.99 | Other | 224 | 10.1977 | 7.8909 | 9.9607 | 13.6 | 4.76 | 3.41 | 0 | 18.4 | 0 |
| Moisture (%AR) | 500.72 | NIR | 187 | 5.5 | 5.9 | 7.7 | 18.4 | 6.09 | 3.41 | 0 | 18.4 | 0 |
|  |  |  |  | Δ9-THC (%DW) |  |  |  |  |  |  |  |  |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, U | 120 | 0.2247 | 0.2215 | 0.2224 | 0.74 | 0.15 | 3 | 0.74 | 6.72 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 172 | 0.2052 | 0.202 | 0.2041 | 0.8 | 0.16 | 3 | 0.74 | 6.72 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 116 | 0.2457 | 0.244 | 0.2399 | 1.23 | 0.25 | 3 | 0.74 | 6.72 | 0 |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, U | 199 | 0.236 | 0.234 | 0.2301 | 1.29 | 0.26 | 3 | 0.74 | 6.72 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 204 | 0.2038 | 0.1987 | 0.2057 | 1.79 | 0.35 | 3 | 0.74 | 6.72 | 0 |
| Δ9-THC (%DW) | 501.01 | AOAC 2018.10 | 210 | 0.192 | 0.185 | 0.193 | 2.29 | 0.45 | 3 | 0.74 | 6.72 | 0 |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, U | 103 | 0.18 | 0.182 | 0.1724 | 2.84 | 0.55 | 3 | 0.74 | 6.72 | 0 |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, U | 218 | 0.192 | 0.201 | 0.191 | 2.83 | 0.55 | 3 | 0.74 | 6.72 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 189 | 0.2271 | 0.2214 | 0.2142 | 2.93 | 0.58 | 3 | 0.74 | 6.72 | 0 |
| Δ9-THC (%DW) | 501.30 | Other LC-MS | 102 | 0.2177 | 0.2162 | 0.2009 | 4.39 | 0.87 | 3 | 0.74 | 6.72 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 162 | 0.1831 | 0.1676 | 0.1673 | 5.23 | 1.01 | 3 | 0.74 | 6.72 | 0 |
| Δ9-THC (%DW) | 501.60 | NIR | 187 | 0.24 | 0.26 | 0.23 | 6.28 | 1.27 | 3 | 0.74 | 6.72 | 0 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%DW) | 501.30 | Other LC-MS | 105 | 0.254 | 0.239 | 0.222 | 6.72 | 1.36 | 3 | 0.74 | 6.72 | 0 |

Lab

Population of Lab RSDr

Δ9-THCA (%DW)

| Δ9-THCA (%DW) | 502.60 | NIR | 187 | <0.0100 | 0.02 | 0.01 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%DW) | 502.30 | Other LC-MS | 102 | <0.0531 | 0.353 | <0.0521 |  |  |  |  |  | 1 |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, U | 199 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, U | 218 | 0.057 | 0.057 | 0.056 | 1.02 | 0.17 | 5.05 | 1.02 | 12.2 | 0 |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, U | 120 | 0.0189 | 0.0185 | 0.019 | 1.41 | 0.19 | 5.05 | 1.02 | 12.2 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 189 | 0.0455 | 0.0455 | 0.047 | 1.88 | 0.3 | 5.05 | 1.02 | 12.2 | 0 |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, U | 103 | 0.02 | 0.02 | 0.0212 | 3.4 | 0.47 | 5.05 | 1.02 | 12.2 | 0 |
| Δ9-THCA (%DW) | 502.01 | AOAC 2018.10 | 210 | 0.027 | 0.025 | 0.025 | 4.5 | 0.65 | 5.05 | 1.02 | 12.2 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 172 | 0.0483 | 0.044 | 0.0483 | 5.3 | 0.84 | 5.05 | 1.02 | 12.2 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 116 | 0.0275 | 0.0308 | 0.0267 | 7.67 | 1.13 | 5.05 | 1.02 | 12.2 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 162 | 0.0166 | 0.0197 | 0.0173 | 9.1 | 1.25 | 5.05 | 1.02 | 12.2 | 0 |
| Δ9-THCA (%DW) | 502.30 | Other LC-MS | 105 | 0.0195 | 0.0153 | 0.0182 | 12.2 | 1.66 | 5.05 | 1.02 | 12.2 | 0 |

CBD (%DW)

| CBD (%DW) | 503.01 | AOAC 2018.10 | 210 | 5.053 | 5.028 | 5.035 | 0.26 | 0.08 | 2.12 | 0.26 | 4.86 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%DW) | 503.10 | Other LC-UV | 189 | 4.7978 | 4.7453 | 4.7487 | 0.62 | 0.2 | 2.12 | 0.26 | 4.86 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 172 | 5.0366 | 4.9786 | 4.9743 | 0.7 | 0.22 | 2.12 | 0.26 | 4.86 | 0 |
| CBD (%DW) | 503.02 | AOAC 2018.11, U | 199 | 5.3295 | 5.2547 | 5.3461 | 0.92 | 0.3 | 2.12 | 0.26 | 4.86 | 0 |
| CBD (%DW) | 503.02 | AOAC 2018.11, U | 218 | 4.92 | 4.956 | 4.853 | 1.06 | 0.34 | 2.12 | 0.26 | 4.86 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%DW) | 503.10 | Other LC-UV | 204 | 5.1593 | 5.0315 | 5.2136 | 1.82 | 0.58 | 2.12 | 0.26 | 4.86 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 116 | 5.476 | 5.2185 | 5.414 | 2.5 | 0.81 | 2.12 | 0.26 | 4.86 | 0 |
| CBD (%DW) | 503.02 | AOAC 2018.11, U | 103 | 4.6739 | 4.505 | 4.7555 | 2.75 | 0.87 | 2.12 | 0.26 | 4.86 | 0 |
| CBD (%DW) | 503.60 | NIR | 187 | 4.72 | 4.6 | 4.9 | 3.19 | 1.01 | 2.12 | 0.26 | 4.86 | 0 |
| CBD (%DW) | 503.30 | Other LC-MS | 105 | 4.99 | 4.75 | 5.09 | 3.53 | 1.13 | 2.12 | 0.26 | 4.86 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 162 | 5.9084 | 5.5424 | 5.5811 | 3.54 | 1.15 | 2.12 | 0.26 | 4.86 | 0 |
| CBD (%DW) | 503.30 | Other LC-MS | 102 | 5.1983 | 4.717 | 4.9412 | 4.86 | 1.55 | 2.12 | 0.26 | 4.86 | 0 |
|  |  |  |  | CBDA (%DW) |  |  |  |  |  |  |  |  |
| CBDA (%DW) | 504.10 | Other LC-UV | 116 | 3.2002 | 3.2287 | 3.2149 | 0.44 | 0.13 | 2.7 | 0.44 | 12.1 | 0 |
| CBDA (%DW) | 504.01 | AOAC 2018.10 | 210 | 3.676 | 3.638 | 3.642 | 0.57 | 0.17 | 2.7 | 0.44 | 12.1 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 189 | 2.998 | 2.9627 | 2.962 | 0.69 | 0.2 | 2.7 | 0.44 | 12.1 | 0 |
| CBDA (%DW) | 504.02 | AOAC 2018.11, U | 199 | 3.4308 | 3.4914 | 3.4886 | 0.99 | 0.3 | 2.7 | 0.44 | 12.1 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 204 | 3.364 | 3.293 | 3.3781 | 1.36 | 0.41 | 2.7 | 0.44 | 12.1 | 0 |
| CBDA (%DW) | 504.02 | AOAC 2018.11, U | 218 | 3.071 | 3.053 | 2.957 | 2.02 | 0.6 | 2.7 | 0.44 | 12.1 | 0 |
| CBDA (%DW) | 504.02 | AOAC 2018.11, U | 103 | 3.1044 | 2.94 | 2.9935 | 2.78 | 0.82 | 2.7 | 0.44 | 12.1 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 172 | 3.337 | 3.1436 | 3.3262 | 3.32 | 1 | 2.7 | 0.44 | 12.1 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 162 | 3.8578 | 3.573 | 3.5662 | 4.54 | 1.38 | 2.7 | 0.44 | 12.1 | 0 |
| CBDA (%DW) | 504.30 | Other LC-MS | 105 | 2.77 | 3.06 | 2.72 | 6.44 | 1.89 | 2.7 | 0.44 | 12.1 | 0 |
| CBDA (%DW) | 504.60 | NIR | 187 | 2.68 | 2.96 | 2.32 | 12.1 | 3.51 | 2.7 | 0.44 | 12.1 | 0 |

Population of Lab RSDr

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

RSDr	HorRat(r)

Population of Lab RSDr	 Rob Mean	min	max

Flag

CBN (%DW)

| CBN (%DW) | 505.20 | Other LC-MS | 102 | <0.0531 | 0.0478 | <0.0521 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%DW) | 505.02 | AOAC 2018.11, U | 199 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| CBN (%DW) | 505.10 | Other LC-UV | 189 | 0.048 | 0.0469 | 0.0474 | 1.16 | 0.18 | 3.27 | 1.16 | 26.6 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 162 | 0.0493 | 0.0493 | 0.0482 | 1.3 | 0.21 | 3.27 | 1.16 | 26.6 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 172 | 0.0451 | 0.044 | 0.0451 | 1.42 | 0.22 | 3.27 | 1.16 | 26.6 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 204 | 0.047 | 0.0456 | 0.0467 | 1.59 | 0.25 | 3.27 | 1.16 | 26.6 | 0 |
| CBN (%DW) | 505.02 | AOAC 2018.11, U | 218 | 0.044 | 0.043 | 0.042 | 2.33 | 0.36 | 3.27 | 1.16 | 26.6 | 0 |
| CBN (%DW) | 505.03 | AOAC 2018.11, M | 210 | 0.039 | 0.039 | 0.037 | 3.01 | 0.46 | 3.27 | 1.16 | 26.6 | 0 |
| CBN (%DW) | 505.20 | Other LC-MS | 105 | 0.0402 | 0.0401 | 0.0379 | 3.3 | 0.51 | 3.27 | 1.16 | 26.6 | 0 |
| CBN (%DW) | 505.30 | GC-FID | 116 | 0.0452 | 0.0426 | 0.0464 | 4.34 | 0.68 | 3.27 | 1.16 | 26.6 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 116 | 0.0621 | 0.0725 | 0.0535 | 15.2 | 2.51 | 3.27 | 1.16 | 26.6 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 192 | 0.0207 | 0.0359 | 0.0308 | 26.6 | 3.92 | 3.27 | 1.16 | 26.6 | 0 |

Total Δ9-THC (%DW)

| Total Δ9-THC (%DW) | 506.30 | Other LC-MS | 102 | 0.2444 | 0.2472 | 0.247 | 0.63 | 0.13 | 2.78 | 0.63 | 10.8 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 116 | 0.2477 | 0.2447 | 0.2475 | 0.68 | 0.14 | 2.78 | 0.63 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, U | 120 | 0.2413 | 0.2377 | 0.239 | 0.76 | 0.15 | 2.78 | 0.63 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, U | 199 | 0.2549 | 0.2572 | 0.2525 | 0.92 | 0.19 | 2.78 | 0.63 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 172 | 0.2476 | 0.2406 | 0.2465 | 1.54 | 0.31 | 2.78 | 0.63 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 116 | 0.2698 | 0.2711 | 0.2633 | 1.56 | 0.32 | 2.78 | 0.63 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 204 | 0.2038 | 0.1987 | 0.2057 | 1.79 | 0.35 | 2.78 | 0.63 | 10.8 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, U | 103 | 0.1974 | 0.1995 | 0.1915 | 2.11 | 0.42 | 2.78 | 0.63 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 189 | 0.267 | 0.2613 | 0.2554 | 2.22 | 0.45 | 2.78 | 0.63 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.01 | AOAC 2018.10 | 210 | 0.216 | 0.207 | 0.215 | 2.32 | 0.46 | 2.78 | 0.63 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, U | 218 | 0.242 | 0.251 | 0.24 | 2.4 | 0.49 | 2.78 | 0.63 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 182 | 0.2008 | 0.2049 | 0.1914 | 3.48 | 0.68 | 2.78 | 0.63 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 138 | 0.2234 | 0.2398 | 0.2265 | 3.79 | 0.76 | 2.78 | 0.63 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 192 | 0.2045 | 0.2212 | 0.2106 | 3.98 | 0.79 | 2.78 | 0.63 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 162 | 0.1976 | 0.1848 | 0.1825 | 4.32 | 0.84 | 2.78 | 0.63 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.30 | Other LC-MS | 105 | 0.271 | 0.252 | 0.238 | 6.53 | 1.33 | 2.78 | 0.63 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.60 | NIR | 187 | 0.243 | 0.274 | 0.237 | 7.9 | 1.61 | 2.78 | 0.63 | 10.8 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 153 | 0.2961 | 0.2559 | 0.2407 | 10.8 | 2.22 | 2.78 | 0.63 | 10.8 | 0 |
|  |  |  |  | Total CBD (%DW) |  |  |  |  |  |  |  |  |
| Total CBD (%DW) | 507.01 | AOAC 2018.10 | 210 | 8.277 | 8.218 | 8.229 | 0.38 | 0.13 | 1.57 | 0.38 | 6.59 | 0 |
| Total CBD (%DW) | 507.30 | Other LC-MS | 105 | 7.42 | 7.43 | 7.48 | 0.43 | 0.15 | 1.57 | 0.38 | 6.59 | 0 |
| Total CBD (%DW) | 507.02 | AOAC 2018.11, U | 199 | 8.3383 | 8.3166 | 8.4056 | 0.56 | 0.19 | 1.57 | 0.38 | 6.59 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 189 | 7.427 | 7.3436 | 7.3463 | 0.64 | 0.22 | 1.57 | 0.38 | 6.59 | 0 |
| Total CBD (%DW) | 507.02 | AOAC 2018.11, U | 218 | 7.613 | 7.634 | 7.446 | 1.36 | 0.46 | 1.57 | 0.38 | 6.59 | 0 |
| Total CBD (%DW) | 507.40 | GC-FID | 116 | 8.887 | 8.6467 | 8.7717 | 1.37 | 0.48 | 1.57 | 0.38 | 6.59 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 116 | 8.2826 | 8.0501 | 8.2335 | 1.5 | 0.51 | 1.57 | 0.38 | 6.59 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 172 | 7.9631 | 7.7355 | 7.8914 | 1.48 | 0.51 | 1.57 | 0.38 | 6.59 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 204 | 8.1095 | 7.9195 | 8.1762 | 1.65 | 0.57 | 1.57 | 0.38 | 6.59 | 0 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%DW) | 507.60 | NIR | 187 | 7.07 | 7.2 | 6.93 | 1.91 | 0.64 | 1.57 | 0.38 | 6.59 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 192 | 5.4364 | 5.6641 | 5.5171 | 2.08 | 0.68 | 1.57 | 0.38 | 6.59 | 0 |
| Total CBD (%DW) | 507.02 | AOAC 2018.11, U | 103 | 7.3964 | 7.0829 | 7.3808 | 2.42 | 0.82 | 1.57 | 0.38 | 6.59 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 162 | 9.2917 | 8.676 | 8.7087 | 3.9 | 1.35 | 1.57 | 0.38 | 6.59 | 0 |
| Total CBD (%DW) | 507.40 | GC-FID | 138 | 8.2654 | 9.122 | 9.3914 | 6.59 | 2.29 | 1.57 | 0.38 | 6.59 | 0 |

Lab

Population of Lab RSDr

CBDV (%DW)

| CBDV (%DW) | 508.02 | AOAC 2018.11, U | 210 | 0.042 | 0.041 | 0.041 | 1.4 | 0.22 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDV (%DW) | 508.99 | Other | 162 | 0.0491 | 0.0468 | 0.0508 | 4.11 | 0.65 | 0 |
| CBDV (%DW) | 508.99 | Other | 189 | 0.0469 | 0.0543 | 0.0472 | 8.47 | 1.35 | 0 |

CBDVA (%DW)

| CBDVA (%DW) | 509.01 | AOAC 2018.10 | 210 | 0.036 | 0.037 | 0.036 | 1.59 | 0.24 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

CBG (%DW)

| CBG (%DW) | 510.99 | Other | 189 | 0.3795 | 0.3764 | 0.3738 | 0.76 | 0.16 | 3.69 | 0.76 | 6.98 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%DW) | 510.99 | Other | 172 | 0.3116 | 0.3277 | 0.3234 | 2.6 | 0.55 | 3.69 | 0.76 | 6.98 | 0 |
| CBG (%DW) | 510.01 | AOAC 2018.10 | 210 | 0.345 | 0.346 | 0.329 | 2.81 | 0.6 | 3.69 | 0.76 | 6.98 | 0 |
| CBG (%DW) | 510.99 | Other | 105 | 0.287 | 0.307 | 0.282 | 4.53 | 0.94 | 3.69 | 0.76 | 6.98 | 0 |
| CBG (%DW) | 510.99 | Other | 204 | 0.3575 | 0.3553 | 0.3848 | 4.49 | 0.97 | 3.69 | 0.76 | 6.98 | 0 |
| CBG (%DW) | 510.99 | Other | 162 | 0.3608 | 0.3138 | 0.3359 | 6.98 | 1.49 | 3.69 | 0.76 | 6.98 | 0 |

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

RSDr	HorRat(r)

Population of Lab RSDr	 Rob Mean	min	max

Flag

CBGA (%DW)

| CBGA (%DW) | 511.01 | AOAC 2018.10 | 210 | 0.419 | 0.417 | 0.416 | 0.37 | 0.08 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%DW) | 511.99 | Other | 189 | 0.4024 | 0.3961 | 0.3974 | 0.83 | 0.18 | 0 |
| CBGA (%DW) | 511.99 | Other | 172 | 0.404 | 0.4104 | 0.3921 | 2.31 | 0.5 | 0 |
| CBGA (%DW) | 511.99 | Other | 204 | 0.424 | 0.42 | 0.4395 | 2.41 | 0.53 | 0 |
| CBGA (%DW) | 511.99 | Other | 162 | 0.5132 | 0.4353 | 0.4504 | 8.86 | 1.98 | 0 |

THCV (%DW)

| THCV (%DW) | 512.02 | AOAC 2018.11, U | 210 | <0.0130 <0.0130 |  | <0.0120 | 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| THCV (%DW) | 512.99 | Other | 189 | 0.0482 | 0.046 | 0.0441 | 4.45 | 0.7 | 0 |

Δ8-THC (%DW)

| Δ8-THC (%DW) | 513.99 | Other | 162 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%DW) | 513.99 | Other | 102 | <0.0212 | <0.0254 | <0.0208 |  |  | 1 |
| Δ8-THC (%DW) | 513.01 | AOAC 2018.10 | 210 | 0.023 | 0.026 | 0.027 | 8.22 | 1.19 | 0 |
| Δ8-THC (%DW) | 513.99 | Other | 189 | 0.0134 | 0.0105 | 0.0094 | 18.6 | 2.37 | 0 |

CBC (%DW)

| CBC (%DW) | 514.01 | AOAC 2018.10 | 210 | 0.236 | 0.24 | 0.238 | 0.84 | 0.17 | 2.16 | 0.84 | 5.93 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%DW) | 514.99 | Other | 189 | 0.2936 | 0.2878 | 0.2919 | 1.02 | 0.21 | 2.16 | 0.84 | 5.93 | 0 |
| CBC (%DW) | 514.99 | Other | 172 | 0.2611 | 0.2557 | 0.2568 | 1.11 | 0.23 | 2.16 | 0.84 | 5.93 | 0 |
| CBC (%DW) | 514.99 | Other | 204 | 0.2469 | 0.2396 | 0.2432 | 1.5 | 0.3 | 2.16 | 0.84 | 5.93 | 0 |

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%DW) | 514.02 | AOAC 2018.11, U | 103 | 0.2262 | 0.2272 | 0.2355 | 2.22 | 0.45 | 2.16 | 0.84 | 5.93 | 0 |
| CBC (%DW) | 514.99 | Other | 162 | 0.2509 | 0.2425 | 0.233 | 3.7 | 0.75 | 2.16 | 0.84 | 5.93 | 0 |
| CBC (%DW) | 514.99 | Other | 105 | 0.278 | 0.261 | 0.247 | 5.93 | 1.21 | 2.16 | 0.84 | 5.93 | 0 |

Lab

Population of Lab RSDr

- Flag definitions: 1 (Lab RSDr not calculated because there were less than 3 numeric lab results)

- Statistical parameters of robust mean, minimum, and maximum for RSDr only shown only if there were 6 or more lab RSDr values.

- AR refers to concentration on as-received basis. DW refers to concentration on dry weight basis.c