| Sample # | Replicate 1 | Replicate 2 |
| --- | --- | --- |
| 1 | 0.2047 | 0.2158 |
| 2 | 0.2080 | 0.2107 |
| 3 | 0.2128 | 0.2085 |
| 4 | 0.2127 | 0.2120 |
| 5 | 0.2107 | 0.2124 |
| 6 | 0.2032 | 0.2096 |
| 7 | 0.2052 | 0.2068 |

| Homogeneity and Stability |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sample ID . . . . . . . . . . . . . . . . . . . . . . . . . |  | HM23SEP-2 |  |  |  |
| Samples shipped . . . . . . . . . . . . . . . . . . . |  | 11-Sep-23 |  |  |  |
| Results due date . . . . . . . . . . . . . . . . . . . . |  | 20-Oct-23 |  |  |  |
| Homogeneity 7 sample jars randomly chosen for analysis. |  |  |  |  |  |
| Duplicate test portions analyzed for analyte and method shown |  | in table caption. |  |  |  |
| Total THC (%AR) via GC-FID (method 006.40) |  |  |  | %RSD | HorRat-r |
| Analysis Date: 9/25/23 |  | Overall average: | 0.2095 |  |  |
|  |  | SD of sample avgs: | 0.002459 | 1.17 |  |
|  |  | repeatability SD: | 0.0037214 | 1.78 | 0.4 |
|  |  | between-sample SD: | 0 | 0.00 |  |
|  |  | reproducibilty SD: | 0.0037214 | 1.78 |  |
|  |  | SD used for Analyte z scores: | 0.0329 |  |  |
|  |  | Check value: | 0.00987 |  |  |
| Total THC (%AR) via GC-FID (method 006.40) Is between-sample SD less than check value? YES Homogeneity test passed |  |  |  |  |  |
| Total CBD (%AR) via GC-FID (method 007.40) |  |  |  | %RSD | HorRat-r |
| Analysis Date: 9/25/23 |  | Overall average: | 7.3657 |  |  |
|  |  | SD of sample avgs: | 0.060154 | 0.82 |  |
|  |  | repeatability SD: | 0.092047 | 1.25 | 0.4 |
|  |  | between-sample SD: | 0 | 0.00 |  |
|  |  | reproducibilty SD: | 0.092047 | 1.25 |  |
|  |  | SD used for Analyte z scores: | 0.6903 |  |  |
|  |  | Check value: | 0.20709 |  |  |
| Total CBD (%AR) via GC-FID (method 007.40) Is between-sample SD less than check value? YES Homogeneity test passed |  |  |  |  |  |

| Sample # | Replicate 1 | Replicate 2 |
| --- | --- | --- |
| 1 | 7.2118 | 7.3822 |
| 2 | 7.3517 | 7.4150 |
| 3 | 7.4656 | 7.3878 |
| 4 | 7.4006 | 7.4510 |
| 5 | 7.3275 | 7.4824 |
| 6 | 7.2267 | 7.4568 |
| 7 | 7.2843 | 7.2770 |

Stability

Sample tested on different days from date sample shipped to results due date. Sample stored at room temperature during period of testing.

| Date of Analysis | Days | Conc. |  |
| --- | --- | --- | --- |
| 9/8/23 | 0 | 0.1980 |  |
| 9/21/23 | 13 | 0.2166 |  |
| 9/25/23 | 17 | 0.2101 |  |
| 9/27/23 | 19 | 0.2100 |  |
| 10/3/23 | 25 | 0.1916 |  |
| 10/9/23 | 31 | 0.2028 |  |
| 10/12/23 | 34 | 0.2179 |  |
| 10/24/23 | 46 | 0.2012 |  |
| 0.24 0.22 0.20 0.18 0.16 y = 2E-06x + 0.206 R² = 1E-05 0.14 0.12 0.10 0 10 20 30 40 50 |  |  |  |

Total THC (%AR) via GC-FID (method 006.40)		Days = x variable, Conc = y variable

SUMMARY OUTPUT

Regression Statistics

Multiple R	0.0038205 R Square		1.46E-05 Adjusted R Squ -0.1666496 Standard Error  0.0099281

Observations	8

| ANOVA |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | df |  | SS | MS | F | Significance F |  |
| Regression |  | 1 | 8.6324E-09 | 8.63238E-09 | 8.75782E-05 | 0.992836649 |  |
| Residual |  | 6 | 0.00059141 | 9.85677E-05 |  |  |  |
| Total |  | 7 | 0.00059142 |  |  |  |  |
|  | Coefficients Standard Error |  |  | t Stat | P-value | Lower 95% | Upper 95% |
| Intercept | 0.2059676 0.00707119 |  |  | 29.12771657 | 1.085E-07 | 0.188664981 | 0.223270128 |
| X Variable 1 | 2.484E-06 0.00026545 |  |  | 0.009358322 | 0.992836649 | -0.000647041 | 0.00065201 |

Coefficient for x variable is not statistically significantly different from 0. No signficant change in result during testing period.

Stability test passed.

%RSD = 4.5	HorRat-R = 0.9

| Total CBD (%AR) via GC-FID (method 006.40) |  |  |  |
| --- | --- | --- | --- |
| Date of Analysis | Days | Conc. |  |
| 9/8/23 | 0 | 7.2692 |  |
| 9/21/23 | 13 | 7.1864 |  |
| 9/25/23 | 17 | 7.2688 |  |
| 9/27/23 | 19 | 7.5574 |  |
| 10/3/23 | 25 | 6.8713 |  |
| 10/9/23 | 31 | 7.4821 |  |
| 10/12/23 | 34 | 7.1984 |  |
| 10/24/23 | 46 | 6.9832 |  |
| 8.0 7.8 7.6 7.4 7.2 7.0 6.8 6.6 6.4 y = -0.0048x + 7.3383 6.2 R² = 0.0884 6.0 0 10 20 30 40 50 |  |  |  |

Days = x variable, Conc = y variable SUMMARY OUTPUT

Regression Statistics

Multiple R	0.2973376 R Square	0.0884096 Adjusted R Squ -0.0635221 Standard Error  0.2358743

Observations	8

| ANOVA |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | df |  | SS | MS | F | Significance F |  |
| Regression |  | 1 | 0.0323752 | 0.032375203 | 0.581903696 | 0.474479903 |  |
| Residual |  | 6 | 0.33382022 | 0.055636703 |  |  |  |
| Total |  | 7 | 0.36619542 |  |  |  |  |
|  | Coefficients Standard Error |  |  | t Stat | P-value | Lower 95% | Upper 95% |
| Intercept | 7.3383496 0.16799861 |  |  | 43.68101359 | 9.63758E-09 | 6.927271827 | 7.749427413 |
| X Variable 1 | -0.0048108 0.00630654 |  |  | -0.76282612 | 0.474479903 | -0.020242345 | 0.010620756 |

Coefficient for x variable is not statistically significantly different from 0. No signficant change in result during testing period.

Stability test passed.

%RSD = 3.2	HorRat-R = 1.1