Hemp Proficiency Testing

All Labs Trueness HM24NOV-2 (hemp)

Sample: HM24NOV-2 (hemp)

All Labs Trueness

Issue Date: December 21, 2024

| Analyte | Code | Method | Lab Num |  | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Δ9-THC (%AR) |  |  |  |  |  |  |  |  |  |  |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 213 | 0.0911 | 0.0884 | 0.0917 | 0.0904 | -3.8 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 141 | 0.1560 | 0.1520 | 0.1486 | 0.152 | -1.63 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 192 | 0.1395 | 0.1603 | 0.1629 | 0.154 | -1.55 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 182 | 0.1756 | 0.1586 | 0.1558 | 0.163 | -1.23 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 |  | 148 | 0.1510 | 0.1720 | 0.1680 | 0.164 | -1.22 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U |  | 161 | 0.1532 | 0.1674 | 0.1808 | 0.167 | -1.1 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 149 | 0.1830 | 0.1720 | 0.1740 | 0.176 | -0.77 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 144 | 0.1926 | 0.1755 | 0.1828 | 0.184 | -0.52 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 170 | 0.1879 | 0.1828 | 0.1842 | 0.185 | -0.47 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U |  | 103 | 0.2037 | 0.1863 | 0.1808 | 0.19 | -0.28 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U |  | 125 | 0.2069 | 0.1921 | 0.1812 | 0.193 | -0.17 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 196 | 0.1970 | 0.1910 | 0.1926 | 0.194 | -0.17 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 106 | 0.1900 | 0.2000 | 0.2000 | 0.197 | -0.06 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 165 | 0.1948 | 0.1957 | 0.2090 | 0.2 | 0.05 | 0.1983 | 29 | 0.0284 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score Rob Mean |  | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 199 | 0.1990 | 0.1990 | 0.2115 | 0.203 | 0.17 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 115 | 0.2059 | 0.2124 | 0.2030 | 0.207 | 0.31 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 116 | 0.2057 | 0.2024 | 0.2130 | 0.207 | 0.31 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 183 | 0.2075 | 0.2122 | 0.2072 | 0.209 | 0.37 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 184 | 0.2092 | 0.2086 | 0.2095 | 0.209 | 0.38 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 102 | 0.2040 | 0.2110 | 0.2110 | 0.209 | 0.36 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 175 | 0.2100 | 0.2100 | 0.2100 | 0.21 | 0.41 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.03 | AOAC 2018.11, M | 139 | 0.2278 | 0.2055 | 0.2185 | 0.217 | 0.67 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 173 | 0.2270 | 0.2010 | 0.2320 | 0.22 | 0.76 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 101 | 0.2360 | 0.2040 | 0.2310 | 0.224 | 0.89 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 122 | 0.2348 | 0.2169 | 0.2270 | 0.226 | 0.98 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 163 | 0.2144 | 0.2226 | 0.2400 | 0.226 | 0.96 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 197 | 0.2411 | 0.2298 | 0.2230 | 0.231 | 1.16 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 189 | 0.2353 | 0.2313 | 0.2308 | 0.232 | 1.2 | 0.1983 | 29 | 0.0284 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 164 | 0.2557 | 0.2414 | 0.2294 | 0.242 | 1.54 | 0.1983 | 29 | 0.0284 | 0 |

Population of Lab Values

| Δ9-THCA (%AR) |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 199 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 161 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 197 | <0.0701 | <0.0739 | <0.0732 | <0.0701 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 192 | 0.0345 | 0.0000 | 0.0000 | 0.0115 | -4.31 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 213 | 0.0129 | 0.0126 | 0.0136 | 0.013 | -4.15 | 0.0521 | 27 | 0.0094 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 141 | 0.0281 | 0.0298 | 0.0297 | 0.0292 | -2.43 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 182 | 0.0373 | 0.0411 | 0.0392 | 0.0392 | -1.37 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 196 | 0.0433 | 0.0418 | 0.0448 | 0.0433 | -0.94 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 173 | 0.0450 | 0.0420 | 0.0440 | 0.0437 | -0.9 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 149 | 0.0510 | 0.0440 | 0.0440 | 0.0463 | -0.61 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 103 | 0.0435 | 0.0402 | 0.0559 | 0.0465 | -0.59 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 175 | 0.0460 | 0.0470 | 0.0530 | 0.0487 | -0.37 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 106 | 0.0500 | 0.0500 | 0.0500 | 0.05 | -0.22 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 144 | 0.0492 | 0.0518 | 0.0495 | 0.0502 | -0.21 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.03 | AOAC 2018.11, M | 139 | 0.0516 | 0.0470 | 0.0501 | 0.0496 | -0.27 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 170 | 0.0505 | 0.0536 | 0.0541 | 0.0527 | 0.07 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 149 | 0.0630 | 0.0500 | 0.0520 | 0.055 | 0.31 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 184 | 0.0539 | 0.0539 | 0.0575 | 0.0551 | 0.32 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 116 | 0.0536 | 0.0549 | 0.0571 | 0.0552 | 0.33 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 163 | 0.0546 | 0.0550 | 0.0594 | 0.0563 | 0.45 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 189 | 0.0563 | 0.0555 | 0.0553 | 0.0557 | 0.38 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 102 | 0.0540 | 0.0580 | 0.0590 | 0.057 | 0.52 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 115 | 0.0590 | 0.0578 | 0.0542 | 0.057 | 0.52 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 164 | 0.0614 | 0.0599 | 0.0548 | 0.0587 | 0.7 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 183 | 0.0604 | 0.0622 | 0.0570 | 0.0599 | 0.82 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 165 | 0.0569 | 0.0588 | 0.0639 | 0.0599 | 0.82 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 148 | 0.0610 | 0.0580 | 0.0630 | 0.0607 | 0.91 | 0.0521 | 27 | 0.0094 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 122 | 0.0646 | 0.0588 | 0.0628 | 0.0621 | 1.06 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 101 | 0.0750 | 0.0610 | 0.0570 | 0.0643 | 1.3 | 0.0521 | 27 | 0.0094 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 125 | 0.0920 | 0.0867 | 0.0785 | 0.0857 | 3.57 | 0.0521 | 27 | 0.0094 | 0 |
|  |  |  |  | CBD (%AR) |  |  |  |  |  |  |  |  |
| CBD (%AR) | 003.30 | Other LC-MS | 149 | 4.7130 | 3.5860 | 4.0300 | 4.11 | -1.51 | 4.8273 | 23 | 0.4758 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 184 | 4.2300 | 4.2950 | 3.8220 | 4.12 | -1.5 | 4.8273 | 23 | 0.4758 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 161 | 4.0937 | 4.2454 | 4.3165 | 4.22 | -1.28 | 4.8273 | 23 | 0.4758 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 192 | 3.8649 | 4.5062 | 4.4303 | 4.27 | -1.18 | 4.8273 | 23 | 0.4758 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 213 | 4.4320 | 4.4030 | 4.4480 | 4.43 | -0.84 | 4.8273 | 23 | 0.4758 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 228 | 4.5030 | 4.4880 | 4.5220 | 4.5 | -0.68 | 4.8273 | 23 | 0.4758 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 149 | 4.5560 | 4.7750 | 4.6660 | 4.67 | -0.34 | 4.8273 | 23 | 0.4758 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 103 | 4.8081 | 4.6547 | 4.6023 | 4.69 | -0.29 | 4.8273 | 23 | 0.4758 | 0 |
| CBD (%AR) | 003.03 | AOAC 2018.11, M | 139 | 4.7843 | 4.6252 | 4.8814 | 4.76 | -0.13 | 4.8273 | 23 | 0.4758 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 173 | 4.8290 | 4.7590 | 4.7340 | 4.77 | -0.11 | 4.8273 | 23 | 0.4758 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 196 | 4.9547 | 4.8187 | 4.7481 | 4.84 | 0.03 | 4.8273 | 23 | 0.4758 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 141 | 5.1014 | 4.7641 | 4.6698 | 4.85 | 0.04 | 4.8273 | 23 | 0.4758 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 197 | 5.0178 | 4.7586 | 4.8468 | 4.87 | 0.1 | 4.8273 | 23 | 0.4758 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 170 | 4.9199 | 4.8153 | 4.9940 | 4.91 | 0.17 | 4.8273 | 23 | 0.4758 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 144 | 4.9471 | 4.8872 | 4.9243 | 4.92 | 0.19 | 4.8273 | 23 | 0.4758 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 199 | 4.9410 | 4.9480 | 5.1015 | 5 | 0.36 | 4.8273 | 23 | 0.4758 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 125 | 5.2471 | 4.9708 | 4.9697 | 5.06 | 0.49 | 4.8273 | 23 | 0.4758 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%AR) | 003.10 | Other LC-UV | 183 | 5.0856 | 5.0096 | 5.3020 | 5.13 | 0.64 | 4.8273 | 23 | 0.4758 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 189 | 5.1862 | 5.1695 | 5.1635 | 5.17 | 0.73 | 4.8273 | 23 | 0.4758 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 116 | 5.1932 | 5.0398 | 5.4180 | 5.22 | 0.82 | 4.8273 | 23 | 0.4758 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 163 | 5.4809 | 5.3907 | 5.4385 | 5.44 | 1.28 | 4.8273 | 23 | 0.4758 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 106 | 5.5300 | 5.8000 | 5.6100 | 5.65 | 1.72 | 4.8273 | 23 | 0.4758 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 164 | 6.6286 | 5.4613 | 5.9590 | 6.02 | 2.5 | 4.8273 | 23 | 0.4758 | 0 |
|  |  |  |  | CBDA (%AR) |  |  |  |  |  |  |  |  |
| CBDA (%AR) | 004.30 | Other LC-MS | 149 | 3.4680 | 2.8630 | 3.5470 | 3.29 | -1.6 | 3.7946 | 22 | 0.313 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 161 | 3.2819 | 3.4102 | 3.4809 | 3.39 | -1.29 | 3.7946 | 22 | 0.313 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 184 | 3.4860 | 3.5540 | 3.1960 | 3.41 | -1.22 | 3.7946 | 22 | 0.313 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 192 | 3.6687 | 3.4993 | 3.4058 | 3.52 | -0.86 | 3.7946 | 22 | 0.313 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 213 | 3.5040 | 3.5120 | 3.6050 | 3.54 | -0.81 | 3.7946 | 22 | 0.313 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 173 | 3.6290 | 3.5720 | 3.5710 | 3.59 | -0.65 | 3.7946 | 22 | 0.313 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 197 | 3.7187 | 3.5232 | 3.6020 | 3.61 | -0.58 | 3.7946 | 22 | 0.313 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 144 | 3.6611 | 3.7378 | 3.7468 | 3.72 | -0.25 | 3.7946 | 22 | 0.313 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 116 | 3.5669 | 3.8160 | 3.8311 | 3.74 | -0.18 | 3.7946 | 22 | 0.313 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 141 | 3.8943 | 3.7300 | 3.6579 | 3.76 | -0.11 | 3.7946 | 22 | 0.313 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 149 | 3.8060 | 3.7810 | 3.7510 | 3.78 | -0.05 | 3.7946 | 22 | 0.313 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 196 | 3.8359 | 3.7303 | 3.8121 | 3.79 | -0.01 | 3.7946 | 22 | 0.313 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 103 | 3.9975 | 3.7852 | 3.6856 | 3.82 | 0.09 | 3.7946 | 22 | 0.313 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 189 | 3.8372 | 3.8230 | 3.8137 | 3.82 | 0.1 | 3.7946 | 22 | 0.313 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDA (%AR) | 004.10 | Other LC-UV | 183 | 3.8499 | 3.9422 | 3.7778 | 3.86 | 0.2 | 3.7946 | 22 | 0.313 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 125 | 4.1762 | 3.8093 | 3.8305 | 3.94 | 0.46 | 3.7946 | 22 | 0.313 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 199 | 3.8905 | 3.9475 | 4.0415 | 3.96 | 0.53 | 3.7946 | 22 | 0.313 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 106 | 3.9600 | 4.1400 | 3.9800 | 4.03 | 0.74 | 3.7946 | 22 | 0.313 | 0 |
| CBDA (%AR) | 004.03 | AOAC 2018.11, M | 139 | 3.9632 | 4.0343 | 4.2306 | 4.08 | 0.9 | 3.7946 | 22 | 0.313 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 170 | 4.3066 | 4.2195 | 4.5350 | 4.35 | 1.79 | 3.7946 | 22 | 0.313 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 163 | 4.4901 | 4.4589 | 4.3976 | 4.45 | 2.09 | 3.7946 | 22 | 0.313 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 164 | 6.7370 | 5.8234 | 6.2997 | 6.29 | 7.96 | 3.7946 | 22 | 0.313 | 0 |
|  |  |  |  | CBN (%AR) |  |  |  |  |  |  |  |  |
| CBN (%AR) | 005.10 | Other LC-UV | 197 | <0.0701 | <0.0739 | <0.0732 | <0.0701 |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 192 | 0.0310 | 0.0337 | 0.0327 | 0.0325 | -1.6 | 0.0505 | 20 | 0.0113 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 141 | 0.0394 | 0.0345 | 0.0356 | 0.0365 | -1.24 | 0.0505 | 20 | 0.0113 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 144 | 0.0394 | 0.0405 | 0.0407 | 0.0402 | -0.91 | 0.0505 | 20 | 0.0113 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 149 | 0.0420 | 0.0410 | 0.0410 | 0.0413 | -0.81 | 0.0505 | 20 | 0.0113 | 0 |
| CBN (%AR) | 005.03 | AOAC 2018.11, M | 139 | 0.0438 | 0.0394 | 0.0430 | 0.0421 | -0.74 | 0.0505 | 20 | 0.0113 | 0 |
| CBN (%AR) | 005.20 | Other LC-MS | 149 | 0.0520 | 0.0420 | 0.0360 | 0.0433 | -0.63 | 0.0505 | 20 | 0.0113 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 213 | 0.0439 | 0.0423 | 0.0436 | 0.0433 | -0.64 | 0.0505 | 20 | 0.0113 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 196 | 0.0446 | 0.0441 | 0.0448 | 0.0445 | -0.53 | 0.0505 | 20 | 0.0113 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 192 | 0.0445 | 0.0447 | 0.0461 | 0.0451 | -0.48 | 0.0505 | 20 | 0.0113 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 199 | 0.0505 | 0.0500 | 0.0525 | 0.051 | 0.05 | 0.0505 | 20 | 0.0113 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 183 | 0.0504 | 0.0529 | 0.0531 | 0.0521 | 0.15 | 0.0505 | 20 | 0.0113 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%AR) | 005.10 | Other LC-UV | 173 | 0.0570 | 0.0510 | 0.0530 | 0.0537 | 0.28 | 0.0505 | 20 | 0.0113 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 189 | 0.0530 | 0.0537 | 0.0549 | 0.0539 | 0.3 | 0.0505 | 20 | 0.0113 | 0 |
| CBN (%AR) | 005.20 | Other LC-MS | 122 | 0.0569 | 0.0541 | 0.0568 | 0.0559 | 0.49 | 0.0505 | 20 | 0.0113 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 116 | 0.0490 | 0.0596 | 0.0622 | 0.0569 | 0.57 | 0.0505 | 20 | 0.0113 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 116 | 0.0609 | 0.0568 | 0.0564 | 0.058 | 0.67 | 0.0505 | 20 | 0.0113 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 138 | 0.0560 | 0.0572 | 0.0675 | 0.0602 | 0.87 | 0.0505 | 20 | 0.0113 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 106 | 0.0700 | 0.0600 | 0.0600 | 0.0633 | 1.14 | 0.0505 | 20 | 0.0113 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 163 | 0.0678 | 0.0677 | 0.0651 | 0.0669 | 1.46 | 0.0505 | 20 | 0.0113 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 125 | 0.1035 | 0.0571 | 0.0583 | 0.073 | 2 | 0.0505 | 20 | 0.0113 | 0 |
|  |  |  |  | Total Δ9-THC (%AR) |  |  |  |  |  |  |  |  |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 213 | 0.1020 | 0.0990 | 0.1030 | 0.101 | -3.89 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 192 | 0.1698 | 0.1603 | 0.1629 | 0.164 | -2.19 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 141 | 0.1807 | 0.1781 | 0.1746 | 0.178 | -1.83 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 182 | 0.2083 | 0.1946 | 0.1901 | 0.198 | -1.3 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 192 | 0.2019 | 0.1951 | 0.2015 | 0.199 | -1.25 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 107 | 0.2050 | 0.1880 | 0.2080 | 0.2 | -1.23 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 149 | 0.2280 | 0.2100 | 0.2120 | 0.217 | -0.79 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 148 | 0.2040 | 0.2220 | 0.2240 | 0.217 | -0.79 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 199 | 0.2209 | 0.2209 | 0.2334 | 0.225 | -0.56 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 114 | 0.2404 | 0.2280 | 0.2164 | 0.228 | -0.48 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 144 | 0.2358 | 0.2209 | 0.2262 | 0.228 | -0.49 | 0.246 | 36 | 0.0372 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 103 | 0.2418 | 0.2216 | 0.2298 | 0.231 | -0.4 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 197 | 0.2411 | 0.2298 | 0.2230 | 0.231 | -0.39 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 196 | 0.2350 | 0.2276 | 0.2319 | 0.231 | -0.39 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 149 | 0.2530 | 0.2350 | 0.2290 | 0.239 | -0.19 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 138 | 0.2353 | 0.2453 | 0.2391 | 0.24 | -0.16 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 116 | 0.2633 | 0.2463 | 0.2439 | 0.251 | 0.14 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 169 | 0.2470 | 0.2506 | 0.2548 | 0.251 | 0.13 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 165 | 0.2447 | 0.2473 | 0.2650 | 0.252 | 0.17 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 175 | 0.2503 | 0.2512 | 0.2565 | 0.253 | 0.18 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 153 | 0.2550 | 0.2641 | 0.2453 | 0.255 | 0.24 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 116 | 0.2527 | 0.2506 | 0.2631 | 0.255 | 0.26 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 173 | 0.2635 | 0.2380 | 0.2710 | 0.257 | 0.31 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 115 | 0.2577 | 0.2631 | 0.2505 | 0.257 | 0.3 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 102 | 0.2510 | 0.2620 | 0.2630 | 0.259 | 0.34 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.03 | AOAC 2018.11, M | 139 | 0.2731 | 0.2467 | 0.2624 | 0.261 | 0.4 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 183 | 0.2610 | 0.2668 | 0.2572 | 0.262 | 0.42 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 125 | 0.2877 | 0.2682 | 0.2500 | 0.269 | 0.61 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.50 | GC-MS | 230 | 0.2870 | 0.2631 | 0.2596 | 0.27 | 0.64 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 101 | 0.3020 | 0.2570 | 0.2810 | 0.28 | 0.91 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 122 | 0.2915 | 0.2685 | 0.2820 | 0.281 | 0.93 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 189 | 0.2847 | 0.2800 | 0.2793 | 0.281 | 0.95 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 164 | 0.3171 | 0.3013 | 0.2842 | 0.301 | 1.47 | 0.246 | 36 | 0.0372 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 190 | 0.3111 | 0.3078 | 0.3086 | 0.309 | 1.7 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 163 | 0.3088 | 0.3062 | 0.3164 | 0.31 | 1.73 | 0.246 | 36 | 0.0372 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 170 | 0.3134 | 0.3124 | 0.3070 | 0.311 | 1.75 | 0.246 | 36 | 0.0372 | 0 |
|  |  |  |  | Total CBD (%AR) |  |  |  |  |  |  |  |  |
| Total CBD (%AR) | 007.40 | GC-FID | 192 | 7.0018 | 6.4649 | 6.8771 | 6.78 | -2.93 | 8.2952 | 22 | 0.5164 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 149 | 7.7550 | 6.0970 | 7.1410 | 7 | -2.51 | 8.2952 | 22 | 0.5164 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 192 | 7.0824 | 7.5751 | 7.4172 | 7.36 | -1.81 | 8.2952 | 22 | 0.5164 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 173 | 8.0120 | 7.8920 | 7.8660 | 7.92 | -0.72 | 8.2952 | 22 | 0.5164 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 149 | 7.8940 | 8.0910 | 7.9550 | 7.98 | -0.61 | 8.2952 | 22 | 0.5164 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 103 | 8.3139 | 7.9743 | 7.8346 | 8.04 | -0.49 | 8.2952 | 22 | 0.5164 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 197 | 8.2790 | 7.8484 | 8.0058 | 8.04 | -0.49 | 8.2952 | 22 | 0.5164 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 141 | 8.5167 | 8.0354 | 7.8778 | 8.14 | -0.29 | 8.2952 | 22 | 0.5164 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 196 | 8.3188 | 8.0902 | 8.0913 | 8.17 | -0.25 | 8.2952 | 22 | 0.5164 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 144 | 8.1586 | 8.1660 | 8.2110 | 8.18 | -0.23 | 8.2952 | 22 | 0.5164 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 170 | 8.2533 | 8.4256 | 7.9849 | 8.22 | -0.14 | 8.2952 | 22 | 0.5164 | 0 |
| Total CBD (%AR) | 007.03 | AOAC 2018.11, M | 139 | 8.2600 | 8.1633 | 8.5916 | 8.34 | 0.08 | 8.2952 | 22 | 0.5164 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 199 | 8.3530 | 8.4100 | 8.6459 | 8.47 | 0.34 | 8.2952 | 22 | 0.5164 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 116 | 8.3214 | 8.3864 | 8.7778 | 8.5 | 0.39 | 8.2952 | 22 | 0.5164 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 125 | 8.9096 | 8.3115 | 8.3291 | 8.52 | 0.43 | 8.2952 | 22 | 0.5164 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 183 | 8.4600 | 8.4700 | 8.6200 | 8.52 | 0.43 | 8.2952 | 22 | 0.5164 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 189 | 8.5515 | 8.5223 | 8.5081 | 8.53 | 0.45 | 8.2952 | 22 | 0.5164 | 0 |

Population of Lab Values

Lab

Lab

Population of Lab Values

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Value | Z score Rob Mean |  | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%AR) | 007.40 | GC-FID | 163 | 8.5279 | 8.5250 | 8.5719 | 8.54 | 0.48 | 8.2952 | 22 | 0.5164 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 116 | 9.0817 | 8.4112 | 8.6032 | 8.7 | 0.78 | 8.2952 | 22 | 0.5164 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 169 | 8.8020 | 8.9290 | 9.2410 | 8.99 | 1.35 | 8.2952 | 22 | 0.5164 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 138 | 9.5589 | 9.5774 | 10.1292 | 9.76 | 2.83 | 8.2952 | 22 | 0.5164 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 164 | 12.5370 | 10.5685 | 11.4839 | 11.5 | 6.26 | 8.2952 | 22 | 0.5164 | 0 |

| CBDV (%AR) |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDV (%AR) | 008.99 | Other | 197 | <0.0467 | <0.0493 | <0.0488 | <0.0467 |  |  |  |  | 1 |
| CBDV (%AR) | 008.99 | Other | 196 | 0.0250 | 0.0240 | 0.0238 | 0.0243 | -0.94 | 0.039 | 6 | 0.0157 | 0 |
| CBDV (%AR) | 008.99 | Other | 149 | 0.0330 | 0.0270 | 0.0320 | 0.0307 | -0.53 | 0.039 | 6 | 0.0157 | 0 |
| CBDV (%AR) | 008.99 | Other | 189 | 0.0306 | 0.0320 | 0.0321 | 0.0316 | -0.48 | 0.039 | 6 | 0.0157 | 0 |
| CBDV (%AR) | 008.02 | AOAC 2018.11, U | 183 | 0.0371 | 0.0393 | 0.0390 | 0.0385 | -0.04 | 0.039 | 6 | 0.0157 | 0 |
| CBDV (%AR) | 008.99 | Other | 106 | 0.0500 | 0.0400 | 0.0500 | 0.0467 | 0.49 | 0.039 | 6 | 0.0157 | 0 |
| CBDV (%AR) | 008.99 | Other | 173 | 0.1020 | 0.1010 | 0.1110 | 0.105 | 4.19 | 0.039 | 6 | 0.0157 | 0 |

| CBDVA (%AR) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDVA (%AR) | 009.99 | Other | 197 | <0.0467 | <0.0493 | <0.0488 | <0.0467 | 1 |
| CBDVA (%AR) | 009.99 | Other | 196 | 0.0385 | 0.0374 | 0.0363 | 0.0374 | 0 |
| CBDVA (%AR) | 009.99 | Other | 149 | 0.0400 | 0.0370 | 0.0370 | 0.038 | 0 |
| CBDVA (%AR) | 009.99 | Other | 173 | 0.0430 | 0.0410 | 0.0440 | 0.0427 | 0 |
| CBDVA (%AR) | 009.02 | AOAC 2018.11, U | 183 | 0.0430 | 0.0630 | 0.0440 | 0.05 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | CBG (%AR) |  |  |  |  |  |  |  |  |
| CBG (%AR) | 010.99 | Other | 197 | 0.0897 | 0.1143 | 0.1161 | 0.107 | -2.39 | 0.2253 | 10 | 0.0497 | 0 |
| CBG (%AR) | 010.99 | Other | 141 | 0.1848 | 0.1810 | 0.1741 | 0.18 | -0.91 | 0.2253 | 10 | 0.0497 | 0 |
| CBG (%AR) | 010.99 | Other | 149 | 0.1970 | 0.1830 | 0.1890 | 0.19 | -0.72 | 0.2253 | 10 | 0.0497 | 0 |
| CBG (%AR) | 010.99 | Other | 196 | 0.2049 | 0.2040 | 0.1985 | 0.202 | -0.46 | 0.2253 | 10 | 0.0497 | 0 |
| CBG (%AR) | 010.99 | Other | 106 | 0.2400 | 0.2200 | 0.2100 | 0.223 | -0.04 | 0.2253 | 10 | 0.0497 | 0 |
| CBG (%AR) | 010.02 | AOAC 2018.11, U | 183 | 0.2414 | 0.2436 | 0.2331 | 0.239 | 0.28 | 0.2253 | 10 | 0.0497 | 0 |
| CBG (%AR) | 010.99 | Other | 173 | 0.2500 | 0.2440 | 0.2430 | 0.246 | 0.41 | 0.2253 | 10 | 0.0497 | 0 |
| CBG (%AR) | 010.99 | Other | 122 | 0.2651 | 0.2571 | 0.2666 | 0.263 | 0.76 | 0.2253 | 10 | 0.0497 | 0 |
| CBG (%AR) | 010.02 | AOAC 2018.11, U | 125 | 0.2798 | 0.2523 | 0.3037 | 0.279 | 1.07 | 0.2253 | 10 | 0.0497 | 0 |
| CBG (%AR) | 010.99 | Other | 189 | 0.2769 | 0.2772 | 0.2861 | 0.28 | 1.1 | 0.2253 | 10 | 0.0497 | 0 |
|  |  |  |  | CBGA (%AR) |  |  |  |  |  |  |  |  |
| CBGA (%AR) | 011.99 | Other | 141 | 0.1321 | 0.1280 | 0.1376 | 0.133 | -2.11 | 0.1965 | 10 | 0.0303 | 0 |
| CBGA (%AR) | 011.99 | Other | 173 | 0.1710 | 0.1660 | 0.1690 | 0.169 | -0.92 | 0.1965 | 10 | 0.0303 | 0 |
| CBGA (%AR) | 011.99 | Other | 106 | 0.1800 | 0.1800 | 0.1800 | 0.18 | -0.54 | 0.1965 | 10 | 0.0303 | 0 |
| CBGA (%AR) | 011.99 | Other | 196 | 0.1883 | 0.1873 | 0.1925 | 0.189 | -0.24 | 0.1965 | 10 | 0.0303 | 0 |
| CBGA (%AR) | 011.99 | Other | 149 | 0.2030 | 0.1840 | 0.1850 | 0.191 | -0.19 | 0.1965 | 10 | 0.0303 | 0 |
| CBGA (%AR) | 011.99 | Other | 197 | 0.2000 | 0.1911 | 0.2039 | 0.198 | 0.06 | 0.1965 | 10 | 0.0303 | 0 |
| CBGA (%AR) | 011.99 | Other | 189 | 0.2130 | 0.2078 | 0.2081 | 0.21 | 0.43 | 0.1965 | 10 | 0.0303 | 0 |
| CBGA (%AR) | 011.02 | AOAC 2018.11, U | 183 | 0.2119 | 0.2150 | 0.2062 | 0.211 | 0.48 | 0.1965 | 10 | 0.0303 | 0 |
| CBGA (%AR) | 011.99 | Other | 122 | 0.2264 | 0.2204 | 0.2261 | 0.224 | 0.92 | 0.1965 | 10 | 0.0303 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%AR) | 011.02 | AOAC 2018.11, U | 125 | 0.2593 | 0.2350 | 0.2405 | 0.245 | 1.6 | 0.1965 | 10 | 0.0303 | 0 |
|  |  |  |  | THCV (%AR) |  |  |  |  |  |  |  |  |
| THCV (%AR) | 012.99 | Other | 189 | <0.0050 | <0.0050 | <0.0050 | <0.0050 |  |  |  |  | 1 |
| THCV (%AR) | 012.02 | AOAC 2018.11, U | 183 | <0.0070 | <0.0070 | <0.0070 | <0.0070 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 | <0.0205 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 106 | <0.0300 | <0.0300 | <0.0300 | <0.0300 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 197 | <0.0467 | <0.0493 | <0.0488 | <0.0467 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 173 | 0.0000 | 0.0000 | 0.0000 | 0 |  |  |  |  | 0 |
|  |  |  |  | Δ8-THC (%AR) |  |  |  |  |  |  |  |  |
| Δ8-THC (%AR) | 013.02 | AOAC 2018.11, U | 183 | <0.0104 | <0.0104 | <0.0104 | <0.0104 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 141 | <0.0120 | <0.0120 | <0.0120 | <0.0120 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 | <0.0205 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 106 | <0.0300 | <0.0300 | <0.0300 | <0.0300 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 197 | <0.0467 | <0.0493 | <0.0488 | <0.0467 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 173 | 0.0000 | 0.0000 | 0.0000 | 0 |  |  |  |  | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 189 | 0.0126 | 0.0090 | 0.0140 | 0.0119 |  |  |  |  | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 138 | 0.0520 | 0.0541 | 0.0587 | 0.0549 |  |  |  |  | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 213 | 0.1770 | 0.1780 | 0.1780 | 0.178 |  |  |  |  | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | CBC (%AR) |  |  |  |  |  |  |  |
| CBC (%AR) | 014.99 | Other | 149 | 0.0510 | 0.0440 | 0.0440 | 0.0463 | -6.9 | 0.2368 | 8 | 0.0276 | 0 |
| CBC (%AR) | 014.99 | Other | 197 | 0.2224 | 0.2138 | 0.2380 | 0.225 | -0.44 | 0.2368 | 8 | 0.0276 | 0 |
| CBC (%AR) | 014.99 | Other | 173 | 0.2310 | 0.2250 | 0.2240 | 0.227 | -0.37 | 0.2368 | 8 | 0.0276 | 0 |
| CBC (%AR) | 014.99 | Other | 196 | 0.2331 | 0.2285 | 0.2286 | 0.23 | -0.24 | 0.2368 | 8 | 0.0276 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, U | 125 | 0.2444 | 0.2309 | 0.2375 | 0.238 | 0.03 | 0.2368 | 8 | 0.0276 | 0 |
| CBC (%AR) | 014.99 | Other | 106 | 0.2400 | 0.2500 | 0.2500 | 0.247 | 0.36 | 0.2368 | 8 | 0.0276 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, U | 183 | 0.2568 | 0.2567 | 0.2520 | 0.255 | 0.66 | 0.2368 | 8 | 0.0276 | 0 |
| CBC (%AR) | 014.99 | Other | 189 | 0.2971 | 0.2810 | 0.2807 | 0.286 | 1.79 | 0.2368 | 8 | 0.0276 | 0 |
|  |  |  |  | Cd (ug/g AR) |  |  |  |  |  |  |  |  |
| Cd (ug/g AR) | 101.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 | <0.2000 |  |  |  |  | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 196 | 0.1290 | 0.1280 | 0.1180 | 0.125 | -1.57 | 0.1426 | 7 | 0.0112 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 141 | 0.1430 | 0.1370 | 0.1310 | 0.137 | -0.5 | 0.1426 | 7 | 0.0112 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 102 | 0.1416 | 0.1413 | 0.1363 | 0.14 | -0.25 | 0.1426 | 7 | 0.0112 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 164 | 0.1388 | 0.1520 | 0.1324 | 0.141 | -0.13 | 0.1426 | 7 | 0.0112 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 106 | 0.1560 | 0.1430 | 0.1440 | 0.148 | 0.46 | 0.1426 | 7 | 0.0112 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 197 | 0.1476 | 0.1464 | 0.1621 | 0.152 | 0.85 | 0.1426 | 7 | 0.0112 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 183 | 0.1530 | 0.1550 | 0.1560 | 0.155 | 1.08 | 0.1426 | 7 | 0.0112 | 0 |
|  |  |  |  | As (ug/g AR) |  |  |  |  |  |  |  |  |
| As (ug/g AR) | 102.04 | ICP-MS | 106 | <0.0800 <0.0800 <0.0800 |  |  | <0.0800 |  |  |  |  | 1 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| As (ug/g AR) | 102.04 | ICP-MS | 197 | <0.0969 | <0.0975 | <0.0982 | <0.0969 |  |  |  |  | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 | <0.2000 |  |  |  |  | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 141 | <0.0400 | <0.0400 | 0.0400 | 0.04 |  |  |  |  | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 196 | <0.0482 | 0.0526 | 0.0563 | 0.0524 |  |  |  |  | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 164 | <0.0500 | 0.0669 | <0.0500 | 0.0556 |  |  |  |  | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 102 | 0.0703 | 0.0600 | 0.0509 | 0.0604 |  |  |  |  | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 183 | 0.0680 | 0.0750 | 0.0780 | 0.0737 |  |  |  |  | 0 |
|  |  |  |  | Hg (ug/g AR) |  |  |  |  |  |  |  |  |
| Hg (ug/g AR) | 103.04 | ICP-MS | 106 | <0.0080 | <0.0080 | <0.0080 | <0.0080 |  |  |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 141 | <0.0300 | <0.0300 | <0.0300 | <0.0300 |  |  |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 164 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 197 | <0.0969 | <0.0975 | <0.0982 | <0.0969 |  |  |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 | <0.2000 |  |  |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 196 | 0.0077 | 0.0076 | 0.0079 | 0.0077 |  |  |  |  | 0 |
| Hg (ug/g AR) | 103.99 | Other | 102 | 0.0080 | 0.0079 | 0.0081 | 0.008 |  |  |  |  | 0 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 183 | 0.0100 | 0.0100 | 0.0100 | 0.01 |  |  |  |  | 0 |
|  |  |  |  | Pb (ug/g AR) |  |  |  |  |  |  |  |  |
| Pb (ug/g AR) | 104.04 | ICP-MS | 106 | 0.1490 | 0.1570 | 0.1380 | 0.148 | -0.97 | 0.1651 | 7 | 0.0177 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 197 | 0.1509 | 0.1364 | 0.1560 | 0.148 | -0.98 | 0.1651 | 7 | 0.0177 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 102 | 0.1560 | 0.1639 | 0.1501 | 0.157 | -0.48 | 0.1651 | 7 | 0.0177 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Pb (ug/g AR) | 104.04 | ICP-MS | 196 | 0.1670 | 0.1590 | 0.1730 | 0.166 | 0.07 | 0.1651 | 7 | 0.0177 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 141 | 0.1560 | 0.1840 | 0.1680 | 0.169 | 0.24 | 0.1651 | 7 | 0.0177 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 183 | 0.1500 | 0.1670 | 0.2150 | 0.177 | 0.69 | 0.1651 | 7 | 0.0177 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 164 | 0.1623 | 0.2532 | 0.1549 | 0.19 | 1.42 | 0.1651 | 7 | 0.0177 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 162 | <0.2000 | <0.2000 | 0.2100 | 0.203 |  |  |  |  | 1 |

Population of Lab Values

Cu (ug/g AR)

Cu (ug/g AR)

105.04	ICP-MS

196

12.8000  12.0000	11.5000	12.1	0

|  |  |  | Ni (ug/g AR) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Ni (ug/g AR) | 106.04 | ICP-MS | 196 | 1.2100 | 1.1800 | 1.0400 | 1.14 | 0 |
| Ni (ug/g AR) | 106.04 | ICP-MS | 141 | 1.2800 | 1.2200 | 1.2500 | 1.25 | 0 |
| Ni (ug/g AR) | 106.04 | ICP-MS | 183 | 1.5100 | 1.3500 | 1.3600 | 1.41 | 0 |

Cr (ug/g AR)

Cr (ug/g AR) Cr (ug/g AR)

107.04	ICP-MS

107.04	ICP-MS

141

196

<0.8100  <0.7980	<0.7920	<0.8100	1

0.4030	0.3940	0.3950	0.397	0

|  |  |  | Se (ug/g AR) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Se (ug/g AR) | 108.04 | ICP-MS | 141 | <0.4050 | <0.3990 | <0.3960 | <0.4050 | 1 |
| Se (ug/g AR) | 108.04 | ICP-MS | 197 | 0.0976 | <0.0975 | 0.1201 | 0.105 | 0 |
| Se (ug/g AR) | 108.04 | ICP-MS | 196 | 0.1160 | 0.1080 | 0.1200 | 0.115 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Zn (ug/g AR) |  |  |  |  |  |  |  |  |
| Zn (ug/g AR) | 111.04 | ICP-MS | 141 | 59.4000 58.0000 59.1000 |  |  | 58.8 |  |  |  |  | 0 |
| Zn (ug/g AR) | 111.04 | ICP-MS | 196 | 69.5000 72.0000 76.3000 |  |  | 72.6 |  |  |  |  | 0 |
|  |  |  |  | Sb (ug/g AR) |  |  |  |  |  |  |  |  |
| Sb (ug/g AR) | 112.04 | ICP-MS | 196 | <0.0241 <0.0241 <0.0241 |  |  | <0.0241 |  |  |  |  | 1 |
|  |  |  |  | (-)-alpha-Bisabolol (%AR) |  |  |  |  |  |  |  |  |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 162 | 0.0324 | 0.0321 | 0.0358 | 0.0334 |  |  |  |  | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 196 | 0.0881 | 0.0846 | 0.0829 | 0.0852 |  |  |  |  | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 183 | 0.1040 | 0.1000 | 0.1110 | 0.105 |  |  |  |  | 0 |
|  |  |  |  | beta-Caryophyllene (%AR) |  |  |  |  |  |  |  |  |
| beta-Caryophyllene (% | 202.02 | GC,MS | 183 | 0.0431 | 0.0404 | 0.0358 | 0.0398 |  |  |  |  | 0 |
| beta-Caryophyllene (% | 202.02 | GC,MS | 196 | 0.0440 | 0.0424 | 0.0397 | 0.042 |  |  |  |  | 0 |
| beta-Caryophyllene (% | 202.02 | GC,MS | 162 | 0.0472 | 0.0481 | 0.0513 | 0.0489 |  |  |  |  | 0 |
|  |  |  | (-)-Caryophyllene oxide (%AR) |  |  |  |  |  |  |  |  |  |
| (-)-Caryophyllene oxide | 203.02 | GC,MS | 196 | <0.0039 <0.0039 <0.0039 |  |  | <0.0039 |  |  |  |  | 1 |
| (-)-Caryophyllene oxide | 203.02 | GC,MS | 183 | 0.0175 | 0.0153 | 0.0203 | 0.0177 |  |  |  |  | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Geraniol (%AR) |  |  |  |  |  |  |  |  |
| Geraniol (%AR) | 204.02 | GC,MS | 183 | <0.0020 | <0.0020 | <0.0020 | <0.0020 |  |  |  |  | 1 |
| Geraniol (%AR) | 204.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Geraniol (%AR) | 204.02 | GC,MS | 196 | <0.2520 | <0.2520 | <0.2520 | <0.2520 |  |  |  |  | 1 |
|  |  |  |  | alpha-Humulene (%AR) |  |  |  |  |  |  |  |  |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 196 | 0.0187 | 0.0180 | 0.0182 | 0.0183 |  |  |  |  | 0 |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 162 | 0.0199 | 0.0204 | 0.0223 | 0.0209 |  |  |  |  | 0 |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 183 | 0.0248 | 0.0212 | 0.0198 | 0.0219 |  |  |  |  | 0 |
|  |  |  |  | d-Limonene (%AR) |  |  |  |  |  |  |  |  |
| d-Limonene (%AR) | 206.02 | GC,MS | 183 | <0.0020 | <0.0020 | <0.0020 | <0.0020 |  |  |  |  | 1 |
| d-Limonene (%AR) | 206.02 | GC,MS | 196 | <0.0039 | <0.0039 | <0.0039 | <0.0039 |  |  |  |  | 1 |
| d-Limonene (%AR) | 206.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
|  |  |  |  | Linalool (%AR) |  |  |  |  |  |  |  |  |
| Linalool (%AR) | 207.02 | GC,MS | 183 | <0.0020 | <0.0020 | <0.0020 | <0.0020 |  |  |  |  | 1 |
| Linalool (%AR) | 207.02 | GC,MS | 196 | <0.0039 | <0.0039 | <0.0039 | <0.0039 |  |  |  |  | 1 |
| Linalool (%AR) | 207.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
|  |  |  |  | beta-Myrcene (%AR) |  |  |  |  |  |  |  |  |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 183 | <0.0020 <0.0020 <0.0020 |  |  | <0.0020 |  |  |  |  | 1 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 196 | <0.0039 | <0.0039 | <0.0039 | <0.0039 |  |  |  |  | 1 |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
|  |  |  |  | Ocimene (%AR) |  |  |  |  |  |  |  |  |
| Ocimene (%AR) | 210.02 | GC,MS | 183 | <0.0020 | <0.0020 | <0.0020 | <0.0020 |  |  |  |  | 1 |
| Ocimene (%AR) | 210.02 | GC,MS | 196 | <0.0039 | <0.0039 | <0.0039 | <0.0039 |  |  |  |  | 1 |
| Ocimene (%AR) | 210.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
|  |  |  |  | alpha-Pinene (%AR) |  |  |  |  |  |  |  |  |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 183 | <0.0020 | <0.0020 | <0.0020 | <0.0020 |  |  |  |  | 1 |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 196 | <0.0039 | <0.0039 | <0.0039 | <0.0039 |  |  |  |  | 1 |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
|  |  |  |  | (-)-beta-Pinene (%AR) |  |  |  |  |  |  |  |  |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 183 | <0.0020 | <0.0020 | <0.0020 | <0.0020 |  |  |  |  | 1 |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 196 | <0.0039 | <0.0039 | <0.0039 | <0.0039 |  |  |  |  | 1 |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
|  |  |  |  | alpha-Terpinene (%AR) |  |  |  |  |  |  |  |  |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 183 | <0.0020 | <0.0020 | <0.0020 | <0.0020 |  |  |  |  | 1 |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 196 | <0.0039 | <0.0039 | <0.0039 | <0.0039 |  |  |  |  | 1 |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |

Population of Lab Values

| Analyte Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

Population of Lab Values

|  |  |  | gamma-Terpinene (%AR) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| gamma-Terpinene (%A | 214.02 | GC,MS | 183 | <0.0020 | <0.0020 | <0.0020 | <0.0020 | 1 |
| gamma-Terpinene (%A | 214.02 | GC,MS | 196 | <0.0039 | <0.0039 | <0.0039 | <0.0039 | 1 |
| gamma-Terpinene (%A | 214.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |

|  |  |  | alpha-Terpinolene (%AR) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 183 | <0.0020 | <0.0020 | <0.0020 | <0.0020 | 1 |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 196 | <0.0039 | <0.0039 | <0.0039 | <0.0039 | 1 |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |

| Moisture (%AR) |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 153 | 3.0300 | 3.1600 | 2.7600 | 2.98 | -1.77 | 6.9029 | 15 | 2.2096 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 189 | 4.6060 | 4.6060 | 4.6060 | 4.61 | -1.04 | 6.9029 | 15 | 2.2096 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 116 | 5.1600 | 4.8600 | 4.8800 | 4.97 | -0.88 | 6.9029 | 15 | 2.2096 | 0 |
| Moisture (%AR) | 500.99 | Other | 102 | 5.0000 | 5.7000 | 5.2000 | 5.3 | -0.73 | 6.9029 | 15 | 2.2096 | 0 |
| Moisture (%AR) | 500.99 | Other | 103 | 5.8700 | 6.0900 | 5.8600 | 5.94 | -0.44 | 6.9029 | 15 | 2.2096 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 106 | 5.9300 | 6.0100 | 6.0500 | 6 | -0.41 | 6.9029 | 15 | 2.2096 | 0 |
| Moisture (%AR) | 500.72 | NIR | 138 | 6.0000 | 5.8000 | 6.2000 | 6 | -0.41 | 6.9029 | 15 | 2.2096 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 196 | 6.5800 | 6.5300 | 6.9200 | 6.68 | -0.1 | 6.9029 | 15 | 2.2096 | 0 |
| Moisture (%AR) | 500.99 | Other | 165 | 6.9600 | 6.9700 | 6.9800 | 6.97 | 0.03 | 6.9029 | 15 | 2.2096 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 210 | 7.6300 | 7.6800 | 8.2000 | 7.84 | 0.42 | 6.9029 | 15 | 2.2096 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 162 | 8.0000 | 8.6500 | 7.8400 | 8.16 | 0.57 | 6.9029 | 15 | 2.2096 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 182 | 8.5200 | 9.0600 | 8.8900 | 8.82 | 0.87 | 6.9029 | 15 | 2.2096 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 192 | 8.7865 | 9.0468 | 8.8719 | 8.9 | 0.9 | 6.9029 | 15 | 2.2096 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 199 | 10.0275 | 9.6195 | 9.9447 | 9.86 | 1.34 | 6.9029 | 15 | 2.2096 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 183 | 10.4200 | 10.6500 | 8.6600 | 9.91 | 1.36 | 6.9029 | 15 | 2.2096 | 0 |
|  |  |  |  | Δ9-THC (%DW) |  |  |  |  |  |  |  |  |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, U | 103 | 0.2164 | 0.1984 | 0.1921 | 0.202 | -1.25 | 0.2246 | 8 | 0.0177 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 165 | 0.2094 | 0.2103 | 0.2246 | 0.215 | -0.55 | 0.2246 | 8 | 0.0177 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 116 | 0.2165 | 0.2130 | 0.2242 | 0.218 | -0.38 | 0.2246 | 8 | 0.0177 | 0 |
| Δ9-THC (%DW) | 501.30 | Other LC-MS | 102 | 0.2140 | 0.2240 | 0.2220 | 0.22 | -0.26 | 0.2246 | 8 | 0.0177 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 162 | 0.2142 | 0.2271 | 0.2239 | 0.222 | -0.16 | 0.2246 | 8 | 0.0177 | 0 |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, U | 199 | 0.2212 | 0.2202 | 0.2349 | 0.225 | 0.05 | 0.2246 | 8 | 0.0177 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 189 | 0.2467 | 0.2425 | 0.2419 | 0.244 | 1.08 | 0.2246 | 8 | 0.0177 | 0 |
| Δ9-THC (%DW) | 501.01 | AOAC 2018.10 | 210 | 0.2570 | 0.2490 | 0.2460 | 0.251 | 1.47 | 0.2246 | 8 | 0.0177 | 0 |
|  |  |  |  | Δ9-THCA (%DW) |  |  |  |  |  |  |  |  |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, U | 199 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, U | 103 | 0.0462 | 0.0428 | 0.0594 | 0.0495 | -1.48 | 0.0572 | 7 | 0.0053 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 162 | 0.0582 | 0.0482 | 0.0579 | 0.0548 | -0.47 | 0.0572 | 7 | 0.0053 | 0 |
| Δ9-THCA (%DW) | 502.01 | AOAC 2018.10 | 210 | 0.0550 | 0.0570 | 0.0550 | 0.0557 | -0.3 | 0.0572 | 7 | 0.0053 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 189 | 0.0590 | 0.0582 | 0.0579 | 0.0584 | 0.21 | 0.0572 | 7 | 0.0053 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 116 | 0.0564 | 0.0578 | 0.0601 | 0.0581 | 0.16 | 0.0572 | 7 | 0.0053 | 0 |
| Δ9-THCA (%DW) | 502.30 | Other LC-MS | 102 | 0.0570 | 0.0610 | 0.0620 | 0.06 | 0.52 | 0.0572 | 7 | 0.0053 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 165 | 0.0612 | 0.0632 | 0.0687 | 0.0644 | 1.35 | 0.0572 | 7 | 0.0053 | 0 |
|  |  |  |  | CBD (%DW) |  |  |  |  |  |  |  |  |
| CBD (%DW) | 503.02 | AOAC 2018.11, U | 103 | 5.1079 | 4.9566 | 4.8888 | 4.98 | -4.08 | 5.4263 | 6 | 0.1083 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 162 | 5.5508 | 5.5126 | 5.1173 | 5.39 | -0.3 | 5.4263 | 6 | 0.1083 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 189 | 5.4367 | 5.4191 | 5.4128 | 5.42 | -0.03 | 5.4263 | 6 | 0.1083 | 0 |
| CBD (%DW) | 503.01 | AOAC 2018.10 | 210 | 5.4510 | 5.5530 | 5.3270 | 5.44 | 0.16 | 5.4263 | 6 | 0.1083 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 116 | 5.4648 | 5.3033 | 5.7013 | 5.49 | 0.59 | 5.4263 | 6 | 0.1083 | 0 |
| CBD (%DW) | 503.02 | AOAC 2018.11, U | 199 | 5.4917 | 5.4746 | 5.6649 | 5.54 | 1.08 | 5.4263 | 6 | 0.1083 | 0 |
|  |  |  |  | CBDA (%DW) |  |  |  |  |  |  |  |  |
| CBDA (%DW) | 504.10 | Other LC-UV | 116 | 3.7535 | 4.0156 | 4.0314 | 3.93 | -0.77 | 4.1835 | 6 | 0.3267 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 189 | 4.0225 | 4.0076 | 3.9978 | 4.01 | -0.53 | 4.1835 | 6 | 0.3267 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 162 | 4.0004 | 4.2261 | 3.8562 | 4.03 | -0.48 | 4.1835 | 6 | 0.3267 | 0 |
| CBDA (%DW) | 504.02 | AOAC 2018.11, U | 103 | 4.2468 | 4.0307 | 3.9150 | 4.06 | -0.37 | 4.1835 | 6 | 0.3267 | 0 |
| CBDA (%DW) | 504.02 | AOAC 2018.11, U | 199 | 4.3241 | 4.3676 | 4.4878 | 4.39 | 0.64 | 4.1835 | 6 | 0.3267 | 0 |
| CBDA (%DW) | 504.01 | AOAC 2018.10 | 210 | 4.9300 | 4.8720 | 4.8190 | 4.87 | 2.11 | 4.1835 | 6 | 0.3267 | 0 |
|  |  |  |  | CBN (%DW) |  |  |  |  |  |  |  |  |
| CBN (%DW) | 505.10 | Other LC-UV | 192 | 0.0340 | 0.0370 | 0.0359 | 0.0356 | -2.78 | 0.0544 | 7 | 0.0067 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%DW) | 505.30 | GC-FID | 192 | 0.0488 | 0.0492 | 0.0506 | 0.0495 | -0.72 | 0.0544 | 7 | 0.0067 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 162 | 0.0506 | 0.0541 | 0.0540 | 0.0529 | -0.22 | 0.0544 | 7 | 0.0067 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 189 | 0.0556 | 0.0563 | 0.0575 | 0.0565 | 0.31 | 0.0544 | 7 | 0.0067 | 0 |
| CBN (%DW) | 505.02 | AOAC 2018.11, U | 199 | 0.0561 | 0.0553 | 0.0583 | 0.0566 | 0.32 | 0.0544 | 7 | 0.0067 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 116 | 0.0515 | 0.0627 | 0.0654 | 0.0599 | 0.81 | 0.0544 | 7 | 0.0067 | 0 |
| CBN (%DW) | 505.30 | GC-FID | 116 | 0.0641 | 0.0598 | 0.0594 | 0.0611 | 1 | 0.0544 | 7 | 0.0067 | 0 |
|  |  |  |  | Total Δ9-THC (%DW) |  |  |  |  |  |  |  |  |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 192 | 0.1862 | 0.1763 | 0.1788 | 0.18 | -2.51 | 0.257 | 14 | 0.0306 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 182 | 0.2277 | 0.2140 | 0.2086 | 0.217 | -1.32 | 0.257 | 14 | 0.0306 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 192 | 0.2214 | 0.2145 | 0.2211 | 0.219 | -1.24 | 0.257 | 14 | 0.0306 | 0 |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, U | 103 | 0.2569 | 0.2359 | 0.2441 | 0.246 | -0.37 | 0.257 | 14 | 0.0306 | 0 |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, U | 199 | 0.2431 | 0.2421 | 0.2568 | 0.247 | -0.32 | 0.257 | 14 | 0.0306 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 138 | 0.2503 | 0.2610 | 0.2543 | 0.255 | -0.06 | 0.257 | 14 | 0.0306 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 153 | 0.2630 | 0.2730 | 0.2520 | 0.263 | 0.18 | 0.257 | 14 | 0.0306 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 116 | 0.2771 | 0.2592 | 0.2567 | 0.264 | 0.24 | 0.257 | 14 | 0.0306 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 116 | 0.2659 | 0.2637 | 0.2768 | 0.269 | 0.38 | 0.257 | 14 | 0.0306 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 162 | 0.2652 | 0.2694 | 0.2747 | 0.27 | 0.42 | 0.257 | 14 | 0.0306 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 165 | 0.2630 | 0.2658 | 0.2849 | 0.271 | 0.46 | 0.257 | 14 | 0.0306 | 0 |
| Total Δ9-THC (%DW) | 506.30 | Other LC-MS | 102 | 0.2640 | 0.2770 | 0.2760 | 0.272 | 0.5 | 0.257 | 14 | 0.0306 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 189 | 0.2984 | 0.2935 | 0.2928 | 0.295 | 1.24 | 0.257 | 14 | 0.0306 | 0 |
| Total Δ9-THC (%DW) | 506.01 | AOAC 2018.10 | 210 | 0.3050 | 0.2990 | 0.2940 | 0.299 | 1.38 | 0.257 | 14 | 0.0306 | 0 |

Population of Lab Values

| Analyte Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score Rob Mean |  | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

Population of Lab Values

| Total CBD (%DW) |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%DW) | 507.40 | GC-FID | 192 | 7.6762 | 7.1079 | 7.5467 | 7.44 | -1.81 | 8.9644 | 10 | 0.8388 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 192 | 7.7646 | 8.3285 | 8.1393 | 8.08 | -1.06 | 8.9644 | 10 | 0.8388 | 0 |
| Total CBD (%DW) | 507.02 | AOAC 2018.11, U | 103 | 8.8324 | 8.4914 | 8.3223 | 8.55 | -0.5 | 8.9644 | 10 | 0.8388 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 162 | 9.0592 | 9.2189 | 8.4992 | 8.93 | -0.05 | 8.9644 | 10 | 0.8388 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 189 | 8.9644 | 8.9338 | 8.9189 | 8.94 | -0.03 | 8.9644 | 10 | 0.8388 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 116 | 8.7567 | 8.8250 | 9.2369 | 8.94 | -0.03 | 8.9644 | 10 | 0.8388 | 0 |
| Total CBD (%DW) | 507.40 | GC-FID | 116 | 9.6090 | 8.8512 | 9.0532 | 9.17 | 0.25 | 8.9644 | 10 | 0.8388 | 0 |
| Total CBD (%DW) | 507.02 | AOAC 2018.11, U | 199 | 9.2839 | 9.3050 | 9.6007 | 9.4 | 0.52 | 8.9644 | 10 | 0.8388 | 0 |
| Total CBD (%DW) | 507.01 | AOAC 2018.10 | 210 | 9.7740 | 9.8250 | 9.5530 | 9.72 | 0.9 | 8.9644 | 10 | 0.8388 | 0 |
| Total CBD (%DW) | 507.40 | GC-FID | 138 | 10.1690 | 10.1887 | 10.7757 | 10.4 | 1.68 | 8.9644 | 10 | 0.8388 | 0 |

| CBDV (%DW) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDV (%DW) | 508.99 | Other | 189 | 0.0321 | 0.0336 | 0.0336 | 0.0331 | 0 |
| CBDV (%DW) | 508.99 | Other | 162 | 0.0396 | 0.0335 | 0.0378 | 0.037 | 0 |
| CBDV (%DW) | 508.02 | AOAC 2018.11, U | 210 | 0.0380 | 0.0380 | 0.0370 | 0.0377 | 0 |

| CBDVA (%DW) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDVA (%DW) | 509.01 | AOAC 2018.10 | 210 | 0.0340 | 0.0300 | 0.0350 | 0.033 | 0 |

| Analyte | Code | Method | Lab Num |  | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | CBG (%DW) |  |  |  |  |  |  |  |  |  |  |
| CBG (%DW) | 510.01 | AOAC 2018.10 |  | 210 | 0.2260 | 0.2380 | 0.2300 | 0.231 | 0 |  |  |  |  |
| CBG (%DW) | 510.99 | Other |  | 162 | 0.2323 | 0.2301 | 0.2351 | 0.233 | 0 |  |  |  |  |
| CBG (%DW) | 510.99 | Other |  | 189 | 0.2903 | 0.2906 | 0.2999 | 0.294 | 0 |  |  |  |  |

Population of Lab Values

| CBGA (%DW) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%DW) | 511.01 | AOAC 2018.10 | 210 | 0.1050 | 0.1150 | 0.1200 | 0.113 | 0 |
| CBGA (%DW) | 511.99 | Other | 162 | 0.2194 | 0.2161 | 0.2076 | 0.214 | 0 |
| CBGA (%DW) | 511.99 | Other | 189 | 0.2233 | 0.2178 | 0.2181 | 0.22 | 0 |

| THCV (%DW) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| THCV (%DW) | 512.99 | Other | 189 | <0.0050 <0.0050 <0.0050 | <0.0050 | 1 |
| THCV (%DW) | 512.02 | AOAC 2018.11, U | 210 | <0.0130 <0.0130 <0.0130 | <0.0130 | 1 |

| Δ8-THC (%DW) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%DW) | 513.99 | Other | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| Δ8-THC (%DW) | 513.99 | Other | 189 | 0.0132 | 0.0094 | 0.0147 | 0.0124 | 0 |
| Δ8-THC (%DW) | 513.01 | AOAC 2018.10 | 210 | 0.0140 | 0.0140 | <0.0140 | 0.014 | 0 |

|  |  |  | CBC (%DW) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%DW) | 514.99 | Other | 162 | 0.2512 | 0.2707 | 0.2636 | 0.262 | 0 |
| CBC (%DW) | 514.99 | Other | 189 | 0.3114 | 0.2946 | 0.2943 | 0.3 | 0 |

| Analyte Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

Population of Lab Values

- Flag definitions: 1 (Lab Value not used because there were less than 2 numeric lab results), 2 (Lab Value not used because there were less than 2 numeric lab results greater than 0).

- Robust statistics performed to determine robust mean and Z scores only if there were 6 or more numeric lab values.

- AR refers to concentration on as-received basis. DW refers to concentration on dry weight basis.