Hemp Proficiency Testing

All Labs Precision HM24SEP-1 (hemp)

Sample: HM24SEP-1 (hemp)

All Labs Precision

Issue Date:

ovember 11, 2024

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

Δ9-THC (%AR)

RSDr	HorRat(r)

Population of Lab RSDr	 Rob Mean	min	max

Flag

| Δ9-THC (%AR) | 001.30 | Other LC-MS | 122 | <0.5000 | <0.5000 | <0.5000 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 183 | 0.1852 | 0.1842 | 0.1837 | 0.41 | 0.08 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 199 | 0.1830 | 0.1845 | 0.1845 | 0.47 | 0.09 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 101 | 0.1793 | 0.1785 | 0.1807 | 0.62 | 0.12 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 213 | 0.0750 | 0.0760 | 0.0750 | 0.77 | 0.13 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 189 | 0.1844 | 0.1869 | 0.1844 | 0.78 | 0.15 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 164 | 0.2167 | 0.2138 | 0.2166 | 0.76 | 0.15 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.99 | Other | 183 | 0.1737 | 0.1763 | 0.1736 | 0.88 | 0.17 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 144 | 0.1634 | 0.1588 | 0.1606 | 1.44 | 0.27 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 125 | 0.1523 | 0.1551 | 0.1507 | 1.46 | 0.28 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 163 | 0.1740 | 0.1740 | 0.1790 | 1.64 | 0.32 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 170 | 0.1602 | 0.1581 | 0.1634 | 1.66 | 0.32 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 149 | 0.1460 | 0.1450 | 0.1410 | 1.84 | 0.34 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 148 | 0.1420 | 0.1370 | 0.1380 | 1.9 | 0.35 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 228 | 0.2421 | 0.2511 | 0.2444 | 1.9 | 0.39 | 3.01 | 0.41 | 29.6 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 196 | 0.1619 | 0.1682 | 0.1676 | 2.1 | 0.4 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 115 | 0.1675 | 0.1594 | 0.1640 | 2.48 | 0.47 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 116 | 0.1585 | 0.1656 | 0.1654 | 2.48 | 0.47 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 217 | 0.1707 | 0.1798 | 0.1758 | 2.6 | 0.5 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 197 | 0.1538 | 0.1600 | 0.1624 | 2.8 | 0.53 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 165 | 0.1521 | 0.1608 | 0.1551 | 2.83 | 0.54 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 149 | 0.1610 | 0.1630 | 0.1530 | 3.33 | 0.63 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 102 | 0.1570 | 0.1550 | 0.1650 | 3.33 | 0.63 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 103 | 0.1415 | 0.1319 | 0.1374 | 3.52 | 0.65 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 106 | 0.1500 | 0.1400 | 0.1500 | 3.94 | 0.74 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 182 | 0.1313 | 0.1219 | 0.1202 | 4.8 | 0.88 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 161 | 0.1360 | 0.1316 | 0.1234 | 4.91 | 0.91 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 184 | 0.1777 | 0.1914 | 0.1959 | 5.03 | 0.98 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.03 | AOAC 2018.11, M | 139 | 0.2111 | 0.2273 | 0.2017 | 6.07 | 1.21 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 175 | 0.1590 | 0.1440 | 0.1660 | 7.19 | 1.36 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 173 | 0.1900 | 0.1790 | 0.1610 | 8.29 | 1.6 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 155 | 0.1740 | 0.1730 | 0.1440 | 10.4 | 1.99 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 141 | 0.1540 | 0.1543 | 0.2320 | 25 | 4.84 | 3.01 | 0.41 | 29.6 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 192 | 0.1632 | 0.1745 | 0.2728 | 29.6 | 5.85 | 3.01 | 0.41 | 29.6 | 0 |
|  |  |  |  | Δ9-THCA (%AR) |  |  |  |  |  |  |  |  |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 102 | <0.0500 <0.0480 |  | <0.0500 |  |  |  |  |  | 1 |

Population of Lab RSDr

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 173 | <0.1590 | <0.1500 | <0.1580 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 197 | <0.0503 | <0.0526 | <0.0508 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 199 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 161 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 122 | <0.5000 | <0.5000 | <0.5000 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 106 | 0.0400 | 0.0400 | 0.0400 | 0 | 0 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 184 | 0.0481 | 0.0492 | 0.0487 | 1.13 | 0.18 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.99 | Other | 183 | 0.0530 | 0.0540 | 0.0529 | 1.14 | 0.18 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 115 | 0.0451 | 0.0437 | 0.0443 | 1.58 | 0.25 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 164 | 0.0586 | 0.0572 | 0.0568 | 1.64 | 0.27 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 213 | 0.0106 | 0.0104 | 0.0100 | 2.96 | 0.37 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 189 | 0.0444 | 0.0468 | 0.0464 | 2.8 | 0.44 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 196 | 0.0419 | 0.0437 | 0.0412 | 3.05 | 0.48 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 103 | 0.0340 | 0.0350 | 0.0362 | 3.14 | 0.48 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 163 | 0.0260 | 0.0270 | 0.0280 | 3.7 | 0.54 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 144 | 0.0478 | 0.0483 | 0.0511 | 3.62 | 0.58 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 149 | 0.0380 | 0.0370 | 0.0350 | 4.17 | 0.64 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 116 | 0.0480 | 0.0465 | 0.0440 | 4.38 | 0.69 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 165 | 0.0403 | 0.0442 | 0.0427 | 4.64 | 0.72 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 148 | 0.0560 | 0.0530 | 0.0510 | 4.72 | 0.76 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 217 | 0.0492 | 0.0443 | 0.0464 | 5.27 | 0.83 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 183 | 0.0571 | 0.0518 | 0.0531 | 5.12 | 0.83 | 5.4 | 0 | 29.8 | 0 |

Lab

Lab

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 155 | 0.0528 | 0.0481 | 0.0530 | 5.41 | 0.87 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 192 | 0.2353 | 0.2187 | 0.2368 | 4.36 | 0.88 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 101 | 0.0548 | 0.0485 | 0.0509 | 6.19 | 0.99 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 149 | 0.0490 | 0.0410 | 0.0470 | 9.12 | 1.44 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 170 | 0.0469 | 0.0507 | 0.0563 | 9.22 | 1.48 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.03 | AOAC 2018.11, M | 139 | 0.0591 | 0.0436 | 0.0522 | 15 | 2.42 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 125 | 0.0512 | 0.0508 | 0.0701 | 19.2 | 3.14 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 182 | 0.0379 | 0.0323 | 0.0221 | 26 | 3.87 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 175 | 0.0628 | 0.0465 | 0.0376 | 26.1 | 4.16 | 5.4 | 0 | 29.8 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 141 | 0.0580 | 0.0543 | 0.0910 | 29.8 | 4.99 | 5.4 | 0 | 29.8 | 0 |
|  |  |  |  | CBD (%AR) |  |  |  |  |  |  |  |  |
| CBD (%AR) | 003.10 | Other LC-UV | 189 | 4.3703 | 4.3600 | 4.3665 | 0.12 | 0.04 | 2.02 | 0.12 | 8.71 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 199 | 4.5375 | 4.5825 | 4.5845 | 0.58 | 0.18 | 2.02 | 0.12 | 8.71 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 163 | 4.6700 | 4.6140 | 4.6320 | 0.62 | 0.19 | 2.02 | 0.12 | 8.71 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 164 | 4.5093 | 4.5215 | 4.5655 | 0.65 | 0.21 | 2.02 | 0.12 | 8.71 | 0 |
| CBD (%AR) | 003.01 | AOAC 2018.10 | 228 | 8.7542 | 8.6955 | 8.6509 | 0.6 | 0.21 | 2.02 | 0.12 | 8.71 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 116 | 4.3713 | 4.3908 | 4.4512 | 0.95 | 0.3 | 2.02 | 0.12 | 8.71 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 213 | 4.6830 | 4.7560 | 4.6530 | 1.13 | 0.36 | 2.02 | 0.12 | 8.71 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 125 | 4.3237 | 4.3981 | 4.4314 | 1.26 | 0.39 | 2.02 | 0.12 | 8.71 | 0 |
| CBD (%AR) | 003.03 | AOAC 2018.11, M | 139 | 4.4814 | 4.3578 | 4.4127 | 1.4 | 0.44 | 2.02 | 0.12 | 8.71 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 155 | 3.9200 | 3.8600 | 3.9800 | 1.53 | 0.47 | 2.02 | 0.12 | 8.71 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%AR) | 003.99 | Other | 183 | 4.5903 | 4.5662 | 4.4530 | 1.62 | 0.51 | 2.02 | 0.12 | 8.71 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 161 | 3.7932 | 3.6942 | 3.6719 | 1.74 | 0.53 | 2.02 | 0.12 | 8.71 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 106 | 4.3400 | 4.3200 | 4.4700 | 1.86 | 0.58 | 2.02 | 0.12 | 8.71 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 197 | 4.4643 | 4.6063 | 4.6264 | 1.94 | 0.61 | 2.02 | 0.12 | 8.71 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 183 | 4.5180 | 4.3470 | 4.3514 | 2.21 | 0.69 | 2.02 | 0.12 | 8.71 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 173 | 3.7660 | 3.7060 | 3.6000 | 2.28 | 0.69 | 2.02 | 0.12 | 8.71 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 149 | 4.0020 | 3.8330 | 3.8340 | 2.5 | 0.77 | 2.02 | 0.12 | 8.71 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 196 | 4.2940 | 4.1306 | 4.3347 | 2.54 | 0.79 | 2.02 | 0.12 | 8.71 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 170 | 4.5568 | 4.3604 | 4.3377 | 2.73 | 0.85 | 2.02 | 0.12 | 8.71 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 184 | 4.3107 | 4.2130 | 4.0723 | 2.85 | 0.89 | 2.02 | 0.12 | 8.71 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 149 | 3.2590 | 3.4340 | 3.4920 | 3.57 | 1.08 | 2.02 | 0.12 | 8.71 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 192 | 6.1850 | 5.7990 | 6.1804 | 3.66 | 1.2 | 2.02 | 0.12 | 8.71 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 144 | 4.5317 | 4.9023 | 4.7898 | 4.01 | 1.27 | 2.02 | 0.12 | 8.71 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 103 | 4.2406 | 3.9418 | 3.7236 | 6.54 | 2.02 | 2.02 | 0.12 | 8.71 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 141 | 4.2530 | 4.1948 | 4.8930 | 8.71 | 2.73 | 2.02 | 0.12 | 8.71 | 0 |
|  |  |  |  | CBDA (%AR) |  |  |  |  |  |  |  |  |
| CBDA (%AR) | 004.10 | Other LC-UV | 189 | 3.2000 | 3.1932 | 3.1987 | 0.11 | 0.03 | 2.55 | 0.11 | 7.55 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 155 | 2.8800 | 2.8800 | 2.8700 | 0.2 | 0.06 | 2.55 | 0.11 | 7.55 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 183 | 3.7252 | 3.7068 | 3.7246 | 0.28 | 0.09 | 2.55 | 0.11 | 7.55 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 199 | 3.3060 | 3.3055 | 3.3305 | 0.43 | 0.13 | 2.55 | 0.11 | 7.55 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 163 | 2.7910 | 2.7580 | 2.7470 | 0.83 | 0.24 | 2.55 | 0.11 | 7.55 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDA (%AR) | 004.10 | Other LC-UV | 213 | 3.2740 | 3.2730 | 3.2230 | 0.9 | 0.27 | 2.55 | 0.11 | 7.55 | 0 |
| CBDA (%AR) | 004.03 | AOAC 2018.11, M | 139 | 4.1585 | 4.1609 | 4.0927 | 0.94 | 0.29 | 2.55 | 0.11 | 7.55 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 106 | 3.3500 | 3.2900 | 3.3000 | 0.97 | 0.29 | 2.55 | 0.11 | 7.55 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 116 | 3.2654 | 3.2187 | 3.1719 | 1.45 | 0.43 | 2.55 | 0.11 | 7.55 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 196 | 3.2921 | 3.1923 | 3.2248 | 1.57 | 0.47 | 2.55 | 0.11 | 7.55 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 173 | 2.7120 | 2.6910 | 2.5970 | 2.3 | 0.67 | 2.55 | 0.11 | 7.55 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 103 | 3.0659 | 3.0094 | 2.9111 | 2.62 | 0.77 | 2.55 | 0.11 | 7.55 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 197 | 3.1849 | 3.3895 | 3.3259 | 3.17 | 0.95 | 2.55 | 0.11 | 7.55 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 161 | 3.0266 | 2.8417 | 2.9901 | 3.32 | 0.98 | 2.55 | 0.11 | 7.55 | 0 |
| CBDA (%AR) | 004.99 | Other | 149 | 2.9070 | 2.9900 | 3.1080 | 3.37 | 0.99 | 2.55 | 0.11 | 7.55 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 192 | 4.0158 | 4.0186 | 4.2494 | 3.27 | 1.01 | 2.55 | 0.11 | 7.55 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 149 | 2.6810 | 2.7420 | 2.8720 | 3.53 | 1.03 | 2.55 | 0.11 | 7.55 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 125 | 3.2155 | 3.4319 | 3.2040 | 3.91 | 1.17 | 2.55 | 0.11 | 7.55 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 164 | 4.9922 | 5.2989 | 5.3701 | 3.85 | 1.24 | 2.55 | 0.11 | 7.55 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 144 | 3.5427 | 3.3691 | 3.7281 | 5.06 | 1.53 | 2.55 | 0.11 | 7.55 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 184 | 3.2905 | 3.6403 | 3.5776 | 5.32 | 1.61 | 2.55 | 0.11 | 7.55 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 170 | 3.2607 | 3.6766 | 3.6880 | 6.87 | 2.08 | 2.55 | 0.11 | 7.55 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 141 | 3.3360 | 3.2480 | 3.7350 | 7.55 | 2.28 | 2.55 | 0.11 | 7.55 | 0 |
|  |  |  |  | CBN (%AR) |  |  |  |  |  |  |  |  |
| CBN (%AR) | 005.10 | Other LC-UV | 173 | <0.1590 <0.1500 |  | <0.1580 |  |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 197 | <0.0503 <0.0526 |  | <0.0508 |  |  |  |  |  | 1 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%AR) | 005.10 | Other LC-UV | 106 | 0.0400 | 0.0400 | 0.0400 | 0 | 0 | 5.68 | 0 | 86.9 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 149 | 0.0350 | 0.0350 | 0.0350 | 0 | 0 | 5.68 | 0 | 86.9 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 189 | 0.0450 | 0.0448 | 0.0446 | 0.45 | 0.07 | 5.68 | 0 | 86.9 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 183 | 0.0401 | 0.0406 | 0.0404 | 0.62 | 0.1 | 5.68 | 0 | 86.9 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 144 | 0.0384 | 0.0376 | 0.0394 | 2.34 | 0.36 | 5.68 | 0 | 86.9 | 0 |
| CBN (%AR) | 005.99 | Other | 183 | 0.0419 | 0.0400 | 0.0414 | 2.4 | 0.37 | 5.68 | 0 | 86.9 | 0 |
| CBN (%AR) | 005.03 | AOAC 2018.11, M | 139 | 0.0419 | 0.0399 | 0.0405 | 2.52 | 0.39 | 5.68 | 0 | 86.9 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 213 | 0.0380 | 0.0400 | 0.0390 | 2.56 | 0.39 | 5.68 | 0 | 86.9 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 163 | 0.0620 | 0.0580 | 0.0620 | 3.81 | 0.63 | 5.68 | 0 | 86.9 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 116 | 0.0533 | 0.0571 | 0.0573 | 4.03 | 0.66 | 5.68 | 0 | 86.9 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 196 | 0.0398 | 0.0362 | 0.0390 | 4.93 | 0.76 | 5.68 | 0 | 86.9 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 125 | 0.0536 | 0.0572 | 0.0519 | 4.99 | 0.81 | 5.68 | 0 | 86.9 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 192 | 0.0709 | 0.0789 | 0.0792 | 6.17 | 1.05 | 5.68 | 0 | 86.9 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 199 | 0.0500 | 0.0500 | 0.0445 | 6.59 | 1.05 | 5.68 | 0 | 86.9 | 0 |
| CBN (%AR) | 005.20 | Other LC-MS | 149 | 0.0420 | 0.0360 | 0.0330 | 12.4 | 1.89 | 5.68 | 0 | 86.9 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 138 | 0.0507 | 0.0645 | 0.0593 | 12 | 1.96 | 5.68 | 0 | 86.9 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 116 | 0.0452 | 0.0582 | 0.0476 | 13.7 | 2.2 | 5.68 | 0 | 86.9 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 141 | 0.0440 | 0.0444 | 0.0560 | 14.2 | 2.25 | 5.68 | 0 | 86.9 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 192 | 0.0000 | 0.0485 | 0.0530 | 86.9 | 13.1 | 5.68 | 0 | 86.9 | 0 |
|  |  |  |  | Total Δ9-THC (%AR) |  |  |  |  |  |  |  |  |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 190 | 0.2413 | 0.2186 |  |  |  |  |  |  | 1 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 114 | 0.2123 | 0.1794 |  |  |  |  |  |  | 1 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 199 | 0.2049 | 0.2064 | 0.2064 | 0.42 | 0.08 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 213 | 0.0840 | 0.0850 | 0.0840 | 0.68 | 0.12 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 164 | 0.2681 | 0.2640 | 0.2665 | 0.78 | 0.16 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.99 | Other | 183 | 0.2201 | 0.2236 | 0.2199 | 0.94 | 0.19 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 189 | 0.2234 | 0.2279 | 0.2250 | 1.01 | 0.2 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 217 | 0.2139 | 0.2187 | 0.2164 | 1.11 | 0.22 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 192 | 0.1866 | 0.1877 | 0.1834 | 1.2 | 0.23 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 144 | 0.2053 | 0.2012 | 0.2054 | 1.17 | 0.23 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 163 | 0.2200 | 0.2150 | 0.2180 | 1.16 | 0.23 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 183 | 0.2353 | 0.2296 | 0.2303 | 1.34 | 0.27 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 116 | 0.2006 | 0.2063 | 0.2040 | 1.41 | 0.28 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 101 | 0.2274 | 0.2211 | 0.2253 | 1.43 | 0.29 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 196 | 0.1986 | 0.2066 | 0.2038 | 2 | 0.39 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 149 | 0.1790 | 0.1770 | 0.1720 | 2.05 | 0.4 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 138 | 0.1948 | 0.1880 | 0.1953 | 2.12 | 0.41 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 115 | 0.2071 | 0.1978 | 0.2028 | 2.3 | 0.45 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 149 | 0.2040 | 0.1990 | 0.1940 | 2.51 | 0.49 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 148 | 0.1920 | 0.1840 | 0.1830 | 2.65 | 0.52 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 103 | 0.1713 | 0.1626 | 0.1691 | 2.7 | 0.52 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 197 | 0.1538 | 0.1600 | 0.1624 | 2.8 | 0.53 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 107 | 0.4540 | 0.4570 | 0.4751 | 2.47 | 0.55 | 3.49 | 0.42 | 26.2 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 102 | 0.1740 | 0.1720 | 0.1820 | 3.01 | 0.58 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 165 | 0.1924 | 0.2050 | 0.1978 | 3.19 | 0.63 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 116 | 0.1923 | 0.2050 | 0.2033 | 3.44 | 0.68 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.03 | AOAC 2018.11, M | 139 | 0.2629 | 0.2655 | 0.2475 | 3.76 | 0.77 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 169 | 0.1773 | 0.1684 | 0.1640 | 3.99 | 0.77 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 125 | 0.1972 | 0.1997 | 0.2122 | 3.96 | 0.78 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 170 | 0.2620 | 0.2767 | 0.3002 | 6.89 | 1.43 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 175 | 0.2146 | 0.1852 | 0.1990 | 7.37 | 1.45 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 155 | 0.2200 | 0.2150 | 0.1910 | 7.43 | 1.47 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 182 | 0.1645 | 0.1502 | 0.1396 | 8.25 | 1.56 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 173 | 0.2391 | 0.2299 | 0.2031 | 8.35 | 1.67 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.50 | GC-MS | 170 | 0.1971 | 0.2257 | 0.2438 | 10.6 | 2.12 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 153 | 0.1655 | 0.2232 | 0.2064 | 15 | 2.94 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 192 | 0.3696 | 0.3664 | 0.4806 | 16 | 3.51 | 3.49 | 0.42 | 26.2 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 141 | 0.2050 | 0.2019 | 0.3120 | 26.2 | 5.29 | 3.49 | 0.42 | 26.2 | 0 |
|  |  |  |  | Total CBD (%AR) |  |  |  |  |  |  |  |  |
| Total CBD (%AR) | 007.10 | Other LC-UV | 189 | 7.1767 | 7.1605 | 7.1717 | 0.12 | 0.04 | 2.3 | 0.12 | 9.09 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 116 | 7.2350 | 7.2136 | 7.2329 | 0.16 | 0.06 | 2.3 | 0.12 | 9.09 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 163 | 7.5680 | 7.5670 | 7.5160 | 0.39 | 0.13 | 2.3 | 0.12 | 9.09 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 199 | 7.4369 | 7.4814 | 7.5053 | 0.46 | 0.16 | 2.3 | 0.12 | 9.09 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 138 | 8.6617 | 8.6419 | 8.7477 | 0.65 | 0.22 | 2.3 | 0.12 | 9.09 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%AR) | 007.40 | GC-FID | 169 | 6.8740 | 6.9100 | 6.8130 | 0.71 | 0.24 | 2.3 | 0.12 | 9.09 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 155 | 6.4400 | 6.3900 | 6.5000 | 0.86 | 0.28 | 2.3 | 0.12 | 9.09 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 149 | 6.5520 | 6.4550 | 6.5600 | 0.9 | 0.3 | 2.3 | 0.12 | 9.09 | 0 |
| Total CBD (%AR) | 007.03 | AOAC 2018.11, M | 139 | 8.1284 | 8.0069 | 8.0020 | 0.89 | 0.3 | 2.3 | 0.12 | 9.09 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 183 | 7.7850 | 7.5978 | 7.6179 | 1.34 | 0.46 | 2.3 | 0.12 | 9.09 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 125 | 7.1436 | 7.4079 | 7.2412 | 1.84 | 0.62 | 2.3 | 0.12 | 9.09 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 196 | 7.1812 | 6.9303 | 7.1629 | 1.97 | 0.66 | 2.3 | 0.12 | 9.09 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 173 | 6.1444 | 6.0660 | 5.8776 | 2.27 | 0.75 | 2.3 | 0.12 | 9.09 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 164 | 8.8874 | 9.1687 | 9.2751 | 2.2 | 0.77 | 2.3 | 0.12 | 9.09 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 197 | 7.2574 | 7.5788 | 7.5432 | 2.36 | 0.8 | 2.3 | 0.12 | 9.09 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 144 | 7.6394 | 7.8577 | 8.0601 | 2.68 | 0.91 | 2.3 | 0.12 | 9.09 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 192 | 9.7069 | 9.3233 | 9.9071 | 3.08 | 1.08 | 2.3 | 0.12 | 9.09 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 149 | 5.6100 | 5.8380 | 6.0100 | 3.45 | 1.13 | 2.3 | 0.12 | 9.09 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 116 | 7.2276 | 7.5062 | 7.8069 | 3.86 | 1.31 | 2.3 | 0.12 | 9.09 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 192 | 7.2623 | 7.2928 | 6.7126 | 4.61 | 1.55 | 2.3 | 0.12 | 9.09 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 103 | 6.9294 | 6.5810 | 6.2766 | 4.95 | 1.65 | 2.3 | 0.12 | 9.09 | 0 |
| Total CBD (%AR) | 007.50 | GC-MS | 170 | 8.3629 | 9.2213 | 8.8792 | 4.9 | 1.7 | 2.3 | 0.12 | 9.09 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 141 | 7.1780 | 7.0433 | 8.1690 | 8.24 | 2.79 | 2.3 | 0.12 | 9.09 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 170 | 7.2625 | 7.5728 | 8.6190 | 9.09 | 3.1 | 2.3 | 0.12 | 9.09 | 0 |
|  |  |  |  | CBDV (%AR) |  |  |  |  |  |  |  |  |
| CBDV (%AR) | 008.99 | Other | 173 | <0.1590 <0.1500 |  | <0.1580 |  |  |  |  |  | 1 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDV (%AR) | 008.99 | Other | 197 | <0.0503 | <0.0526 | <0.0508 |  |  |  |  |  | 1 |
| CBDV (%AR) | 008.02 | AOAC 2018.11, U | 183 | 0.0301 | 0.0288 | 0.0287 | 2.67 | 0.39 | 7.41 | 2.67 | 12.4 | 0 |
| CBDV (%AR) | 008.99 | Other | 183 | 0.0306 | 0.0309 | 0.0291 | 3.19 | 0.47 | 7.41 | 2.67 | 12.4 | 0 |
| CBDV (%AR) | 008.99 | Other | 196 | 0.0209 | 0.0218 | 0.0230 | 4.81 | 0.68 | 7.41 | 2.67 | 12.4 | 0 |
| CBDV (%AR) | 008.99 | Other | 149 | 0.0250 | 0.0260 | 0.0300 | 9.8 | 1.43 | 7.41 | 2.67 | 12.4 | 0 |
| CBDV (%AR) | 008.99 | Other | 189 | 0.0312 | 0.0368 | 0.0296 | 11.6 | 1.74 | 7.41 | 2.67 | 12.4 | 0 |
| CBDV (%AR) | 008.99 | Other | 106 | 0.0400 | 0.0500 | 0.0500 | 12.4 | 1.96 | 7.41 | 2.67 | 12.4 | 0 |

Lab

Population of Lab RSDr

CBDVA (%AR)

| CBDVA (%AR) | 009.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDVA (%AR) | 009.99 | Other | 197 | <0.0503 | <0.0526 | <0.0508 |  |  | 1 |
| CBDVA (%AR) | 009.99 | Other | 173 | <0.1590 | <0.1500 | <0.1580 |  |  | 1 |
| CBDVA (%AR) | 009.02 | AOAC 2018.11, U | 183 | 0.0330 | 0.0320 | 0.0330 | 1.77 | 0.27 | 0 |
| CBDVA (%AR) | 009.99 | Other | 149 | 0.0340 | 0.0340 | 0.0300 | 7.07 | 1.06 | 0 |

CBG (%AR)

| CBG (%AR) | 010.99 | Other | 173 | <0.1590 | <0.1500 | <0.1580 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%AR) | 010.99 | Other | 106 | 0.1000 | 0.1000 | 0.1000 | 0 | 0 | 3.65 | 0 | 22.8 | 0 |
| CBG (%AR) | 010.99 | Other | 189 | 0.1394 | 0.1382 | 0.1399 | 0.63 | 0.12 | 3.65 | 0 | 22.8 | 0 |
| CBG (%AR) | 010.02 | AOAC 2018.11, U | 183 | 0.1416 | 0.1365 | 0.1375 | 1.95 | 0.36 | 3.65 | 0 | 22.8 | 0 |
| CBG (%AR) | 010.99 | Other | 149 | 0.0950 | 0.1000 | 0.0980 | 2.58 | 0.46 | 3.65 | 0 | 22.8 | 0 |
| CBG (%AR) | 010.99 | Other | 197 | 0.1075 | 0.1116 | 0.1137 | 2.84 | 0.51 | 3.65 | 0 | 22.8 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%AR) | 010.99 | Other | 155 | 0.2579 | 0.2613 | 0.2707 | 2.52 | 0.52 | 3.65 | 0 | 22.8 | 0 |
| CBG (%AR) | 010.99 | Other | 196 | 0.1088 | 0.1016 | 0.1110 | 4.59 | 0.82 | 3.65 | 0 | 22.8 | 0 |
| CBG (%AR) | 010.99 | Other | 183 | 0.1468 | 0.1338 | 0.1368 | 4.89 | 0.91 | 3.65 | 0 | 22.8 | 0 |
| CBG (%AR) | 010.02 | AOAC 2018.11, U | 125 | 0.1733 | 0.1489 | 0.1694 | 8 | 1.53 | 3.65 | 0 | 22.8 | 0 |
| CBG (%AR) | 010.99 | Other | 141 | 0.1010 | 0.1465 | 0.1600 | 22.8 | 4.23 | 3.65 | 0 | 22.8 | 0 |
|  |  |  |  | CBGA (%AR) |  |  |  |  |  |  |  |  |
| CBGA (%AR) | 011.99 | Other | 173 | <0.1590 | 0.1550 | <0.1580 |  |  |  |  |  | 1 |
| CBGA (%AR) | 011.99 | Other | 106 | 0.1200 | 0.1200 | 0.1200 | 0 | 0 | 4.19 | 0 | 14.1 | 0 |
| CBGA (%AR) | 011.99 | Other | 189 | 0.1548 | 0.1538 | 0.1538 | 0.38 | 0.07 | 4.19 | 0 | 14.1 | 0 |
| CBGA (%AR) | 011.02 | AOAC 2018.11, U | 125 | 0.1439 | 0.1485 | 0.1438 | 1.85 | 0.35 | 4.19 | 0 | 14.1 | 0 |
| CBGA (%AR) | 011.02 | AOAC 2018.11, U | 183 | 0.1799 | 0.1736 | 0.1715 | 2.5 | 0.48 | 4.19 | 0 | 14.1 | 0 |
| CBGA (%AR) | 011.99 | Other | 197 | 0.1487 | 0.1579 | 0.1574 | 3.34 | 0.63 | 4.19 | 0 | 14.1 | 0 |
| CBGA (%AR) | 011.99 | Other | 196 | 0.1574 | 0.1447 | 0.1443 | 5.01 | 0.94 | 4.19 | 0 | 14.1 | 0 |
| CBGA (%AR) | 011.99 | Other | 141 | 0.1890 | 0.1655 | 0.1990 | 9.32 | 1.81 | 4.19 | 0 | 14.1 | 0 |
| CBGA (%AR) | 011.99 | Other | 149 | 0.1100 | 0.1340 | 0.1030 | 14.1 | 2.55 | 4.19 | 0 | 14.1 | 0 |
|  |  |  |  | THCV (%AR) |  |  |  |  |  |  |  |  |
| THCV (%AR) | 012.99 | Other | 189 | <0.0050 | <0.0050 | <0.0050 |  |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 183 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 197 | <0.0503 | <0.0526 | <0.0508 |  |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 106 | <0.0300 | <0.0300 | <0.0300 |  |  |  |  |  | 1 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr HorRat(r) Rob Mean min max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| THCV (%AR) | 012.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 |  | 1 |
| THCV (%AR) | 012.99 | Other | 173 | 0.0000 | 0.0000 | 0.0000 | 0 | 0 |
| THCV (%AR) | 012.02 | AOAC 2018.11, U | 183 | <0.0020 | <0.0020 | <0.0020 |  | 1 |

Lab

Population of Lab RSDr

Δ8-THC (%AR)

| Δ8-THC (%AR) | 013.99 | Other | 183 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%AR) | 013.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 |  |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 173 | 0.0000 | 0.0000 | 0.0000 | 0 |  | 4.04 | 0 | 9.65 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 197 | <0.0503 | <0.0526 | <0.0508 |  |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 141 | 0.0000 | 0.0000 | 0.0000 | 0 |  | 4.04 | 0 | 9.65 | 0 |
| Δ8-THC (%AR) | 013.03 | AOAC 2018.11, M | 139 | <0.0125 | 0.0125 | 0.0125 |  |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 106 | <0.0300 | <0.0300 | <0.0300 |  |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.02 | AOAC 2018.11, U | 183 | <0.0010 | <0.0010 | <0.0010 |  |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.01 | AOAC 2018.10 | 228 | 4.4951 | 4.5042 | 4.4888 | 0.17 | 0.05 | 4.04 | 0 | 9.65 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 138 | 0.0771 | 0.0871 | 0.0795 | 6.43 | 1.11 | 4.04 | 0 | 9.65 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 189 | 0.0093 | 0.0105 | 0.0087 | 9.65 | 1.2 | 4.04 | 0 | 9.65 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 213 | 0.1350 | 0.1310 | 0.1520 | 8 | 1.49 | 4.04 | 0 | 9.65 | 0 |

CBC (%AR)

| CBC (%AR) | 014.99 | Other | 189 | 0.2317 | 0.2344 | 0.2337 | 0.6 | 0.12 | 1.94 | 0.6 | 10.6 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%AR) | 014.02 | AOAC 2018.11, U | 125 | 0.2087 | 0.2115 | 0.2084 | 0.82 | 0.16 | 1.94 | 0.6 | 10.6 | 0 |
| CBC (%AR) | 014.99 | Other | 183 | 0.2242 | 0.2241 | 0.2189 | 1.36 | 0.27 | 1.94 | 0.6 | 10.6 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%AR) | 014.99 | Other | 196 | 0.1995 | 0.1942 | 0.1961 | 1.37 | 0.27 | 1.94 | 0.6 | 10.6 | 0 |
| CBC (%AR) | 014.99 | Other | 197 | 0.1920 | 0.1979 | 0.1980 | 1.75 | 0.34 | 1.94 | 0.6 | 10.6 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, U | 183 | 0.2339 | 0.2274 | 0.2268 | 1.72 | 0.34 | 1.94 | 0.6 | 10.6 | 0 |
| CBC (%AR) | 014.99 | Other | 149 | 0.1730 | 0.1740 | 0.1680 | 1.87 | 0.36 | 1.94 | 0.6 | 10.6 | 0 |
| CBC (%AR) | 014.99 | Other | 106 | 0.1900 | 0.1900 | 0.2000 | 2.99 | 0.58 | 1.94 | 0.6 | 10.6 | 0 |
| CBC (%AR) | 014.99 | Other | 173 | 0.1810 | 0.1780 | 0.1700 | 3.22 | 0.62 | 1.94 | 0.6 | 10.6 | 0 |
| CBC (%AR) | 014.99 | Other | 155 | 0.1980 | 0.1970 | 0.2360 | 10.6 | 2.1 | 1.94 | 0.6 | 10.6 | 0 |
|  |  |  |  | Cd (ug/g AR) |  |  |  |  |  |  |  |  |
| Cd (ug/g AR) | 101.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 |  |  |  |  |  | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 102 | 0.1354 | 0.1388 | 0.1353 | 1.46 | 0.27 | 4.88 | 1.46 | 7.33 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 164 | 0.1249 | 0.1195 | 0.1194 | 2.6 | 0.47 | 4.88 | 1.46 | 7.33 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 197 | 0.1535 | 0.1644 | 0.1544 | 3.84 | 0.73 | 4.88 | 1.46 | 7.33 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 196 | 0.1310 | 0.1230 | 0.1170 | 5.68 | 1.04 | 4.88 | 1.46 | 7.33 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 106 | 0.1640 | 0.1440 | 0.1520 | 6.57 | 1.24 | 4.88 | 1.46 | 7.33 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 183 | 0.1600 | 0.1500 | 0.1400 | 6.67 | 1.26 | 4.88 | 1.46 | 7.33 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 141 | 0.1610 | 0.1500 | 0.1390 | 7.33 | 1.38 | 4.88 | 1.46 | 7.33 | 0 |
|  |  |  |  | As (ug/g AR) |  |  |  |  |  |  |  |  |
| As (ug/g AR) | 102.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 |  |  |  |  |  | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 102 | 0.1594 | 0.1564 | 0.1630 | 2.07 | 0.39 | 7.68 | 2.07 | 13.8 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 183 | 0.2000 | 0.2100 | 0.2100 | 2.79 | 0.55 | 7.68 | 2.07 | 13.8 | 0 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| As (ug/g AR) | 102.04 | ICP-MS | 106 | 0.1790 | 0.1620 | 0.1610 | 6.05 | 1.16 | 7.68 | 2.07 | 13.8 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 164 | 0.1593 | 0.1595 | 0.1422 | 6.46 | 1.22 | 7.68 | 2.07 | 13.8 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 196 | 0.1440 | 0.1630 | 0.1380 | 8.8 | 1.66 | 7.68 | 2.07 | 13.8 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 197 | 0.1709 | 0.1432 | 0.1309 | 13.8 | 2.6 | 7.68 | 2.07 | 13.8 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 141 | 0.1890 | 0.1540 | 0.1470 | 13.8 | 2.63 | 7.68 | 2.07 | 13.8 | 0 |

Lab

Population of Lab RSDr

Hg (ug/g AR)

| Hg (ug/g AR) | 103.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Hg (ug/g AR) | 103.04 | ICP-MS | 197 | <0.0978 | <0.0963 | <0.0988 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 141 | <0.0300 | <0.0300 | <0.0290 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 196 | <0.0072 | <0.0072 | <0.0072 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 164 | <0.0500 | <0.0500 | <0.0500 |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 183 | 0.0100 | 0.0100 | 0.0100 | 0 | 0 | 0 |
| Hg (ug/g AR) | 103.99 | Other | 102 | 0.0450 | 0.0450 | 0.0440 | 1.29 | 0.2 | 0 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 106 | 0.0090 | 0.0080 | 0.0080 | 6.93 | 0.85 | 0 |

Pb (ug/g AR)

| Pb (ug/g AR) | 104.04 | ICP-MS | 197 | 0.7603 | 0.7433 | 0.7380 | 1.56 | 0.37 | 12.46 | 1.56 | 46.1 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Pb (ug/g AR) | 104.04 | ICP-MS | 196 | 0.7630 | 0.7480 | 0.7330 | 2.01 | 0.48 | 12.46 | 1.56 | 46.1 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 102 | 0.7071 | 0.8060 | 0.7070 | 7.72 | 1.85 | 12.46 | 1.56 | 46.1 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 162 | 0.9249 | 1.0050 | 0.8265 | 9.73 | 2.41 | 12.46 | 1.56 | 46.1 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 183 | 0.8700 | 0.8100 | 0.6900 | 11.6 | 2.81 | 12.46 | 1.56 | 46.1 | 0 |

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Pb (ug/g AR) | 104.04 | ICP-MS | 141 | 0.8590 | 1.0600 | 0.7630 | 17 | 4.18 | 12.46 | 1.56 | 46.1 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 106 | 0.8170 | 0.7800 | 1.1390 | 21.7 | 5.35 | 12.46 | 1.56 | 46.1 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 164 | 0.6784 | 1.5626 | 0.8338 | 46.1 | 11.58 | 12.46 | 1.56 | 46.1 | 0 |

Lab

Population of Lab RSDr

Cu (ug/g AR)

Cu (ug/g AR)

105.04

ICP-MS

196

11.4000	11.5000	10.4000	5.48	1.97	0

Ni (ug/g AR)

| Ni (ug/g AR) | 106.04 | ICP-MS | 183 | 2.2800 | 2.3200 | 2.3500 | 1.52 | 0.43 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Ni (ug/g AR) | 106.04 | ICP-MS | 196 | 1.9300 | 1.9400 | 2.0000 | 1.93 | 0.54 | 0 |
| Ni (ug/g AR) | 106.04 | ICP-MS | 141 | 2.5900 | 2.5100 | 2.3800 | 4.25 | 1.22 | 0 |

Cr (ug/g AR)

| Cr (ug/g AR) | 107.04 | ICP-MS | 196 | 2.5700 | 2.5100 | 2.7100 | 3.95 | 1.14 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cr (ug/g AR) | 107.04 | ICP-MS | 141 | 3.4700 | 3.5600 | 3.1500 | 6.35 | 1.91 | 0 |

Se (ug/g AR)

| Se (ug/g AR) | 108.04 | ICP-MS | 141 | <0.3990 | <0.4030 | <0.3920 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Se (ug/g AR) | 108.04 | ICP-MS | 197 | 0.1189 | 0.1564 | <0.0988 |  |  | 1 |
| Se (ug/g AR) | 108.04 | ICP-MS | 196 | 0.1340 | 0.1260 | 0.1210 | 5.16 | 0.95 | 0 |

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

RSDr	HorRat(r)

Population of Lab RSDr	 Rob Mean	min	max

Flag

Zn (ug/g AR)

| Zn (ug/g AR) | 111.04 | ICP-MS | 141 | 76.9000 70.2000 | 68.7000 | 6.07 | 2.89 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Zn (ug/g AR) | 111.04 | ICP-MS | 196 | 72.5000 68.2000 | 63.9000 | 6.3 | 2.98 | 0 |

Sb (ug/g AR)

Sb (ug/g AR)

112.04

ICP-MS

196

<0.0241	<0.0241	<0.0241	1

(-)-alpha-Bisabolol (%AR)

| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 196 | 0.0653 | 0.0675 | 0.0664 | 1.66 | 0.28 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 162 | 0.0304 | 0.0303 | 0.0338 | 6.33 | 0.94 | 0 |

beta-Caryophyllene (%AR)

| beta-Caryophyllene (% | 202.02 | GC,MS | 196 | 0.0361 | 0.0379 | 0.0341 | 5.28 | 0.8 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| beta-Caryophyllene (% | 202.02 | GC,MS | 162 | 0.0474 | 0.0436 | 0.0432 | 5.18 | 0.81 | 0 |

(-)-Caryophyllene oxide (%AR)

(-)-Caryophyllene oxide

203.02

GC,MS

196

<0.0158	<0.0158	<0.0158	1

Geraniol (%AR)

| Geraniol (%AR) | 204.02 | GC,MS | 162 | <0.0100 <0.0100 | <0.0100 | 1 |
| --- | --- | --- | --- | --- | --- | --- |
| Geraniol (%AR) | 204.02 | GC,MS | 196 | <0.2520 <0.2520 | <0.2520 | 1 |

|  |  |  | Lab |  |  |  | Lab |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) |  | Rob Mean | min | max | Flag |
|  |  |  |  | alpha-Humulene (%AR) |  |  |  |  |  |  |  |  |  |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 196 | 0.0143 | 0.0151 | 0.0145 | 2.85 | 0.38 |  | 0 |  |  |  |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 162 | 0.0189 | 0.0176 | 0.0174 | 4.53 | 0.62 |  | 0 |  |  |  |
|  |  |  |  | d-Limonene (%AR) |  |  |  |  |  |  |  |  |  |
| d-Limonene (%AR) | 206.02 | GC,MS | 196 | <0.0039 <0.0039 |  | <0.0039 |  | 1 |  |  |  |  |  |
| d-Limonene (%AR) | 206.02 | GC,MS | 162 | <0.0100 <0.0100 |  | <0.0100 |  | 1 |  |  |  |  |  |
|  |  |  |  | Linalool (%AR) |  |  |  |  |  |  |  |  |  |
| Linalool (%AR) | 207.02 | GC,MS | 162 | <0.0100 <0.0100 |  | <0.0100 |  | 1 |  |  |  |  |  |
| Linalool (%AR) | 207.02 | GC,MS | 196 | <0.0039 <0.0039 |  | <0.0039 |  | 1 |  |  |  |  |  |
|  |  |  |  | beta-Myrcene (%AR) |  |  |  |  |  |  |  |  |  |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 162 | <0.0100 <0.0100 |  | <0.0100 |  | 1 |  |  |  |  |  |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 196 | <0.0039 <0.0039 |  | <0.0039 |  | 1 |  |  |  |  |  |
|  |  |  |  | Ocimene (%AR) |  |  |  |  |  |  |  |  |  |
| Ocimene (%AR) | 210.02 | GC,MS | 162 | <0.0100 <0.0100 |  | <0.0100 |  | 1 |  |  |  |  |  |
| Ocimene (%AR) | 210.02 | GC,MS | 196 | <0.0039 <0.0039 |  | <0.0039 |  | 1 |  |  |  |  |  |

Population of Lab RSDr

|  |  |  | Lab |  |  |  | Lab |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) |  | Rob Mean | min | max | Flag |
|  |  |  |  | alpha-Pinene (%AR) |  |  |  |  |  |  |  |  |  |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 162 | <0.0100 <0.0100 |  | <0.0100 |  | 1 |  |  |  |  |  |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 196 | <0.0039 <0.0039 |  | <0.0039 |  | 1 |  |  |  |  |  |
|  |  |  |  | (-)-beta-Pinene (%AR) |  |  |  |  |  |  |  |  |  |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 196 | <0.0039 <0.0039 |  | <0.0039 |  | 1 |  |  |  |  |  |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 162 | <0.0100 <0.0100 |  | <0.0100 |  | 1 |  |  |  |  |  |
|  |  |  |  | alpha-Terpinene (%AR) |  |  |  |  |  |  |  |  |  |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 162 | <0.0100 <0.0100 |  | <0.0100 |  | 1 |  |  |  |  |  |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 196 | <0.0039 <0.0039 |  | <0.0039 |  | 1 |  |  |  |  |  |
|  |  |  |  | gamma-Terpinene (%AR) |  |  |  |  |  |  |  |  |  |
| gamma-Terpinene (%A | 214.02 | GC,MS | 196 | <0.0039 <0.0039 |  | <0.0039 |  | 1 |  |  |  |  |  |
| gamma-Terpinene (%A | 214.02 | GC,MS | 162 | <0.0100 <0.0100 |  | <0.0100 |  | 1 |  |  |  |  |  |
|  |  |  |  | alpha-Terpinolene (%AR) |  |  |  |  |  |  |  |  |  |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 162 | <0.0100 <0.0100 |  | <0.0100 |  | 1 |  |  |  |  |  |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 196 | <0.0039 <0.0039 |  | <0.0039 |  | 1 |  |  |  |  |  |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Moisture (%AR) |  |  |  |  |  |  |  |  |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 189 | 6.4990 | 6.4990 | 6.4990 | 0 | 0 | 2.44 | 0 | 8.27 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 116 | 6.6400 | 6.7700 | 6.6100 | 1.27 | 0.42 | 2.44 | 0 | 8.27 | 0 |
| Moisture (%AR) | 500.99 | Other | 103 | 6.9100 | 7.1000 | 7.1000 | 1.56 | 0.52 | 2.44 | 0 | 8.27 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 182 | 8.6700 | 8.8900 | 8.9600 | 1.71 | 0.59 | 2.44 | 0 | 8.27 | 0 |
| Moisture (%AR) | 500.99 | Other | 165 | 8.7000 | 8.4000 | 8.6000 | 1.78 | 0.62 | 2.44 | 0 | 8.27 | 0 |
| Moisture (%AR) | 500.72 | NIR | 138 | 7.5000 | 7.3500 | 7.2000 | 2.04 | 0.69 | 2.44 | 0 | 8.27 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 199 | 10.5314 | 10.2499 | 10.6504 | 1.96 | 0.7 | 2.44 | 0 | 8.27 | 0 |
| Moisture (%AR) | 500.99 | Other | 102 | 6.9000 | 7.2000 | 7.0000 | 2.17 | 0.73 | 2.44 | 0 | 8.27 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 106 | 7.5200 | 7.4100 | 7.7400 | 2.22 | 0.75 | 2.44 | 0 | 8.27 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 192 | 9.8250 | 10.1240 | 10.2645 | 2.23 | 0.79 | 2.44 | 0 | 8.27 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 196 | 8.3300 | 8.4600 | 8.7700 | 2.65 | 0.92 | 2.44 | 0 | 8.27 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 153 | 5.0100 | 5.3200 | 5.3600 | 3.66 | 1.18 | 2.44 | 0 | 8.27 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 162 | 11.6500 | 11.0000 | 11.8800 | 3.97 | 1.43 | 2.44 | 0 | 8.27 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 210 | 8.0800 | 8.8600 | 9.0300 | 5.85 | 2.03 | 2.44 | 0 | 8.27 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 183 | 8.6400 | 9.8200 | 8.4500 | 8.27 | 2.88 | 2.44 | 0 | 8.27 | 0 |
|  |  |  |  | Δ9-THC (%DW) |  |  |  |  |  |  |  |  |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, U | 199 | 0.2045 | 0.2056 | 0.2065 | 0.49 | 0.1 | 2.29 | 0.49 | 4.43 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 189 | 0.1972 | 0.1998 | 0.1972 | 0.76 | 0.15 | 2.29 | 0.49 | 4.43 | 0 |
| Δ9-THC (%DW) | 501.01 | AOAC 2018.10 | 210 | 0.1800 | 0.1800 | 0.1770 | 0.97 | 0.19 | 2.29 | 0.49 | 4.43 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 116 | 0.1698 | 0.1774 | 0.1772 | 2.48 | 0.48 | 2.29 | 0.49 | 4.43 | 0 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 165 | 0.1666 | 0.1756 | 0.1697 | 2.68 | 0.52 | 2.29 | 0.49 | 4.43 | 0 |
| Δ9-THC (%DW) | 501.30 | Other LC-MS | 102 | 0.1690 | 0.1670 | 0.1770 | 3.09 | 0.59 | 2.29 | 0.49 | 4.43 | 0 |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, U | 103 | 0.1520 | 0.1420 | 0.1479 | 3.41 | 0.64 | 2.29 | 0.49 | 4.43 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 162 | 0.1958 | 0.1879 | 0.1792 | 4.43 | 0.86 | 2.29 | 0.49 | 4.43 | 0 |

Lab

Population of Lab RSDr

Δ9-THCA (%DW)

| Δ9-THCA (%DW) | 502.30 | Other LC-MS | 102 | <0.0530 | <0.0520 | <0.0530 |  |  |  |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, U | 199 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  |  | 1 |
| Δ9-THCA (%DW) | 502.01 | AOAC 2018.10 | 228 | 0.0611 | 0.0594 | 0.0604 | 1.42 | 0.23 | 3.22 | 1.42 | 4.45 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 162 | 0.0429 | 0.0450 | 0.0432 | 2.6 | 0.41 | 3.22 | 1.42 | 4.45 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 189 | 0.0475 | 0.0500 | 0.0496 | 2.74 | 0.44 | 3.22 | 1.42 | 4.45 | 0 |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, U | 103 | 0.0365 | 0.0377 | 0.0390 | 3.31 | 0.51 | 3.22 | 1.42 | 4.45 | 0 |
| Δ9-THCA (%DW) | 502.01 | AOAC 2018.10 | 210 | 0.0490 | 0.0460 | 0.0490 | 3.61 | 0.57 | 3.22 | 1.42 | 4.45 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 116 | 0.0514 | 0.0498 | 0.0471 | 4.4 | 0.7 | 3.22 | 1.42 | 4.45 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 165 | 0.0442 | 0.0483 | 0.0467 | 4.45 | 0.7 | 3.22 | 1.42 | 4.45 | 0 |

CBD (%DW)

| CBD (%DW) | 503.10 | Other LC-UV | 189 | 4.6741 | 4.6630 | 4.6700 | 0.12 | 0.04 | 1.6 | 0.12 | 6.43 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%DW) | 503.02 | AOAC 2018.11, U | 199 | 5.0716 | 5.1058 | 5.1310 | 0.58 | 0.19 | 1.6 | 0.12 | 6.43 | 0 |
| CBD (%DW) | 503.01 | AOAC 2018.10 | 210 | 4.7980 | 4.7440 | 4.7320 | 0.74 | 0.23 | 1.6 | 0.12 | 6.43 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 116 | 4.6837 | 4.7046 | 4.7693 | 0.95 | 0.3 | 1.6 | 0.12 | 6.43 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 162 | 4.6677 | 4.7383 | 4.9251 | 2.78 | 0.88 | 1.6 | 0.12 | 6.43 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%DW) | 503.02 | AOAC 2018.11, U | 103 | 4.5554 | 4.2431 | 4.0082 | 6.43 | 2 | 1.6 | 0.12 | 6.43 | 0 |
|  |  |  |  | CBDA (%DW) |  |  |  |  |  |  |  |  |
| CBDA (%DW) | 504.10 | Other LC-UV | 189 | 3.4224 | 3.4152 | 3.4210 | 0.11 | 0.03 | 1.28 | 0.11 | 3.07 | 0 |
| CBDA (%DW) | 504.01 | AOAC 2018.10 | 228 | 5.0015 | 5.0121 | 5.0009 | 0.13 | 0.04 | 1.28 | 0.11 | 3.07 | 0 |
| CBDA (%DW) | 504.02 | AOAC 2018.11, U | 199 | 3.6952 | 3.6830 | 3.7275 | 0.62 | 0.19 | 1.28 | 0.11 | 3.07 | 0 |
| CBDA (%DW) | 504.01 | AOAC 2018.10 | 210 | 4.1210 | 4.0370 | 4.0890 | 1.04 | 0.32 | 1.28 | 0.11 | 3.07 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 116 | 3.4988 | 3.4487 | 3.3986 | 1.45 | 0.44 | 1.28 | 0.11 | 3.07 | 0 |
| CBDA (%DW) | 504.02 | AOAC 2018.11, U | 103 | 3.2935 | 3.2394 | 3.1336 | 2.52 | 0.75 | 1.28 | 0.11 | 3.07 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 162 | 3.5313 | 3.5963 | 3.7483 | 3.07 | 0.93 | 1.28 | 0.11 | 3.07 | 0 |
|  |  |  |  | CBN (%DW) |  |  |  |  |  |  |  |  |
| CBN (%DW) | 505.10 | Other LC-UV | 189 | 0.0481 | 0.0479 | 0.0477 | 0.42 | 0.07 | 4.75 | 0.42 | 86.9 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 162 | 0.0473 | 0.0482 | 0.0473 | 1.09 | 0.17 | 4.75 | 0.42 | 86.9 | 0 |
| CBN (%DW) | 505.02 | AOAC 2018.11, U | 199 | 0.0514 | 0.0507 | 0.0500 | 1.38 | 0.22 | 4.75 | 0.42 | 86.9 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 116 | 0.0571 | 0.0612 | 0.0614 | 4.05 | 0.67 | 4.75 | 0.42 | 86.9 | 0 |
| CBN (%DW) | 505.30 | GC-FID | 192 | 0.0787 | 0.0877 | 0.0883 | 6.33 | 1.1 | 4.75 | 0.42 | 86.9 | 0 |
| CBN (%DW) | 505.30 | GC-FID | 116 | 0.0484 | 0.0565 | 0.0510 | 7.96 | 1.28 | 4.75 | 0.42 | 86.9 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 192 | 0.0000 | 0.0540 | 0.0591 | 86.9 | 13.31 | 4.75 | 0.42 | 86.9 | 0 |
|  |  |  |  | Total Δ9-THC (%DW) |  |  |  |  |  |  |  |  |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, U | 199 | 0.2264 | 0.2275 | 0.2284 | 0.44 | 0.09 | 3.25 | 0.44 | 16.2 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%DW) | 506.01 | AOAC 2018.10 | 210 | 0.2230 | 0.2200 | 0.2200 | 0.78 | 0.16 | 3.25 | 0.44 | 16.2 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 189 | 0.2389 | 0.2437 | 0.2407 | 1.01 | 0.2 | 3.25 | 0.44 | 16.2 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 192 | 0.2070 | 0.2088 | 0.2044 | 1.07 | 0.21 | 3.25 | 0.44 | 16.2 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 116 | 0.2149 | 0.2211 | 0.2185 | 1.43 | 0.28 | 3.25 | 0.44 | 16.2 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 138 | 0.2103 | 0.2029 | 0.2108 | 2.13 | 0.42 | 3.25 | 0.44 | 16.2 | 0 |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, U | 103 | 0.1840 | 0.1750 | 0.1821 | 2.63 | 0.51 | 3.25 | 0.44 | 16.2 | 0 |
| Total Δ9-THC (%DW) | 506.30 | Other LC-MS | 102 | 0.1880 | 0.1850 | 0.1960 | 3 | 0.59 | 3.25 | 0.44 | 16.2 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 165 | 0.2108 | 0.2239 | 0.2165 | 3.03 | 0.6 | 3.25 | 0.44 | 16.2 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 116 | 0.2061 | 0.2196 | 0.2179 | 3.43 | 0.68 | 3.25 | 0.44 | 16.2 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 162 | 0.2332 | 0.2270 | 0.2167 | 3.69 | 0.74 | 3.25 | 0.44 | 16.2 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 182 | 0.1801 | 0.1649 | 0.1533 | 8.09 | 1.55 | 3.25 | 0.44 | 16.2 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 153 | 0.1750 | 0.2360 | 0.2180 | 14.9 | 2.96 | 3.25 | 0.44 | 16.2 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 192 | 0.4099 | 0.4076 | 0.5355 | 16.2 | 3.61 | 3.25 | 0.44 | 16.2 | 0 |
|  |  |  |  | Total CBD (%DW) |  |  |  |  |  |  |  |  |
| Total CBD (%DW) | 507.10 | Other LC-UV | 189 | 7.6756 | 7.6582 | 7.6702 | 0.12 | 0.04 | 2.15 | 0.12 | 4.85 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 116 | 7.7521 | 7.7291 | 7.7498 | 0.16 | 0.06 | 2.15 | 0.12 | 4.85 | 0 |
| Total CBD (%DW) | 507.02 | AOAC 2018.11, U | 199 | 8.3123 | 8.3358 | 8.4000 | 0.54 | 0.19 | 2.15 | 0.12 | 4.85 | 0 |
| Total CBD (%DW) | 507.40 | GC-FID | 138 | 9.3488 | 9.3275 | 9.4564 | 0.74 | 0.26 | 2.15 | 0.12 | 4.85 | 0 |
| Total CBD (%DW) | 507.01 | AOAC 2018.10 | 210 | 8.4120 | 8.2840 | 8.3180 | 0.8 | 0.27 | 2.15 | 0.12 | 4.85 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 162 | 7.7646 | 7.8921 | 8.2127 | 2.9 | 0.99 | 2.15 | 0.12 | 4.85 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 192 | 10.7645 | 10.3736 | 11.0404 | 3.12 | 1.12 | 2.15 | 0.12 | 4.85 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%DW) | 507.40 | GC-FID | 116 | 7.7441 | 8.0427 | 8.3648 | 3.86 | 1.32 | 2.15 | 0.12 | 4.85 | 0 |
| Total CBD (%DW) | 507.40 | GC-FID | 192 | 8.0536 | 8.1143 | 7.4805 | 4.44 | 1.52 | 2.15 | 0.12 | 4.85 | 0 |
| Total CBD (%DW) | 507.02 | AOAC 2018.11, U | 103 | 7.4438 | 7.0840 | 6.7563 | 4.85 | 1.63 | 2.15 | 0.12 | 4.85 | 0 |
|  |  |  |  | CBDV (%DW) |  |  |  |  |  |  |  |  |
| CBDV (%DW) | 508.02 | AOAC 2018.11, U | 210 | 0.0350 | 0.0350 | 0.0340 | 1.67 | 0.25 |  |  |  | 0 |
| CBDV (%DW) | 508.99 | Other | 189 | 0.0334 | 0.0393 | 0.0316 | 11.6 | 1.75 |  |  |  | 0 |
| CBDV (%DW) | 508.99 | Other | 162 | 0.0375 | 0.0311 | 0.0302 | 12.1 | 1.82 |  |  |  | 0 |
|  |  |  |  | CBDVA (%DW) |  |  |  |  |  |  |  |  |
| CBDVA (%DW) | 509.01 | AOAC 2018.10 | 210 | 0.0280 | 0.0250 | 0.0240 | 8.11 | 1.17 |  |  |  | 0 |
|  |  |  |  | CBG (%DW) |  |  |  |  |  |  |  |  |
| CBG (%DW) | 510.99 | Other | 189 | 0.1491 | 0.1478 | 0.1497 | 0.65 | 0.12 |  |  |  | 0 |
| CBG (%DW) | 510.99 | Other | 162 | 0.1251 | 0.1221 | 0.1208 | 1.8 | 0.33 |  |  |  | 0 |
| CBG (%DW) | 510.01 | AOAC 2018.10 | 210 | 0.1360 | 0.1370 | 0.1410 | 1.92 | 0.36 |  |  |  | 0 |
|  |  |  |  | CBGA (%DW) |  |  |  |  |  |  |  |  |
| CBGA (%DW) | 511.99 | Other | 189 | 0.1656 | 0.1645 | 0.1645 | 0.38 | 0.07 |  |  |  | 0 |
| CBGA (%DW) | 511.99 | Other | 162 | 0.1900 | 0.1771 | 0.1807 | 3.65 | 0.71 |  |  |  | 0 |
| CBGA (%DW) | 511.01 | AOAC 2018.10 | 210 | 0.0900 | 0.0740 | 0.0690 | 14.1 | 2.41 |  |  |  | 0 |

Population of Lab RSDr

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

RSDr	HorRat(r)

Population of Lab RSDr	 Rob Mean	min	max

Flag

THCV (%DW)

| THCV (%DW) | 512.02 | AOAC 2018.11, U | 210 | <0.0130 <0.0130 | <0.0130 | 1 |
| --- | --- | --- | --- | --- | --- | --- |
| THCV (%DW) | 512.99 | Other | 189 | <0.0050 <0.0050 | <0.0050 | 1 |

Δ8-THC (%DW)

| Δ8-THC (%DW) | 513.99 | Other | 162 | <0.0100 | <0.0100 | <0.0100 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%DW) | 513.01 | AOAC 2018.10 | 210 | 0.0240 | 0.0240 | 0.0220 | 4.95 | 0.71 | 0 |
| Δ8-THC (%DW) | 513.99 | Other | 189 | 0.0099 | 0.0112 | 0.0093 | 9.58 | 1.21 | 0 |

CBC (%DW)

| CBC (%DW) | 514.99 | Other | 189 | 0.2478 | 0.2507 | 0.2500 | 0.61 | 0.12 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%DW) | 514.99 | Other | 162 | 0.2231 | 0.2259 | 0.2286 | 1.22 | 0.24 | 0 |

- Flag definitions: 1 (Lab RSDr not calculated because there were less than 3 numeric lab results)

- Statistical parameters of robust mean, minimum, and maximum for RSDr only shown only if there were 6 or more lab RSDr values.

- AR refers to concentration on as-received basis. DW refers to concentration on dry weight basis.c