# Homogeneity and Stability

Sample ID . . . . . . . . . . . . . . . . . . . . . . . .

Samples shipped . . . . . . . . . . . . . . . . . . .

Results due date . . . . . . . . . . . . . . . . . . . .

HM24SEP-1

9-Sep-24

18-Oct-24

## Homogeneity

8 sample jars randomly chosen for analysis.

| Sample # | Replicate 1 | Replicate 2 |
| --- | --- | --- |
| 1 | 0.1909 | 0.1908 |
| 2 | 0.2052 | 0.2011 |
| 3 | 0.1991 | 0.1959 |
| 4 | 0.1918 | 0.1976 |
| 5 | 0.1932 | 0.2025 |
| 6 | 0.1939 | 0.1906 |
| 7 | 0.1906 | 0.1942 |
| 8 | 0.1904 | 0.1923 |

Duplicate test portions analyzed for analyte and method shown in table caption.

| Total THC (%AR) via GC-FID (method 006.40) |  |  | %RSD | HorRat-r |
| --- | --- | --- | --- | --- |
| Analysis Date: 9/25/23 | Overall average: | 0.1950 |  |  |
|  | SD of sample avgs: | 0.004238 | 2.17 |  |
|  | repeatability SD: | 0.003304 | 1.69 | 0.3 |
|  | between-sample SD: | 0.00354 | 1.82 |  |
|  | reproducibilty SD: | 0.00484 | 2.48 |  |
|  | SD used for Analyte z scores: | 0.0309 |  |  |
|  | Check value: | 0.00927 |  |  |
|  | Total THC (%AR) via GC-FID (method 006.40) Is between-sample SD less than check value? YES Homogeneity test passed |  |  |  |
| Total CBD (%AR) via GC-FID (method 007.40) |  |  | %RSD | HorRat-r |
| Analysis Date: 9/25/23 | Overall average: | 7.8749 |  |  |
|  | SD of sample avgs: | 0.1337 | 1.70 |  |
|  | repeatability SD: | 0.1376 | 1.75 | 0.6 |
|  | between-sample SD: | 0.09177 | 1.17 |  |
|  | reproducibilty SD: | 0.1654 | 2.10 |  |
|  | SD used for Analyte z scores: | 0.9137 |  |  |
|  | Check value: | 0.27411 |  |  |
|  | Total CBD (%AR) via GC-FID (method 007.40) Is between-sample SD less than check value? YES Homogeneity test passed |  |  |  |

| Sample # | Replicate 1 | Replicate 2 |
| --- | --- | --- |
| 1 | 7.5457 | 7.7217 |
| 2 | 8.0394 | 8.0554 |
| 3 | 7.8393 | 7.8980 |
| 4 | 7.6803 | 7.9986 |
| 5 | 7.8521 | 8.2496 |
| 6 | 7.8518 | 7.7866 |
| 7 | 7.9296 | 7.8712 |
| 8 | 7.8221 | 7.8567 |

Stability

Sample tested on different days from date sample shipped to results due date. Sample stored at room temperature during period of testing.

| Total THC (%AR) via GC-FID (method 006.40) |  |  |
| --- | --- | --- |
| Date of Analysis | Days | Conc. |
| 9/10/24 | 0 | 0.1942 |
| 9/16/24 | 6 | 0.1941 |
| 9/17/24 | 7 | 0.2336 |
| 9/25/24 | 15 | 0.1947 |
| 9/30/24 | 20 | 0.1971 |
| 10/7/24 | 27 | 0.1846 |
| 10/10/24 | 46 | 0.2254 |

Days = x variable, Conc = y variable SUMMARY OUTPUT

Regression Statistics

Multiple R	 0.2347614 R Square	0.0551129 Adjusted R Squ -0.1338645 Standard Error  0.0196152

Observations	7

ANOVA

df	SS	MS	F	Significance F

Regression	1  0.00011221  0.0001122  0.2916	0.61234777

y = 0.0003x + 0.1986 R² = 0.0551

0.26

0.24

0.22

0.20

0.18

0.16

0.14

0.12

0.10

0

10

20

30

40

50

Residual	5  0.00192378  0.0003848

Total	6 0.00203599

Coefficients tandard Erro	t Stat	P-value	Lower 95%	Upper 95%

Intercept	0.1985992	0.0115552  17.186994	1E-05  0.168895639  0.228302838

Coefficient for x variable is not statistically significantly different from 0. No signficant change in result during testing period.

Stability test passed.

X Variable 1	0.0002769 0.00051275  0.5400347  0.6123  -0.001041166  0.001594973

%RSD = 9.1	HorRat-R = 1.8

| Total CBD (%AR) via GC-FID (method 006.40) |  |  |
| --- | --- | --- |
| Date of Analysis | Days | Conc. |
| 9/10/24 | 0 | 8.6617 |
| 9/16/24 | 6 | 8.6419 |
| 9/17/24 | 7 | 8.7614 |
| 9/25/24 | 15 | 8.7477 |
| 9/30/24 | 20 | 8.4618 |
| 10/1/24 | 21 | 8.0677 |
| 10/7/24 | 27 | 8.6254 |
| 10/10/24 | 46 | 8.2989 |

Days = x variable, Conc = y variable SUMMARY OUTPUT

Regression Statistics

Multiple R	0.5697501 R Square	0.3246151 Adjusted R Squ 0.212051 Standard Error  0.215247

Observations	8

ANOVA

df	SS	MS	F	Significance F

| Regression | 1 | 0.13361123 | 0.1336112 2.8838 0.140391466 |  |
| --- | --- | --- | --- | --- |
| Residual | 6 | 0.27798764 | 0.0463313 |  |
| Total | 7 | 0.41159887 |  |  |

y = -0.0057x + 8.6655 R² = 0.1367

8.9

8.7

8.5

8.3

8.1

7.9

7.7

7.5

0

10

20

30

40

50

Coefficients tandard Erro	t Stat	P-value	Lower 95%	Upper 95%

Intercept	8.7022205  0.12523775  69.485602	6E-10  8.395774764  9.008666239

X Variable 1	-0.0095159	0.0056036 -1.6981825  0.1404  -0.023227471  0.004195584

Coefficient for x variable is not statistically significantly different from 0. No signficant change in result during testing period.

Stability test passed.

%RSD = 2.8	HorRat-R = 1