Hemp Proficiency Testing

All Labs Trueness HM24SEP-2 (hemp)

Sample: HM24SEP-2 (hemp)

All Labs Trueness

Issue Date: November 11, 2024

| Analyte | Code | Method | Lab Num |  | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Δ9-THC (%AR) |  |  |  |  |  |  |  |  |  |  |
| Δ9-THC (%AR) | 001.30 | Other LC-MS |  | 122 | <0.5000 | <0.5000 | <0.5000 | <0.5000 |  |  |  |  | 1 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 213 | 0.0602 | 0.0601 | 0.0600 | 0.0601 | -3.49 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 182 | 0.1070 | 0.0988 | 0.1009 | 0.102 | -1.57 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 192 | 0.0000 | 0.1585 | 0.1578 | 0.105 | -1.43 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U |  | 161 | 0.1153 | 0.1035 | 0.1003 | 0.106 | -1.38 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U |  | 103 | 0.1145 | 0.1102 | 0.1108 | 0.112 | -1.13 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 106 | 0.1100 | 0.1200 | 0.1200 | 0.117 | -0.91 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 197 | 0.1150 | 0.1192 | 0.1237 | 0.119 | -0.79 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U |  | 125 | 0.1234 | 0.1257 | 0.1153 | 0.121 | -0.7 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 |  | 148 | 0.1260 | 0.1290 | 0.1250 | 0.127 | -0.46 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 149 | 0.1330 | 0.1280 | 0.1270 | 0.129 | -0.34 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS |  | 149 | 0.1400 | 0.1250 | 0.1270 | 0.131 | -0.28 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS |  | 102 | 0.1320 | 0.1320 | 0.1320 | 0.132 | -0.22 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 196 | 0.1289 | 0.1389 | 0.1340 | 0.134 | -0.13 | 0.1367 | 32 | 0.022 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 155 | 0.1410 | 0.1380 | 0.1250 | 0.135 | -0.09 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 165 | 0.1383 | 0.1338 | 0.1336 | 0.135 | -0.07 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 217 | 0.1387 | 0.1370 | 0.1367 | 0.137 | 0.03 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 175 | 0.1407 | 0.1364 | 0.1340 | 0.137 | 0.01 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 144 | 0.1382 | 0.1358 | 0.1397 | 0.138 | 0.05 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 115 | 0.1420 | 0.1318 | 0.1389 | 0.138 | 0.04 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 116 | 0.1378 | 0.1379 | 0.1409 | 0.139 | 0.1 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 170 | 0.1410 | 0.1433 | 0.1388 | 0.141 | 0.2 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 189 | 0.1454 | 0.1445 | 0.1446 | 0.145 | 0.37 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.99 | Other | 183 | 0.1479 | 0.1441 | 0.1471 | 0.146 | 0.44 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 101 | 0.1497 | 0.1474 | 0.1523 | 0.15 | 0.59 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 141 | 0.1365 | 0.1650 | 0.1570 | 0.153 | 0.73 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 163 | 0.1540 | 0.1550 | 0.1490 | 0.153 | 0.73 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 184 | 0.1536 | 0.1522 | 0.1578 | 0.155 | 0.81 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 173 | 0.1650 | <0.1540 | <0.1540 | 0.158 |  |  |  |  | 1 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 183 | 0.1611 | 0.1611 | 0.1573 | 0.16 | 1.05 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 199 | 0.1640 | 0.1655 | 0.1650 | 0.165 | 1.28 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.03 | AOAC 2018.11, M | 139 | 0.1776 | 0.1640 | 0.1944 | 0.179 | 1.91 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 164 | 0.1983 | 0.1904 | 0.1905 | 0.193 | 2.57 | 0.1367 | 32 | 0.022 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 228 | 0.2004 | 0.1998 | 0.1929 | 0.198 | 2.78 | 0.1367 | 32 | 0.022 | 0 |

Population of Lab Values

| Analyte | Code |  | Method | Lab Num |  | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Δ9-THCA (%AR) |  |  |  |  |  |  |  |  |  |  |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV |  | 163 | <0.0200 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  | 1 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV |  | 149 | <0.0250 | <0.0250 | <0.0250 | <0.0250 |  |  |  |  | 1 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV |  | 170 | <0.0300 | <0.0300 | <0.0300 | <0.0300 |  |  |  |  | 1 |
| Δ9-THCA (%AR) |  | 002.30 | Other LC-MS |  | 102 | <0.0480 | <0.0490 | <0.0480 | <0.0480 |  |  |  |  | 1 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV |  | 197 | <0.0483 | <0.0469 | <0.0483 | <0.0483 |  |  |  |  | 1 |
| Δ9-THCA (%AR) |  | 002.02 | AOAC 2018.11, U |  | 161 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| Δ9-THCA (%AR) |  | 002.02 | AOAC 2018.11, U |  | 199 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| Δ9-THCA (%AR) |  | 002.30 | Other LC-MS |  | 122 | <0.5000 | <0.5000 | <0.5000 | <0.5000 |  |  |  |  | 1 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV |  | 213 | 0.0061 | 0.0059 | 0.0063 | 0.0061 | -2.86 | 0.0325 | 23 | 0.0092 | 0 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV |  | 182 | 0.0204 | 0.0167 | 0.0142 | 0.0171 | -1.66 | 0.0325 | 23 | 0.0092 | 0 |
| Δ9-THCA (%AR) |  | 002.02 | AOAC 2018.11, U |  | 103 | 0.0207 | 0.0207 | 0.0217 | 0.021 | -1.24 | 0.0325 | 23 | 0.0092 | 0 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV |  | 196 | 0.0250 | 0.0252 | 0.0221 | 0.0241 | -0.91 | 0.0325 | 23 | 0.0092 | 0 |
| Δ9-THCA (%AR) |  | 002.03 | AOAC 2018.11, M |  | 139 | 0.0260 | 0.0224 | 0.0300 | 0.0261 | -0.69 | 0.0325 | 23 | 0.0092 | 0 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV |  | 165 | 0.0253 | 0.0291 | 0.0271 | 0.0272 | -0.57 | 0.0325 | 23 | 0.0092 | 0 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV |  | 184 | 0.0274 | 0.0287 | 0.0270 | 0.0277 | -0.52 | 0.0325 | 23 | 0.0092 | 0 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV |  | 217 | 0.0310 | 0.0257 | 0.0290 | 0.0286 | -0.42 | 0.0325 | 23 | 0.0092 | 0 |
| Δ9-THCA (%AR) |  | 002.01 | AOAC 2018.10 |  | 115 | 0.0296 | 0.0278 | 0.0289 | 0.0288 | -0.4 | 0.0325 | 23 | 0.0092 | 0 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV |  | 106 | <0.0300 | <0.0300 | 0.0300 | 0.03 |  |  |  |  | 1 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV |  | 189 | 0.0341 | 0.0193 | 0.0357 | 0.0297 | -0.3 | 0.0325 | 23 | 0.0092 | 0 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV |  | 116 | 0.0311 | 0.0320 | 0.0302 | 0.0311 | -0.15 | 0.0325 | 23 | 0.0092 | 0 |
| Δ9-THCA (%AR) |  | 002.10 | Other LC-UV |  | 144 | 0.0319 | 0.0331 | 0.0333 | 0.0328 | 0.03 | 0.0325 | 23 | 0.0092 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.99 | Other | 183 | 0.0345 | 0.0338 | 0.0353 | 0.0345 | 0.22 | 0.0325 | 23 | 0.0092 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 183 | 0.0350 | 0.0341 | 0.0353 | 0.0348 | 0.25 | 0.0325 | 23 | 0.0092 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 149 | 0.0440 | 0.0330 | 0.0290 | 0.0353 | 0.31 | 0.0325 | 23 | 0.0092 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 175 | 0.0425 | 0.0404 | 0.0211 | 0.0347 | 0.24 | 0.0325 | 23 | 0.0092 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 101 | 0.0363 | 0.0376 | 0.0335 | 0.0358 | 0.36 | 0.0325 | 23 | 0.0092 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 164 | 0.0379 | 0.0384 | 0.0382 | 0.0382 | 0.62 | 0.0325 | 23 | 0.0092 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 155 | 0.0435 | 0.0345 | 0.0394 | 0.0391 | 0.72 | 0.0325 | 23 | 0.0092 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 148 | 0.0470 | 0.0440 | 0.0390 | 0.0433 | 1.18 | 0.0325 | 23 | 0.0092 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 125 | 0.0379 | 0.0298 | 0.0641 | 0.0439 | 1.24 | 0.0325 | 23 | 0.0092 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 192 | 0.2273 | 0.0000 | 0.0000 | 0.0758 | 4.69 | 0.0325 | 23 | 0.0092 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 173 | <0.1550 | <0.1540 | 0.0000 | 0.103 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 141 | 0.1389 | 0.1250 | 0.1184 | 0.127 | 10.29 | 0.0325 | 23 | 0.0092 | 0 |
|  |  |  |  | CBD (%AR) |  |  |  |  |  |  |  |  |
| CBD (%AR) | 003.10 | Other LC-UV | 173 | 4.0600 | 3.7900 | 3.8730 | 3.91 | -1.56 | 4.6033 | 25 | 0.4465 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 149 | 3.8560 | 3.9580 | 3.9440 | 3.92 | -1.53 | 4.6033 | 25 | 0.4465 | 0 |
| CBD (%AR) | 003.01 | AOAC 2018.10 | 228 | 4.0078 | 4.0128 | 4.0009 | 4.01 | -1.33 | 4.6033 | 25 | 0.4465 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 155 | 4.0700 | 4.1200 | 3.9400 | 4.04 | -1.25 | 4.6033 | 25 | 0.4465 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 161 | 4.2798 | 4.0807 | 4.0285 | 4.13 | -1.06 | 4.6033 | 25 | 0.4465 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 149 | 4.2370 | 4.3870 | 4.1560 | 4.26 | -0.77 | 4.6033 | 25 | 0.4465 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 196 | 4.2711 | 4.5391 | 4.2866 | 4.37 | -0.53 | 4.6033 | 25 | 0.4465 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 184 | 4.4215 | 4.3321 | 4.4416 | 4.4 | -0.46 | 4.6033 | 25 | 0.4465 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 103 | 4.3318 | 4.5149 | 4.3911 | 4.41 | -0.43 | 4.6033 | 25 | 0.4465 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 189 | 4.5735 | 4.5673 | 4.5750 | 4.57 | -0.07 | 4.6033 | 25 | 0.4465 | 0 |
| CBD (%AR) | 003.03 | AOAC 2018.11, M | 139 | 4.5433 | 4.4266 | 4.7895 | 4.59 | -0.04 | 4.6033 | 25 | 0.4465 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 125 | 4.6101 | 4.8279 | 4.4691 | 4.64 | 0.07 | 4.6033 | 25 | 0.4465 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 170 | 4.6373 | 4.8409 | 4.5174 | 4.67 | 0.14 | 4.6033 | 25 | 0.4465 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 183 | 4.7799 | 4.7198 | 4.6984 | 4.73 | 0.29 | 4.6033 | 25 | 0.4465 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 141 | 4.5943 | 4.6190 | 5.1212 | 4.78 | 0.39 | 4.6033 | 25 | 0.4465 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 106 | 4.6800 | 4.8700 | 4.8000 | 4.78 | 0.4 | 4.6033 | 25 | 0.4465 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 197 | 4.7932 | 4.8244 | 4.7826 | 4.8 | 0.44 | 4.6033 | 25 | 0.4465 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 199 | 4.7080 | 4.8350 | 4.8995 | 4.81 | 0.47 | 4.6033 | 25 | 0.4465 | 0 |
| CBD (%AR) | 003.99 | Other | 183 | 4.8819 | 4.7907 | 4.8007 | 4.82 | 0.5 | 4.6033 | 25 | 0.4465 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 116 | 4.7869 | 4.7544 | 4.9038 | 4.82 | 0.47 | 4.6033 | 25 | 0.4465 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 163 | 4.9000 | 4.9000 | 4.9000 | 4.9 | 0.66 | 4.6033 | 25 | 0.4465 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 144 | 5.0030 | 4.8186 | 5.3966 | 5.07 | 1.05 | 4.6033 | 25 | 0.4465 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 164 | 5.0805 | 5.1536 | 5.0671 | 5.1 | 1.11 | 4.6033 | 25 | 0.4465 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 213 | 5.3400 | 5.2220 | 5.1720 | 5.24 | 1.44 | 4.6033 | 25 | 0.4465 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 192 | 6.6752 | 7.3058 | 6.1795 | 6.72 | 4.74 | 4.6033 | 25 | 0.4465 | 0 |
|  |  |  |  | CBDA (%AR) |  |  |  |  |  |  |  |  |
| CBDA (%AR) | 004.10 | Other LC-UV | 173 | 2.1440 | 1.9690 | 2.0500 | 2.05 | -1.89 | 2.7468 | 22 | 0.3669 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 155 | 2.3300 | 2.3100 | 2.1400 | 2.26 | -1.33 | 2.7468 | 22 | 0.3669 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 163 | 2.3300 | 2.3600 | 2.3500 | 2.35 | -1.09 | 2.7468 | 22 | 0.3669 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDA (%AR) | 004.30 | Other LC-MS | 149 | 2.2960 | 2.3760 | 2.5800 | 2.42 | -0.9 | 2.7468 | 22 | 0.3669 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 149 | 2.4250 | 2.4650 | 2.4680 | 2.45 | -0.8 | 2.7468 | 22 | 0.3669 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 161 | 2.5207 | 2.5010 | 2.5312 | 2.52 | -0.62 | 2.7468 | 22 | 0.3669 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 103 | 2.5252 | 2.5770 | 2.6022 | 2.57 | -0.49 | 2.7468 | 22 | 0.3669 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 196 | 2.5555 | 2.6863 | 2.5460 | 2.6 | -0.41 | 2.7468 | 22 | 0.3669 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 197 | 2.6783 | 2.7587 | 2.6512 | 2.7 | -0.14 | 2.7468 | 22 | 0.3669 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 213 | 2.7340 | 2.6900 | 2.7060 | 2.71 | -0.1 | 2.7468 | 22 | 0.3669 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 189 | 2.7201 | 2.7159 | 2.7209 | 2.72 | -0.08 | 2.7468 | 22 | 0.3669 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 125 | 2.7516 | 2.8227 | 2.6684 | 2.75 | 0 | 2.7468 | 22 | 0.3669 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 106 | 2.7100 | 2.7800 | 2.8100 | 2.77 | 0.05 | 2.7468 | 22 | 0.3669 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 116 | 2.8307 | 2.7744 | 2.7682 | 2.79 | 0.12 | 2.7468 | 22 | 0.3669 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 199 | 2.7975 | 2.8255 | 2.8050 | 2.81 | 0.17 | 2.7468 | 22 | 0.3669 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 141 | 2.7997 | 2.7410 | 2.9971 | 2.85 | 0.27 | 2.7468 | 22 | 0.3669 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 184 | 2.9388 | 2.9554 | 2.8584 | 2.92 | 0.47 | 2.7468 | 22 | 0.3669 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 144 | 3.0096 | 2.9633 | 3.1908 | 3.05 | 0.84 | 2.7468 | 22 | 0.3669 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 183 | 3.1511 | 3.1181 | 3.1065 | 3.13 | 1.03 | 2.7468 | 22 | 0.3669 | 0 |
| CBDA (%AR) | 004.03 | AOAC 2018.11, M | 139 | 3.5560 | 3.2987 | 3.6937 | 3.52 | 2.1 | 2.7468 | 22 | 0.3669 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 192 | 3.4919 | 4.0959 | 3.2800 | 3.62 | 2.39 | 2.7468 | 22 | 0.3669 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 164 | 4.8172 | 4.2628 | 4.2028 | 4.43 | 4.58 | 2.7468 | 22 | 0.3669 | 0 |
|  |  |  |  | CBN (%AR) |  |  |  |  |  |  |  |  |
| CBN (%AR) | 005.10 | Other LC-UV | 197 | <0.0725 <0.0704 <0.0725 |  |  | <0.0725 |  |  |  |  | 1 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%AR) | 005.10 | Other LC-UV | 173 | <0.1550 | <0.1540 | <0.1540 | <0.1550 |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 149 | 0.0490 | 0.0480 | 0.0500 | 0.049 | -1.14 | 0.062 | 19 | 0.0115 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 196 | 0.0498 | 0.0520 | 0.0494 | 0.0504 | -1.01 | 0.062 | 19 | 0.0115 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 213 | 0.0500 | 0.0510 | 0.0510 | 0.0507 | -0.99 | 0.062 | 19 | 0.0115 | 0 |
| CBN (%AR) | 005.20 | Other LC-MS | 149 | 0.0590 | 0.0530 | 0.0450 | 0.0523 | -0.85 | 0.062 | 19 | 0.0115 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 144 | 0.0533 | 0.0535 | 0.0543 | 0.0537 | -0.73 | 0.062 | 19 | 0.0115 | 0 |
| CBN (%AR) | 005.03 | AOAC 2018.11, M | 139 | 0.0575 | 0.0517 | 0.0521 | 0.0538 | -0.72 | 0.062 | 19 | 0.0115 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 141 | 0.0543 | 0.0580 | 0.0591 | 0.0571 | -0.43 | 0.062 | 19 | 0.0115 | 0 |
| CBN (%AR) | 005.99 | Other | 183 | 0.0614 | 0.0552 | 0.0573 | 0.058 | -0.36 | 0.062 | 19 | 0.0115 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 192 | 0.0000 | 0.0846 | 0.0910 | 0.0585 | -0.31 | 0.062 | 19 | 0.0115 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 189 | 0.0586 | 0.0597 | 0.0584 | 0.0589 | -0.27 | 0.062 | 19 | 0.0115 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 183 | 0.0596 | 0.0589 | 0.0596 | 0.0594 | -0.23 | 0.062 | 19 | 0.0115 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 199 | 0.0600 | 0.0625 | 0.0610 | 0.0612 | -0.08 | 0.062 | 19 | 0.0115 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 116 | 0.0659 | 0.0688 | 0.0658 | 0.0668 | 0.42 | 0.062 | 19 | 0.0115 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 106 | 0.0700 | 0.0700 | 0.0700 | 0.07 | 0.69 | 0.062 | 19 | 0.0115 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 125 | 0.0710 | 0.0724 | 0.0655 | 0.0696 | 0.66 | 0.062 | 19 | 0.0115 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 116 | 0.0716 | 0.0737 | 0.0743 | 0.0732 | 0.97 | 0.062 | 19 | 0.0115 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 138 | 0.0665 | 0.0887 | 0.0780 | 0.0777 | 1.37 | 0.062 | 19 | 0.0115 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 163 | 0.0830 | 0.0830 | 0.0810 | 0.0823 | 1.77 | 0.062 | 19 | 0.0115 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 192 | 0.0839 | 0.0961 | 0.0894 | 0.0898 | 2.42 | 0.062 | 19 | 0.0115 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Total Δ9-THC (%AR) |  |  |  |  |  |  |  |  |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 213 | 0.0628 | 0.0627 | 0.0625 | 0.0627 | -4.24 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 182 | 0.1249 | 0.1135 | 0.1134 | 0.117 | -2.01 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 197 | 0.1150 | 0.1192 | 0.1237 | 0.119 | -1.93 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 169 | 0.1304 | 0.1211 | 0.1270 | 0.126 | -1.65 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 103 | 0.1327 | 0.1284 | 0.1298 | 0.13 | -1.48 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 192 | 0.1381 | 0.1596 | 0.1455 | 0.148 | -0.77 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 149 | 0.1530 | 0.1480 | 0.1470 | 0.149 | -0.7 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 102 | 0.1490 | 0.1490 | 0.1490 | 0.149 | -0.72 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 196 | 0.1508 | 0.1610 | 0.1533 | 0.155 | -0.47 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 116 | 0.1591 | 0.1574 | 0.1520 | 0.156 | -0.42 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 114 | 0.1705 | 0.1428 |  | 0.157 | -0.4 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 153 | 0.1440 | 0.1735 | 0.1582 | 0.159 | -0.33 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 190 | 0.1628 | 0.1551 |  | 0.159 | -0.31 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 125 | 0.1566 | 0.1518 | 0.1715 | 0.16 | -0.27 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 138 | 0.1760 | 0.1500 | 0.1564 | 0.161 | -0.24 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 165 | 0.1636 | 0.1629 | 0.1607 | 0.162 | -0.17 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 149 | 0.1790 | 0.1540 | 0.1520 | 0.162 | -0.2 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 115 | 0.1679 | 0.1561 | 0.1642 | 0.163 | -0.16 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 217 | 0.1659 | 0.1595 | 0.1622 | 0.163 | -0.16 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 148 | 0.1670 | 0.1680 | 0.1590 | 0.165 | -0.08 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 116 | 0.1651 | 0.1660 | 0.1674 | 0.166 | -0.02 | 0.1666 | 38 | 0.0245 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 175 | 0.1780 | 0.1719 | 0.1522 | 0.167 | 0.03 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 144 | 0.1662 | 0.1648 | 0.1689 | 0.167 | 0 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 155 | 0.1790 | 0.1680 | 0.1590 | 0.169 | 0.09 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 189 | 0.1752 | 0.1614 | 0.1759 | 0.171 | 0.17 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 192 | 0.1993 | 0.1585 | 0.1578 | 0.172 | 0.22 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.99 | Other | 183 | 0.1781 | 0.1738 | 0.1780 | 0.177 | 0.41 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 173 | 0.2036 | 0.1843 | 0.1510 | 0.18 | 0.53 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 101 | 0.1816 | 0.1804 | 0.1817 | 0.181 | 0.6 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.50 | GC-MS | 170 | 0.1690 | 0.1754 | 0.1980 | 0.181 | 0.58 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 199 | 0.1859 | 0.1874 | 0.1869 | 0.187 | 0.82 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 183 | 0.1918 | 0.1910 | 0.1883 | 0.19 | 0.97 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 163 | 0.1900 | 0.1900 | 0.1960 | 0.192 | 1.04 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.03 | AOAC 2018.11, M | 139 | 0.2004 | 0.1836 | 0.2207 | 0.202 | 1.43 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 164 | 0.2315 | 0.2241 | 0.2241 | 0.227 | 2.45 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 170 | 0.2322 | 0.2201 | 0.2502 | 0.234 | 2.76 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 141 | 0.2583 | 0.2740 | 0.2608 | 0.264 | 3.99 | 0.1666 | 38 | 0.0245 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 107 | 0.4020 | 0.3720 | 0.4120 | 0.395 | 9.33 | 0.1666 | 38 | 0.0245 | 0 |
|  |  |  |  | Total CBD (%AR) |  |  |  |  |  |  |  |  |
| Total CBD (%AR) | 007.10 | Other LC-UV | 173 | 5.9403 | 5.5168 | 5.6709 | 5.71 | -1.63 | 7.2794 | 24 | 0.9634 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 149 | 5.9830 | 6.1200 | 6.1080 | 6.07 | -1.25 | 7.2794 | 24 | 0.9634 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 155 | 6.1100 | 6.2400 | 5.9600 | 6.1 | -1.22 | 7.2794 | 24 | 0.9634 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%AR) | 007.30 | Other LC-MS | 149 | 6.2510 | 6.4710 | 6.4190 | 6.38 | -0.93 | 7.2794 | 24 | 0.9634 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 169 | 6.3290 | 6.4620 | 6.3920 | 6.39 | -0.92 | 7.2794 | 24 | 0.9634 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 196 | 6.5123 | 6.8950 | 6.5195 | 6.64 | -0.66 | 7.2794 | 24 | 0.9634 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 103 | 6.5464 | 6.7749 | 6.6732 | 6.66 | -0.64 | 7.2794 | 24 | 0.9634 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 189 | 6.9590 | 6.9492 | 6.9612 | 6.96 | -0.34 | 7.2794 | 24 | 0.9634 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 125 | 7.0233 | 7.3033 | 6.8093 | 7.05 | -0.24 | 7.2794 | 24 | 0.9634 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 192 | 6.9566 | 7.5679 | 6.6380 | 7.05 | -0.23 | 7.2794 | 24 | 0.9634 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 197 | 7.1421 | 7.2438 | 7.1077 | 7.16 | -0.12 | 7.2794 | 24 | 0.9634 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 116 | 7.2695 | 7.1875 | 7.3314 | 7.26 | -0.02 | 7.2794 | 24 | 0.9634 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 141 | 7.0496 | 7.0230 | 7.7497 | 7.27 | -0.01 | 7.2794 | 24 | 0.9634 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 199 | 7.1614 | 7.3130 | 7.3595 | 7.28 | 0 | 7.2794 | 24 | 0.9634 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 183 | 7.5434 | 7.4544 | 7.4227 | 7.47 | 0.2 | 7.2794 | 24 | 0.9634 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 116 | 7.4572 | 7.3210 | 7.8472 | 7.54 | 0.27 | 7.2794 | 24 | 0.9634 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 170 | 7.3106 | 7.3881 | 8.1695 | 7.62 | 0.36 | 7.2794 | 24 | 0.9634 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 163 | 7.6000 | 7.5600 | 7.7100 | 7.62 | 0.36 | 7.2794 | 24 | 0.9634 | 0 |
| Total CBD (%AR) | 007.03 | AOAC 2018.11, M | 139 | 7.6619 | 7.3196 | 8.0289 | 7.67 | 0.41 | 7.2794 | 24 | 0.9634 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 144 | 7.6430 | 7.4180 | 8.1956 | 7.75 | 0.49 | 7.2794 | 24 | 0.9634 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 138 | 8.4776 | 8.8145 | 8.9560 | 8.75 | 1.53 | 7.2794 | 24 | 0.9634 | 0 |
| Total CBD (%AR) | 007.50 | GC-MS | 170 | 8.6887 | 8.5791 | 9.0421 | 8.77 | 1.55 | 7.2794 | 24 | 0.9634 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 164 | 9.3052 | 8.8921 | 8.7530 | 8.98 | 1.77 | 7.2794 | 24 | 0.9634 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 192 | 9.7376 | 10.8979 | 9.0560 | 9.9 | 2.72 | 7.2794 | 24 | 0.9634 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num |  | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | CBDV (%AR) |  |  |  |  |  |  |  |  |  |  |
| CBDV (%AR) | 008.99 | Other |  | 197 | <0.0483 | <0.0469 | <0.0483 | <0.0483 |  |  |  |  | 1 |
| CBDV (%AR) | 008.99 | Other |  | 173 | <0.1550 | <0.1540 | <0.1540 | <0.1550 |  |  |  |  | 1 |
| CBDV (%AR) | 008.99 | Other |  | 196 | 0.0254 | 0.0272 | 0.0248 | 0.0258 | -1.12 | 0.0354 | 6 | 0.0085 | 0 |
| CBDV (%AR) | 008.99 | Other |  | 189 | 0.0321 | 0.0316 | 0.0278 | 0.0305 | -0.57 | 0.0354 | 6 | 0.0085 | 0 |
| CBDV (%AR) | 008.99 | Other |  | 149 | 0.0390 | 0.0310 | 0.0350 | 0.035 | -0.05 | 0.0354 | 6 | 0.0085 | 0 |
| CBDV (%AR) | 008.02 | AOAC 2018.11, U |  | 183 | 0.0367 | 0.0356 | 0.0348 | 0.0357 | 0.04 | 0.0354 | 6 | 0.0085 | 0 |
| CBDV (%AR) | 008.99 | Other |  | 183 | 0.0377 | 0.0377 | 0.0361 | 0.0372 | 0.21 | 0.0354 | 6 | 0.0085 | 0 |
| CBDV (%AR) | 008.99 | Other |  | 106 | 0.0600 | 0.0600 | 0.0500 | 0.0567 | 2.49 | 0.0354 | 6 | 0.0085 | 0 |

Population of Lab Values

| CBDVA (%AR) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDVA (%AR) | 009.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 | <0.0205 | 1 |
| CBDVA (%AR) | 009.99 | Other | 197 | <0.0483 | <0.0469 | <0.0483 | <0.0483 | 1 |
| CBDVA (%AR) | 009.99 | Other | 173 | <0.1550 | <0.1540 | <0.1540 | <0.1550 | 1 |
| CBDVA (%AR) | 009.02 | AOAC 2018.11, U | 183 | 0.0330 | 0.0320 | 0.0330 | 0.0327 | 0 |
| CBDVA (%AR) | 009.99 | Other | 149 | 0.0370 | 0.0350 | 0.0340 | 0.0353 | 0 |

|  |  |  | CBG (%AR) |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%AR) | 010.99 | Other | 197 | 0.1314 | 0.1249 | 0.1256 | 0.127 | -1.34 | 0.1856 | 11 | 0.0437 | 0 |
| CBG (%AR) | 010.99 | Other | 196 | 0.1403 | 0.1495 | 0.1425 | 0.144 | -0.95 | 0.1856 | 11 | 0.0437 | 0 |
| CBG (%AR) | 010.99 | Other | 149 | 0.1480 | 0.1430 | 0.1470 | 0.146 | -0.91 | 0.1856 | 11 | 0.0437 | 0 |
| CBG (%AR) | 010.99 | Other | 106 | 0.1600 | 0.1700 | 0.1600 | 0.163 | -0.51 | 0.1856 | 11 | 0.0437 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%AR) | 010.99 | Other | 173 | 0.2090 | 0.1620 | 0.1780 | 0.183 | -0.06 | 0.1856 | 11 | 0.0437 | 0 |
| CBG (%AR) | 010.99 | Other | 183 | 0.2068 | 0.1863 | 0.1899 | 0.194 | 0.2 | 0.1856 | 11 | 0.0437 | 0 |
| CBG (%AR) | 010.99 | Other | 189 | 0.1944 | 0.1948 | 0.1958 | 0.195 | 0.21 | 0.1856 | 11 | 0.0437 | 0 |
| CBG (%AR) | 010.02 | AOAC 2018.11, U | 183 | 0.1977 | 0.1955 | 0.1958 | 0.196 | 0.25 | 0.1856 | 11 | 0.0437 | 0 |
| CBG (%AR) | 010.99 | Other | 141 | 0.2169 | 0.1710 | 0.2226 | 0.204 | 0.41 | 0.1856 | 11 | 0.0437 | 0 |
| CBG (%AR) | 010.02 | AOAC 2018.11, U | 125 | 0.2390 | 0.2319 | 0.2428 | 0.238 | 1.2 | 0.1856 | 11 | 0.0437 | 0 |
| CBG (%AR) | 010.99 | Other | 155 | 0.3000 | 0.2751 | 0.2752 | 0.283 | 2.24 | 0.1856 | 11 | 0.0437 | 0 |
|  |  |  |  | CBGA (%AR) |  |  |  |  |  |  |  |  |
| CBGA (%AR) | 011.99 | Other | 173 | <0.1550 | <0.1540 | <0.1540 | <0.1550 |  |  |  |  | 1 |
| CBGA (%AR) | 011.99 | Other | 149 | 0.1010 | 0.0970 | 0.1200 | 0.106 | -1.14 | 0.1339 | 8 | 0.0245 | 0 |
| CBGA (%AR) | 011.99 | Other | 106 | 0.1100 | 0.1100 | 0.1100 | 0.11 | -0.97 | 0.1339 | 8 | 0.0245 | 0 |
| CBGA (%AR) | 011.99 | Other | 196 | 0.1238 | 0.1151 | 0.1210 | 0.12 | -0.57 | 0.1339 | 8 | 0.0245 | 0 |
| CBGA (%AR) | 011.02 | AOAC 2018.11, U | 125 | 0.1348 | 0.1371 | 0.1271 | 0.133 | -0.03 | 0.1339 | 8 | 0.0245 | 0 |
| CBGA (%AR) | 011.99 | Other | 189 | 0.1381 | 0.1399 | 0.1390 | 0.139 | 0.21 | 0.1339 | 8 | 0.0245 | 0 |
| CBGA (%AR) | 011.99 | Other | 197 | 0.1420 | 0.1408 | 0.1411 | 0.141 | 0.3 | 0.1339 | 8 | 0.0245 | 0 |
| CBGA (%AR) | 011.02 | AOAC 2018.11, U | 183 | 0.1528 | 0.1503 | 0.1499 | 0.151 | 0.7 | 0.1339 | 8 | 0.0245 | 0 |
| CBGA (%AR) | 011.99 | Other | 141 | 0.2002 | 0.1810 | 0.2100 | 0.197 | 2.58 | 0.1339 | 8 | 0.0245 | 0 |
|  |  |  |  | THCV (%AR) |  |  |  |  |  |  |  |  |
| THCV (%AR) | 012.02 | AOAC 2018.11, U | 183 | <0.0020 <0.0020 <0.0020 |  |  | <0.0020 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 189 | <0.0050 <0.0050 <0.0050 |  |  | <0.0050 |  |  |  |  | 1 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score Rob Mean |  | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| THCV (%AR) | 012.99 | Other | 183 | <0.0200 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 | <0.0205 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 106 | <0.0300 | <0.0300 | <0.0300 | <0.0300 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 197 | <0.0483 | <0.0469 | <0.0483 | <0.0483 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 173 | 0.0000 | 0.0000 | 0.0000 | 0 |  |  |  |  | 0 |

Population of Lab Values

| Δ8-THC (%AR) |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%AR) | 013.02 | AOAC 2018.11, U | 183 | <0.0010 | <0.0010 | <0.0010 | <0.0010 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.03 | AOAC 2018.11, M | 139 | <0.0125 | <0.0125 | <0.0125 | <0.0125 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 183 | <0.0200 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 196 | <0.0205 | <0.0205 | <0.0205 | <0.0205 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 106 | <0.0300 | <0.0300 | <0.0300 | <0.0300 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 197 | <0.0483 | <0.0469 | <0.0483 | <0.0483 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 141 | 0.0000 | 0.0000 | 0.0000 | 0 | -0.77 | 0.0847 | 6 | 0.1097 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 173 | 0.0000 | 0.0000 | 0.0000 | 0 | -0.77 | 0.0847 | 6 | 0.1097 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 189 | 0.0127 | 0.0126 | 0.0126 | 0.0126 | -0.66 | 0.0847 | 6 | 0.1097 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 213 | 0.1220 | 0.1250 | 0.1270 | 0.125 | 0.36 | 0.0847 | 6 | 0.1097 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 138 | 0.1269 | 0.1395 | 0.1347 | 0.134 | 0.45 | 0.0847 | 6 | 0.1097 | 0 |
| Δ8-THC (%AR) | 013.01 | AOAC 2018.10 | 228 | 0.2397 | 0.2429 | 0.2289 | 0.237 | 1.39 | 0.0847 | 6 | 0.1097 | 0 |

|  |  |  | CBC (%AR) |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%AR) | 014.99 | Other | 149 | 0.1850 | 0.1800 | 0.1810 | 0.182 | -0.93 | 0.2077 | 10 | 0.0275 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%AR) | 014.99 | Other | 173 | 0.1880 | 0.1790 | 0.1800 | 0.182 | -0.92 | 0.2077 | 10 | 0.0275 | 0 |
| CBC (%AR) | 014.99 | Other | 155 | 0.1923 | 0.1816 | 0.1794 | 0.184 | -0.85 | 0.2077 | 10 | 0.0275 | 0 |
| CBC (%AR) | 014.99 | Other | 196 | 0.1953 | 0.1986 | 0.1871 | 0.194 | -0.51 | 0.2077 | 10 | 0.0275 | 0 |
| CBC (%AR) | 014.99 | Other | 197 | 0.2000 | 0.2009 | 0.2000 | 0.2 | -0.27 | 0.2077 | 10 | 0.0275 | 0 |
| CBC (%AR) | 014.99 | Other | 106 | 0.2000 | 0.2100 | 0.2000 | 0.203 | -0.16 | 0.2077 | 10 | 0.0275 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, U | 125 | 0.2167 | 0.2209 | 0.2051 | 0.214 | 0.24 | 0.2077 | 10 | 0.0275 | 0 |
| CBC (%AR) | 014.99 | Other | 183 | 0.2325 | 0.2253 | 0.2262 | 0.228 | 0.74 | 0.2077 | 10 | 0.0275 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, U | 183 | 0.2423 | 0.2379 | 0.2389 | 0.24 | 1.16 | 0.2077 | 10 | 0.0275 | 0 |
| CBC (%AR) | 014.99 | Other | 189 | 0.2467 | 0.2754 | 0.2467 | 0.256 | 1.77 | 0.2077 | 10 | 0.0275 | 0 |
|  |  |  |  | Cd (ug/g AR) |  |  |  |  |  |  |  |  |
| Cd (ug/g AR) | 101.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 | <0.2000 |  |  |  |  | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 196 | 0.1440 | 0.1390 | 0.1450 | 0.143 | -1.47 | 0.1686 | 7 | 0.0177 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 164 | 0.1406 | 0.1481 | 0.1735 | 0.154 | -0.83 | 0.1686 | 7 | 0.0177 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 102 | 0.1697 | 0.1642 | 0.1595 | 0.164 | -0.24 | 0.1686 | 7 | 0.0177 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 106 | 0.1710 | 0.1790 | 0.1780 | 0.176 | 0.42 | 0.1686 | 7 | 0.0177 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 197 | 0.1780 | 0.1789 | 0.1730 | 0.177 | 0.45 | 0.1686 | 7 | 0.0177 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 183 | 0.1800 | 0.1900 | 0.1800 | 0.183 | 0.83 | 0.1686 | 7 | 0.0177 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 141 | 0.1850 | 0.1850 | 0.1800 | 0.183 | 0.83 | 0.1686 | 7 | 0.0177 | 0 |
|  |  |  |  | As (ug/g AR) |  |  |  |  |  |  |  |  |
| As (ug/g AR) | 102.04 | ICP-MS | 197 | 0.2178 | 0.2479 | 0.2528 | 0.24 | -1.26 | 0.2592 | 8 | 0.0156 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| As (ug/g AR) | 102.04 | ICP-MS | 106 | 0.2300 | 0.2630 | 0.2480 | 0.247 | -0.78 | 0.2592 | 8 | 0.0156 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 102 | 0.2350 | 0.2680 | 0.2510 | 0.251 | -0.51 | 0.2592 | 8 | 0.0156 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 141 | 0.2760 | 0.2520 | 0.2350 | 0.254 | -0.31 | 0.2592 | 8 | 0.0156 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 164 | 0.2571 | 0.2695 | 0.2728 | 0.266 | 0.46 | 0.2592 | 8 | 0.0156 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 162 | 0.2619 | 0.3074 | 0.2276 | 0.266 | 0.41 | 0.2592 | 8 | 0.0156 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 196 | 0.2710 | 0.2600 | 0.2700 | 0.267 | 0.5 | 0.2592 | 8 | 0.0156 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 183 | 0.3200 | 0.3900 | 0.3200 | 0.343 | 5.38 | 0.2592 | 8 | 0.0156 | 0 |

Population of Lab Values

|  |  |  | Hg (ug/g AR) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Hg (ug/g AR) | 103.04 | ICP-MS | 141 | <0.0300 | <0.0300 | <0.0300 | <0.0300 | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 164 | <0.0500 | <0.0500 | <0.0500 | <0.0500 | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 197 | <0.0975 | <0.0952 | <0.0984 | <0.0975 | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 | <0.2000 | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 183 | 0.0050 | 0.0050 | 0.0040 | 0.0047 | 0 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 196 | 0.0084 | 0.0081 | 0.0080 | 0.0082 | 0 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 106 | 0.0080 | 0.0090 | 0.0090 | 0.0087 | 0 |
| Hg (ug/g AR) | 103.99 | Other | 102 | 0.0300 | 0.0520 | 0.0495 | 0.0438 | 0 |

|  |  |  | Pb (ug/g AR) |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Pb (ug/g AR) | 104.04 | ICP-MS | 197 | 0.7048 | 0.7490 | 0.7672 | 0.74 | -1.22 | 0.8486 | 8 | 0.0886 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 102 | 0.7420 | 0.7460 | 0.7420 | 0.743 | -1.19 | 0.8486 | 8 | 0.0886 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 141 | 0.9010 | 0.7710 | 0.7880 | 0.82 | -0.32 | 0.8486 | 8 | 0.0886 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Pb (ug/g AR) | 104.04 | ICP-MS | 183 | 0.8100 | 0.8400 | 0.8900 | 0.847 | -0.02 | 0.8486 | 8 | 0.0886 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 106 | 0.8380 | 0.8740 | 0.8670 | 0.86 | 0.13 | 0.8486 | 8 | 0.0886 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 164 | 0.8561 | 0.8866 | 0.9983 | 0.914 | 0.73 | 0.8486 | 8 | 0.0886 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 162 | 0.8316 | 0.9551 | 0.9561 | 0.914 | 0.74 | 0.8486 | 8 | 0.0886 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 196 | 0.9370 | 0.9880 | 0.9270 | 0.951 | 1.15 | 0.8486 | 8 | 0.0886 | 0 |

Population of Lab Values

Cu (ug/g AR)

Cu (ug/g AR)

105.04	ICP-MS

196

11.7000  10.1000	10.7000	10.8	0

|  |  |  | Ni (ug/g AR) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Ni (ug/g AR) | 106.04 | ICP-MS | 196 | 2.9600 | 2.8500 | 3.0100 | 2.94 | 0 |
| Ni (ug/g AR) | 106.04 | ICP-MS | 183 | 3.6700 | 3.7400 | 3.8300 | 3.75 | 0 |
| Ni (ug/g AR) | 106.04 | ICP-MS | 141 | 4.0400 | 4.1500 | 3.9300 | 4.04 | 0 |

|  |  |  | Cr (ug/g AR) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cr (ug/g AR) | 107.04 | ICP-MS | 196 | 5.1000 | 4.9500 | 5.3000 | 5.12 | 0 |
| Cr (ug/g AR) | 107.04 | ICP-MS | 141 | 6.6900 | 7.4700 | 6.6700 | 6.94 | 0 |

|  |  |  | Se (ug/g AR) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Se (ug/g AR) | 108.04 | ICP-MS | 141 | <0.4030 | <0.3990 | <0.3990 | <0.4030 | 1 |
| Se (ug/g AR) | 108.04 | ICP-MS | 197 | 0.1529 | 0.1522 | 0.1478 | 0.151 | 0 |
| Se (ug/g AR) | 108.04 | ICP-MS | 196 | 0.2480 | 0.2430 | 0.2010 | 0.231 | 0 |

| Analyte Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

Population of Lab Values

|  |  |  | Zn (ug/g AR) |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Zn (ug/g AR) | 111.04 | ICP-MS | 196 | 64.9000 55.3000 57.6000 | 59.3 | 0 |
| Zn (ug/g AR) | 111.04 | ICP-MS | 141 | 61.0000 61.4000 57.9000 | 60.1 | 0 |

Sb (ug/g AR)

Sb (ug/g AR)

112.04	ICP-MS

196

<0.0241  <0.0241	<0.0241	<0.0241	1

|  |  |  | (-)-alpha-Bisabolol (%AR) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 162 | 0.0285 | 0.0298 | 0.0296 | 0.0293 | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 196 | 0.0662 | 0.0721 | 0.0703 | 0.0695 | 0 |

|  |  |  | beta-Caryophyllene (%AR) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| beta-Caryophyllene (% | 202.02 | GC,MS | 196 | 0.0322 | 0.0338 | 0.0312 | 0.0324 | 0 |
| beta-Caryophyllene (% | 202.02 | GC,MS | 162 | 0.0349 | 0.0354 | 0.0334 | 0.0346 | 0 |

(-)-Caryophyllene oxide (%AR)

(-)-Caryophyllene oxide

203.02	GC,MS

196

<0.0158  <0.0158	<0.0158	<0.0158	1

|  |  |  | Geraniol (%AR) |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Geraniol (%AR) | 204.02 | GC,MS | 162 | <0.0100 <0.0100 <0.0100 | <0.0100 | 1 |
| Geraniol (%AR) | 204.02 | GC,MS | 196 | <0.2520 <0.2520 <0.2520 | <0.2520 | 1 |

|  |  |  | Lab |  |  |  | Lab |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Value | Z score |  | Rob Mean | n | Rob StDev | Flag |
|  |  |  |  | alpha-Humulene (%AR) |  |  |  |  |  |  |  |  |  |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 196 | 0.0131 | 0.0139 | 0.0135 | 0.0135 | 0 |  |  |  |  |  |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 162 | 0.0143 | 0.0145 | 0.0137 | 0.0142 | 0 |  |  |  |  |  |
|  |  |  |  | d-Limonene (%AR) |  |  |  |  |  |  |  |  |  |
| d-Limonene (%AR) | 206.02 | GC,MS | 196 | <0.0039 <0.0039 <0.0039 |  |  | <0.0039 | 1 |  |  |  |  |  |
| d-Limonene (%AR) | 206.02 | GC,MS | 162 | <0.0100 <0.0100 <0.0100 |  |  | <0.0100 | 1 |  |  |  |  |  |
|  |  |  |  | Linalool (%AR) |  |  |  |  |  |  |  |  |  |
| Linalool (%AR) | 207.02 | GC,MS | 196 | <0.0039 <0.0039 <0.0039 |  |  | <0.0039 | 1 |  |  |  |  |  |
| Linalool (%AR) | 207.02 | GC,MS | 162 | <0.0100 <0.0100 <0.0100 |  |  | <0.0100 | 1 |  |  |  |  |  |
|  |  |  |  | beta-Myrcene (%AR) |  |  |  |  |  |  |  |  |  |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 196 | <0.0039 <0.0039 <0.0039 |  |  | <0.0039 | 1 |  |  |  |  |  |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 162 | <0.0100 <0.0100 <0.0100 |  |  | <0.0100 | 1 |  |  |  |  |  |
|  |  |  |  | Ocimene (%AR) |  |  |  |  |  |  |  |  |  |
| Ocimene (%AR) | 210.02 | GC,MS | 196 | <0.0039 <0.0039 <0.0039 |  |  | <0.0039 | 1 |  |  |  |  |  |
| Ocimene (%AR) | 210.02 | GC,MS | 162 | <0.0100 <0.0100 <0.0100 |  |  | <0.0100 | 1 |  |  |  |  |  |

Population of Lab Values

|  |  |  | Lab |  |  |  |  | Lab |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Analyte | Code | Method | Num |  | Result1 | Result2 | Result3 | Value | Z score |  | Rob Mean | n | Rob StDev | Flag |
|  |  |  |  |  | alpha-Pinene (%AR) |  |  |  |  |  |  |  |  |  |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 196 |  | <0.0039 <0.0039 <0.0039 |  |  | <0.0039 | 1 |  |  |  |  |  |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 162 |  | <0.0100 <0.0100 <0.0100 |  |  | <0.0100 | 1 |  |  |  |  |  |
|  |  |  |  |  | (-)-beta-Pinene (%AR) |  |  |  |  |  |  |  |  |  |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 196 |  | <0.0039 <0.0039 <0.0039 |  |  | <0.0039 | 1 |  |  |  |  |  |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 162 |  | <0.0100 <0.0100 <0.0100 |  |  | <0.0100 | 1 |  |  |  |  |  |
|  |  |  |  |  | alpha-Terpinene (%AR) |  |  |  |  |  |  |  |  |  |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 196 |  | <0.0039 <0.0039 <0.0039 |  |  | <0.0039 | 1 |  |  |  |  |  |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 162 |  | <0.0100 <0.0100 <0.0100 |  |  | <0.0100 | 1 |  |  |  |  |  |
|  |  |  |  |  | gamma-Terpinene (%AR) |  |  |  |  |  |  |  |  |  |
| gamma-Terpinene (%A | 214.02 | GC,MS | 196 |  | <0.0039 <0.0039 <0.0039 |  |  | <0.0039 | 1 |  |  |  |  |  |
| gamma-Terpinene (%A | 214.02 | GC,MS | 162 |  | <0.0100 <0.0100 <0.0100 |  |  | <0.0100 | 1 |  |  |  |  |  |
|  |  |  |  |  | alpha-Terpinolene (%AR) |  |  |  |  |  |  |  |  |  |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 196 |  | <0.0039 <0.0039 <0.0039 |  |  | <0.0039 | 1 |  |  |  |  |  |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 162 |  | <0.0100 <0.0100 <0.0100 |  |  | <0.0100 | 1 |  |  |  |  |  |
|  |  |  |  |  | Moisture (%AR) |  |  |  |  |  |  |  |  |  |
| Moisture (%AR) | 500.70 | Commerc. Anal. |  | 153 | 4.8100 | 5.3100 | 5.3400 | 5.15 | -1.64 | 7.9314 |  | 15 | 1.6965 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 189 | 6.2650 | 6.2650 | 6.2650 | 6.26 | -0.98 | 7.9314 | 15 | 1.6965 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 116 | 6.2100 | 6.5900 | 6.6500 | 6.48 | -0.85 | 7.9314 | 15 | 1.6965 | 0 |
| Moisture (%AR) | 500.99 | Other | 103 | 6.5100 | 6.6900 | 6.6700 | 6.62 | -0.77 | 7.9314 | 15 | 1.6965 | 0 |
| Moisture (%AR) | 500.99 | Other | 102 | 6.9000 | 6.6000 | 6.8000 | 6.77 | -0.69 | 7.9314 | 15 | 1.6965 | 0 |
| Moisture (%AR) | 500.72 | NIR | 138 | 7.5500 | 7.2000 | 7.3500 | 7.37 | -0.33 | 7.9314 | 15 | 1.6965 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 106 | 7.3500 | 7.4100 | 7.5200 | 7.43 | -0.3 | 7.9314 | 15 | 1.6965 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 210 | 8.2800 | 7.8900 | 7.5800 | 7.92 | -0.01 | 7.9314 | 15 | 1.6965 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 196 | 8.0500 | 7.9600 | 8.0500 | 8.02 | 0.05 | 7.9314 | 15 | 1.6965 | 0 |
| Moisture (%AR) | 500.99 | Other | 165 | 8.4000 | 8.1000 | 8.5000 | 8.33 | 0.24 | 7.9314 | 15 | 1.6965 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 183 | 9.1600 | 8.1000 | 9.4100 | 8.89 | 0.57 | 7.9314 | 15 | 1.6965 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 182 | 8.4950 | 9.9600 | 10.2350 | 9.56 | 0.96 | 7.9314 | 15 | 1.6965 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 192 | 9.7013 | 9.6971 | 10.0280 | 9.81 | 1.11 | 7.9314 | 15 | 1.6965 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 199 | 9.9085 | 9.9908 | 10.1716 | 10 | 1.23 | 7.9314 | 15 | 1.6965 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 162 | 9.7100 | 10.7800 | 9.8000 | 10.1 | 1.28 | 7.9314 | 15 | 1.6965 | 0 |
|  |  |  |  | Δ9-THC (%DW) |  |  |  |  |  |  |  |  |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, U | 103 | 0.1225 | 0.1181 | 0.1187 | 0.12 | -1.67 | 0.1514 | 8 | 0.0189 | 0 |
| Δ9-THC (%DW) | 501.30 | Other LC-MS | 102 | 0.1420 | 0.1410 | 0.1420 | 0.142 | -0.51 | 0.1514 | 8 | 0.0189 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 165 | 0.1510 | 0.1455 | 0.1460 | 0.147 | -0.2 | 0.1514 | 8 | 0.0189 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 116 | 0.1474 | 0.1475 | 0.1507 | 0.149 | -0.15 | 0.1514 | 8 | 0.0189 | 0 |
| Δ9-THC (%DW) | 501.01 | AOAC 2018.10 | 210 | 0.1500 | 0.1510 | 0.1490 | 0.15 | -0.07 | 0.1514 | 8 | 0.0189 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 189 | 0.1551 | 0.1542 | 0.1543 | 0.155 | 0.17 | 0.1514 | 8 | 0.0189 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score Rob Mean |  | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 162 | 0.1670 | 0.1686 | 0.1625 | 0.166 | 0.77 | 0.1514 | 8 | 0.0189 | 0 |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, U | 199 | 0.1820 | 0.1839 | 0.1837 | 0.183 | 1.68 | 0.1514 | 8 | 0.0189 | 0 |

Population of Lab Values

| Δ9-THCA (%DW) |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, U | 199 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| Δ9-THCA (%DW) | 502.30 | Other LC-MS | 102 | <0.0510 | <0.0530 | <0.0520 | <0.0510 |  |  |  |  | 1 |
| Δ9-THCA (%DW) | 502.02 | AOAC 2018.11, U | 103 | 0.0221 | 0.0222 | 0.0233 | 0.0225 | -1.46 | 0.0312 | 7 | 0.0059 | 0 |
| Δ9-THCA (%DW) | 502.01 | AOAC 2018.10 | 210 | 0.0300 | 0.0280 | 0.0300 | 0.0293 | -0.31 | 0.0312 | 7 | 0.0059 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 165 | 0.0277 | 0.0317 | 0.0296 | 0.0297 | -0.25 | 0.0312 | 7 | 0.0059 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 162 | 0.0340 | 0.0294 | 0.0316 | 0.0317 | 0.09 | 0.0312 | 7 | 0.0059 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 189 | 0.0363 | 0.0206 | 0.0381 | 0.0317 | 0.09 | 0.0312 | 7 | 0.0059 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 116 | 0.0333 | 0.0342 | 0.0323 | 0.0333 | 0.36 | 0.0312 | 7 | 0.0059 | 0 |
| Δ9-THCA (%DW) | 502.01 | AOAC 2018.10 | 228 | 0.0549 | 0.0556 | 0.0537 | 0.0547 | 3.99 | 0.0312 | 7 | 0.0059 | 0 |

| CBD (%DW) |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%DW) | 503.02 | AOAC 2018.11, U | 103 | 4.6334 | 4.8386 | 4.7049 | 4.73 | -1.13 | 5.0397 | 6 | 0.2792 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 162 | 5.0016 | 4.8083 | 4.8412 | 4.88 | -0.56 | 5.0397 | 6 | 0.2792 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 189 | 4.8792 | 4.8726 | 4.8807 | 4.88 | -0.58 | 5.0397 | 6 | 0.2792 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 116 | 5.1186 | 5.0839 | 5.2435 | 5.15 | 0.39 | 5.0397 | 6 | 0.2792 | 0 |
| CBD (%DW) | 503.01 | AOAC 2018.10 | 210 | 5.2780 | 5.2500 | 5.2290 | 5.25 | 0.76 | 5.0397 | 6 | 0.2792 | 0 |
| CBD (%DW) | 503.02 | AOAC 2018.11, U | 199 | 5.2258 | 5.3717 | 5.4543 | 5.35 | 1.11 | 5.0397 | 6 | 0.2792 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | CBDA (%DW) |  |  |  |  |  |  |  |  |
| CBDA (%DW) | 504.02 | AOAC 2018.11, U | 103 | 2.7010 | 2.7618 | 2.7882 | 2.75 | -0.96 | 3.0906 | 7 | 0.3556 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 162 | 2.9218 | 2.7896 | 2.8764 | 2.86 | -0.64 | 3.0906 | 7 | 0.3556 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 189 | 2.9019 | 2.8974 | 2.9028 | 2.9 | -0.53 | 3.0906 | 7 | 0.3556 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 116 | 3.0269 | 2.9666 | 2.9600 | 2.98 | -0.3 | 3.0906 | 7 | 0.3556 | 0 |
| CBDA (%DW) | 504.02 | AOAC 2018.11, U | 199 | 3.1052 | 3.1391 | 3.1226 | 3.12 | 0.09 | 3.0906 | 7 | 0.3556 | 0 |
| CBDA (%DW) | 504.01 | AOAC 2018.10 | 210 | 3.4030 | 3.4010 | 3.3650 | 3.39 | 0.84 | 3.0906 | 7 | 0.3556 | 0 |
| CBDA (%DW) | 504.01 | AOAC 2018.10 | 228 | 4.1399 | 4.1482 | 4.1376 | 4.14 | 2.96 | 3.0906 | 7 | 0.3556 | 0 |
|  |  |  |  | CBN (%DW) |  |  |  |  |  |  |  |  |
| CBN (%DW) | 505.10 | Other LC-UV | 162 | 0.0610 | 0.0605 | 0.0601 | 0.0605 | -0.96 | 0.0702 | 7 | 0.0101 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 189 | 0.0625 | 0.0637 | 0.0623 | 0.0628 | -0.73 | 0.0702 | 7 | 0.0101 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 192 | 0.0000 | 0.0937 | 0.1012 | 0.065 | -0.52 | 0.0702 | 7 | 0.0101 | 0 |
| CBN (%DW) | 505.02 | AOAC 2018.11, U | 199 | 0.0666 | 0.0694 | 0.0679 | 0.068 | -0.22 | 0.0702 | 7 | 0.0101 | 0 |
| CBN (%DW) | 505.30 | GC-FID | 116 | 0.0705 | 0.0735 | 0.0703 | 0.0714 | 0.12 | 0.0702 | 7 | 0.0101 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 116 | 0.0766 | 0.0788 | 0.0795 | 0.0783 | 0.8 | 0.0702 | 7 | 0.0101 | 0 |
| CBN (%DW) | 505.30 | GC-FID | 192 | 0.0929 | 0.1064 | 0.0993 | 0.0995 | 2.9 | 0.0702 | 7 | 0.0101 | 0 |
|  |  |  |  | Total Δ9-THC (%DW) |  |  |  |  |  |  |  |  |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 182 | 0.1365 | 0.1135 | 0.1263 | 0.125 | -2.49 | 0.1728 | 14 | 0.019 | 0 |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, U | 103 | 0.1419 | 0.1376 | 0.1391 | 0.14 | -1.75 | 0.1728 | 14 | 0.019 | 0 |
| Total Δ9-THC (%DW) | 506.30 | Other LC-MS | 102 | 0.1600 | 0.1600 | 0.1600 | 0.16 | -0.67 | 0.1728 | 14 | 0.019 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 192 | 0.1530 | 0.1768 | 0.1617 | 0.164 | -0.47 | 0.1728 | 14 | 0.019 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 153 | 0.1510 | 0.1840 | 0.1670 | 0.167 | -0.29 | 0.1728 | 14 | 0.019 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 116 | 0.1701 | 0.1683 | 0.1626 | 0.167 | -0.3 | 0.1728 | 14 | 0.019 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 138 | 0.1899 | 0.1619 | 0.1688 | 0.174 | 0.04 | 0.1728 | 14 | 0.019 | 0 |
| Total Δ9-THC (%DW) | 506.01 | AOAC 2018.10 | 210 | 0.1760 | 0.1760 | 0.1750 | 0.176 | 0.15 | 0.1728 | 14 | 0.019 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 165 | 0.1786 | 0.1772 | 0.1756 | 0.177 | 0.23 | 0.1728 | 14 | 0.019 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 116 | 0.1766 | 0.1775 | 0.1790 | 0.178 | 0.26 | 0.1728 | 14 | 0.019 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 189 | 0.1869 | 0.1722 | 0.1876 | 0.182 | 0.5 | 0.1728 | 14 | 0.019 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 192 | 0.2207 | 0.1756 | 0.1754 | 0.191 | 0.94 | 0.1728 | 14 | 0.019 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 162 | 0.1971 | 0.1939 | 0.1907 | 0.194 | 1.11 | 0.1728 | 14 | 0.019 | 0 |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, U | 199 | 0.2039 | 0.2058 | 0.2056 | 0.205 | 1.7 | 0.1728 | 14 | 0.019 | 0 |
|  |  |  |  | Total CBD (%DW) |  |  |  |  |  |  |  |  |
| Total CBD (%DW) | 507.02 | AOAC 2018.11, U | 103 | 7.0022 | 7.2607 | 7.1501 | 7.14 | -1.07 | 8.0646 | 10 | 0.8679 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 162 | 7.5645 | 7.2551 | 7.3636 | 7.39 | -0.77 | 8.0646 | 10 | 0.8679 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 189 | 7.4242 | 7.4136 | 7.4265 | 7.42 | -0.74 | 8.0646 | 10 | 0.8679 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 116 | 7.7732 | 7.6856 | 7.8394 | 7.77 | -0.34 | 8.0646 | 10 | 0.8679 | 0 |
| Total CBD (%DW) | 507.40 | GC-FID | 192 | 7.7040 | 8.3806 | 7.3779 | 7.82 | -0.28 | 8.0646 | 10 | 0.8679 | 0 |
| Total CBD (%DW) | 507.40 | GC-FID | 116 | 7.9739 | 7.8283 | 8.3909 | 8.06 | 0 | 8.0646 | 10 | 0.8679 | 0 |
| Total CBD (%DW) | 507.02 | AOAC 2018.11, U | 199 | 7.9491 | 8.1247 | 8.1928 | 8.09 | 0.03 | 8.0646 | 10 | 0.8679 | 0 |
| Total CBD (%DW) | 507.01 | AOAC 2018.10 | 210 | 8.2620 | 8.2330 | 8.1800 | 8.23 | 0.18 | 8.0646 | 10 | 0.8679 | 0 |
| Total CBD (%DW) | 507.40 | GC-FID | 138 | 9.1518 | 9.5155 | 9.6342 | 9.43 | 1.58 | 8.0646 | 10 | 0.8679 | 0 |

Population of Lab Values

|  |  |  | Lab |  |  |  | Lab |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Analyte | Code | Method | Num | Result1 Result2 Result3 |  |  | Value | Z score |  | Rob Mean | n | Rob StDev | Flag |
| Total CBD (%DW) | 507.10 | Other LC-UV | 192 | 10.7838 12.0681 10.0654 |  |  | 11 | 3.35 |  | 8.0646 | 10 | 0.8679 | 0 |
|  |  |  |  | CBDV (%DW) |  |  |  |  |  |  |  |  |  |
| CBDV (%DW) | 508.99 | Other | 189 | 0.0343 | 0.0337 | 0.0297 | 0.0326 | 0 |  |  |  |  |  |
| CBDV (%DW) | 508.99 | Other | 162 | 0.0377 | 0.0327 | 0.0335 | 0.0346 | 0 |  |  |  |  |  |
| CBDV (%DW) | 508.02 | AOAC 2018.11, U | 210 | 0.0420 | 0.0420 | 0.0410 | 0.0417 | 0 |  |  |  |  |  |
|  |  |  |  | CBDVA (%DW) |  |  |  |  |  |  |  |  |  |
| CBDVA (%DW) | 509.01 | AOAC 2018.10 | 210 | 0.0280 | 0.0190 | 0.0240 | 0.0237 | 0 |  |  |  |  |  |
|  |  |  |  | CBG (%DW) |  |  |  |  |  |  |  |  |  |
| CBG (%DW) | 510.99 | Other | 162 | 0.1786 | 0.1762 | 0.1712 | 0.175 | 0 |  |  |  |  |  |
| CBG (%DW) | 510.01 | AOAC 2018.10 | 210 | 0.1880 | 0.1890 | 0.1920 | 0.19 | 0 |  |  |  |  |  |
| CBG (%DW) | 510.99 | Other | 189 | 0.2074 | 0.2078 | 0.2088 | 0.208 | 0 |  |  |  |  |  |
|  |  |  |  | CBGA (%DW) |  |  |  |  |  |  |  |  |  |
| CBGA (%DW) | 511.01 | AOAC 2018.10 | 210 | 0.0640 | 0.0540 | 0.0620 | 0.06 | 0 |  |  |  |  |  |
| CBGA (%DW) | 511.99 | Other | 189 | 0.1473 | 0.1493 | 0.1483 | 0.148 | 0 |  |  |  |  |  |
| CBGA (%DW) | 511.99 | Other | 162 | 0.1477 | 0.1735 | 0.1565 | 0.159 | 0 |  |  |  |  |  |
|  |  |  |  | THCV (%DW) |  |  |  |  |  |  |  |  |  |
| THCV (%DW) | 512.99 | Other | 189 | <0.0050 <0.0050 <0.0050 |  |  | <0.0050 | 1 |  |  |  |  |  |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 Result2 Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| THCV (%DW) | 512.02 | AOAC 2018.11, U | 210 | <0.0130 <0.0130 <0.0130 | <0.0130 |  |  |  |  | 1 |

Population of Lab Values

| Δ8-THC (%DW) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%DW) | 513.99 | Other | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| Δ8-THC (%DW) | 513.99 | Other | 189 | 0.0135 | 0.0135 | 0.0135 | 0.0135 | 0 |
| Δ8-THC (%DW) | 513.01 | AOAC 2018.10 | 210 | 0.0150 | 0.0150 | 0.0130 | 0.0143 | 0 |

|  |  |  | CBC (%DW) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%DW) | 514.99 | Other | 162 | 0.2333 | 0.2378 | 0.2383 | 0.236 | 0 |
| CBC (%DW) | 514.99 | Other | 189 | 0.2631 | 0.2938 | 0.2631 | 0.273 | 0 |

- Flag definitions: 1 (Lab Value not used because there were less than 2 numeric lab results), 2 (Lab Value not used because there were less than 2 numeric lab results greater than 0).

- Robust statistics performed to determine robust mean and Z scores only if there were 6 or more numeric lab values.

- AR refers to concentration on as-received basis. DW refers to concentration on dry weight basis.