Hemp Proficiency Testing

All Labs Precision HM25NOV-1 (Hemp)

| Sample: HM25NOV-1 (Hemp) |  |  |  | All Labs Precision |  |  |  |  | Issue Date: |  | January 27, 2026 |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  |  |  |  | Population of Lab RSDr |  |  |  |
| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
|  |  |  |  | Δ9-THC (%AR) |  |  |  |  |  |  |  |  |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 189 | 0.1687 | 0.1687 | 0.1686 | 0.03 | 0.01 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 183 | 0.18 | 0.182 | 0.181 | 0.55 | 0.11 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 116 | 0.1651 | 0.1673 | 0.1652 | 0.75 | 0.14 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 148 | 0.148 | 0.145 | 0.145 | 1.19 | 0.22 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 199 | 0.1635 | 0.1665 | 0.1625 | 1.27 | 0.24 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 164 | 0.1434 | 0.1397 | 0.1398 | 1.5 | 0.28 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 154 | 0.1408 | 0.1442 | 0.1453 | 1.64 | 0.31 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 144 | 0.1449 | 0.1431 | 0.1483 | 1.82 | 0.34 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 161 | 0.1317 | 0.1371 | 0.1332 | 2.08 | 0.39 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 101 | 0.173 | 0.166 | 0.171 | 2.12 | 0.41 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 142 | 0.1529 | 0.1578 | 0.1514 | 2.17 | 0.41 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 102 | 0.166 | 0.159 | 0.164 | 2.21 | 0.42 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 152 | 0.2128 | 0.2179 | 0.2248 | 2.76 | 0.55 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 173 | 0.16 | 0.167 | 0.158 | 2.92 | 0.56 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 184 | 0.1661 | 0.1592 | 0.1561 | 3.19 | 0.61 | 3.45 | 0.03 | 21.2 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 106 | 0.16 | 0.15 | 0.16 | 3.69 | 0.7 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 213 | 0.0955 | 0.0945 | 0.102 | 4.18 | 0.74 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 175 | 0.16 | 0.162 | 0.15 | 4.09 | 0.78 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 149 | 0.141 | 0.144 | 0.153 | 4.28 | 0.8 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 165 | 0.1513 | 0.1409 | 0.1559 | 5.14 | 0.97 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 163 | 0.149 | 0.163 | 0.167 | 5.92 | 1.13 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 192 | 0.1475 | 0.14 | 0.159 | 6.43 | 1.21 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 103 | 0.1398 | 0.1602 | 0.1396 | 8.08 | 1.52 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 115 | 0.1416 | 0.1379 | 0.117 | 10 | 1.86 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.03 | AOAC 2018.11, M | 139 | 0.1516 | 0.1786 | 0.1426 | 11.9 | 2.26 | 3.45 | 0.03 | 21.2 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 204 | 0.228 | 0.162 | 0.1596 | 21.2 | 4.12 | 3.45 | 0.03 | 21.2 | 0 |
|  |  |  |  | Δ9-THCA (%AR) |  |  |  |  |  |  |  |  |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 173 | 0 | 0 | 0 |  |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 106 | 0.04 | 0.04 | 0.04 | 0 | 0 | 8.9 | 0 | 98.3 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 116 | 0.0154 | 0.0151 | 0.0154 | 1.13 | 0.15 | 8.9 | 0 | 98.3 | 0 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 164 | 0.0151 | 0.0148 | 0.0147 | 1.4 | 0.19 | 8.9 | 0 | 98.3 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 148 | 0.028 | 0.029 | 0.028 | 2.04 | 0.3 | 8.9 | 0 | 98.3 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 144 | 0.017 | 0.0166 | 0.0174 | 2.35 | 0.32 | 8.9 | 0 | 98.3 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 189 | 0.0312 | 0.0326 | 0.0328 | 2.71 | 0.41 | 8.9 | 0 | 98.3 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 165 | 0.0196 | 0.0195 | 0.0207 | 3.34 | 0.47 | 8.9 | 0 | 98.3 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 152 | 0.1077 | 0.1021 | 0.1065 | 2.8 | 0.5 | 8.9 | 0 | 98.3 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 142 | 0.0082 | 0.009 | 0.0087 | 4.68 | 0.58 | 8.9 | 0 | 98.3 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 192 | 0.0974 | 0.0905 | 0.0952 | 3.74 | 0.66 | 8.9 | 0 | 98.3 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 184 | 0.0156 | 0.0161 | 0.0185 | 9.26 | 1.26 | 8.9 | 0 | 98.3 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 101 | 0.015 | 0.017 | 0.019 | 11.8 | 1.6 | 8.9 | 0 | 98.3 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 115 | 0.0151 | 0.0132 | 0.0117 | 12.8 | 1.68 | 8.9 | 0 | 98.3 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 183 | 0.013 | 0.017 | 0.014 | 14.2 | 1.89 | 8.9 | 0 | 98.3 | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 103 | 0.0113 | 0.0141 | 0.0165 | 18.6 | 2.46 | 8.9 | 0 | 98.3 | 0 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 175 | 0.018 | 0.01 | 0.012 | 31.2 | 4.09 | 8.9 | 0 | 98.3 | 0 |
| Δ9-THCA (%AR) | 002.03 | AOAC 2018.11, M | 139 | 0.0125 | 0.0162 | 0.0078 | 34.6 | 4.48 | 8.9 | 0 | 98.3 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 204 | 0.0524 | 0.0119 | 0.0094 | 98.3 | 14.12 | 8.9 | 0 | 98.3 | 0 |
|  |  |  |  | CBD (%AR) |  |  |  |  |  |  |  |  |
| CBD (%AR) | 003.10 | Other LC-UV | 189 | 5.7148 | 5.7194 | 5.7203 | 0.05 | 0.02 | 2.75 | 0.05 | 10.4 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 204 | 5.8516 | 5.9024 | 5.8547 | 0.48 | 0.16 | 2.75 | 0.05 | 10.4 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 173 | 5.454 | 5.442 | 5.513 | 0.7 | 0.22 | 2.75 | 0.05 | 10.4 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 149 | 5.363 | 5.43 | 5.364 | 0.71 | 0.23 | 2.75 | 0.05 | 10.4 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 106 | 6.22 | 6.19 | 6.31 | 1 | 0.33 | 2.75 | 0.05 | 10.4 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 163 | 6.536 | 6.648 | 6.679 | 1.14 | 0.38 | 2.75 | 0.05 | 10.4 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 199 | 5.6465 | 5.5755 | 5.509 | 1.23 | 0.4 | 2.75 | 0.05 | 10.4 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 164 | 5.6587 | 5.5525 | 5.5127 | 1.35 | 0.44 | 2.75 | 0.05 | 10.4 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 144 | 5.9242 | 5.753 | 5.9039 | 1.6 | 0.52 | 2.75 | 0.05 | 10.4 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 142 | 5.2965 | 5.5265 | 5.3265 | 2.32 | 0.75 | 2.75 | 0.05 | 10.4 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 161 | 4.9728 | 5.1782 | 5.2026 | 2.47 | 0.79 | 2.75 | 0.05 | 10.4 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 154 | 4.8868 | 4.936 | 4.6893 | 2.7 | 0.86 | 2.75 | 0.05 | 10.4 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 183 | 5.95 | 6.333 | 6.223 | 3.2 | 1.05 | 2.75 | 0.05 | 10.4 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 213 | 4.027 | 3.968 | 4.234 | 3.43 | 1.06 | 2.75 | 0.05 | 10.4 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 116 | 5.6989 | 6.1571 | 5.9145 | 3.87 | 1.27 | 2.75 | 0.05 | 10.4 | 0 |
| CBD (%AR) | 003.01 | AOAC 2018.10 | 152 | 5.5469 | 5.0422 | 4.9314 | 6.34 | 2.03 | 2.75 | 0.05 | 10.4 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 184 | 6.097 | 5.865 | 5.331 | 6.81 | 2.22 | 2.75 | 0.05 | 10.4 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 192 | 6.9895 | 6.101 | 6.6238 | 6.8 | 2.26 | 2.75 | 0.05 | 10.4 | 0 |
| CBD (%AR) | 003.03 | AOAC 2018.11, M | 139 | 5.0837 | 5.7605 | 4.6996 | 10.4 | 3.32 | 2.75 | 0.05 | 10.4 | 0 |
|  |  |  |  | CBDA (%AR) |  |  |  |  |  |  |  |  |
| CBDA (%AR) | 004.10 | Other LC-UV | 189 | 1.9892 | 1.988 | 1.9892 | 0.03 | 0.01 | 2.75 | 0.03 | 13.8 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 199 | 2.089 | 2.1065 | 2.0855 | 0.54 | 0.15 | 2.75 | 0.03 | 13.8 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 204 | 2.1466 | 2.1745 | 2.1514 | 0.69 | 0.19 | 2.75 | 0.03 | 13.8 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 173 | 1.843 | 1.831 | 1.817 | 0.71 | 0.2 | 2.75 | 0.03 | 13.8 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 164 | 1.8853 | 1.8567 | 1.8608 | 0.83 | 0.23 | 2.75 | 0.03 | 13.8 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 152 | 1.8449 | 1.8242 | 1.8054 | 1.08 | 0.3 | 2.75 | 0.03 | 13.8 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 106 | 2.11 | 2.16 | 2.13 | 1.18 | 0.33 | 2.75 | 0.03 | 13.8 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 163 | 2.122 | 2.177 | 2.205 | 1.95 | 0.55 | 2.75 | 0.03 | 13.8 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 149 | 1.983 | 1.958 | 1.881 | 2.74 | 0.76 | 2.75 | 0.03 | 13.8 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 116 | 1.8643 | 1.9724 | 1.9175 | 2.82 | 0.78 | 2.75 | 0.03 | 13.8 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 184 | 2.133 | 2.242 | 2.128 | 2.97 | 0.84 | 2.75 | 0.03 | 13.8 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDA (%AR) | 004.10 | Other LC-UV | 183 | 2.109 | 2.256 | 2.178 | 3.37 | 0.95 | 2.75 | 0.03 | 13.8 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 144 | 2.0247 | 1.9224 | 2.0581 | 3.53 | 0.98 | 2.75 | 0.03 | 13.8 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 213 | 1.2755 | 1.271 | 1.366 | 4.11 | 1.07 | 2.75 | 0.03 | 13.8 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 154 | 1.422 | 1.4227 | 1.3228 | 4.14 | 1.09 | 2.75 | 0.03 | 13.8 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 142 | 1.6874 | 1.8263 | 1.6931 | 4.53 | 1.23 | 2.75 | 0.03 | 13.8 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 161 | 1.7696 | 1.7922 | 1.6299 | 5.08 | 1.38 | 2.75 | 0.03 | 13.8 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 192 | 2.3697 | 2.0141 | 2.1704 | 8.16 | 2.3 | 2.75 | 0.03 | 13.8 | 0 |
| CBDA (%AR) | 004.03 | AOAC 2018.11, M | 139 | 2.1094 | 2.4615 | 1.8716 | 13.8 | 3.88 | 2.75 | 0.03 | 13.8 | 0 |
|  |  |  |  | CBN (%AR) |  |  |  |  |  |  |  |  |
| CBN (%AR) | 005.10 | Other LC-UV | 189 | 0.0832 | 0.0828 | 0.0828 | 0.28 | 0.05 | 4.05 | 0.28 | 11.2 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 183 | 0.087 | 0.089 | 0.088 | 1.14 | 0.2 | 4.05 | 0.28 | 11.2 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 163 | 0.333 | 0.337 | 0.33 | 1.05 | 0.22 | 4.05 | 0.28 | 11.2 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 173 | 0.074 | 0.073 | 0.072 | 1.37 | 0.23 | 4.05 | 0.28 | 11.2 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 199 | 0.077 | 0.0775 | 0.0755 | 1.36 | 0.23 | 4.05 | 0.28 | 11.2 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 142 | 0.069 | 0.0708 | 0.069 | 1.49 | 0.25 | 4.05 | 0.28 | 11.2 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 213 | 0.0463 | 0.0459 | 0.0474 | 1.67 | 0.26 | 4.05 | 0.28 | 11.2 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 144 | 0.068 | 0.0689 | 0.0659 | 2.28 | 0.38 | 4.05 | 0.28 | 11.2 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 149 | 0.068 | 0.071 | 0.072 | 2.96 | 0.5 | 4.05 | 0.28 | 11.2 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 116 | 0.0951 | 0.0916 | 0.0897 | 2.97 | 0.52 | 4.05 | 0.28 | 11.2 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 154 | 0.0671 | 0.0731 | 0.0687 | 4.46 | 0.75 | 4.05 | 0.28 | 11.2 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 116 | 0.0815 | 0.0728 | 0.0789 | 5.75 | 0.98 | 4.05 | 0.28 | 11.2 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%AR) | 005.10 | Other LC-UV | 204 | 0.0875 | 0.0797 | 0.0785 | 5.97 | 1.03 | 4.05 | 0.28 | 11.2 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 192 | 0.1097 | 0.0952 | 0.1038 | 7.09 | 1.26 | 4.05 | 0.28 | 11.2 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 138 | 0.1093 | 0.0917 | 0.0922 | 10.3 | 1.81 | 4.05 | 0.28 | 11.2 | 0 |
| CBN (%AR) | 005.03 | AOAC 2018.11, M | 139 | 0.0567 | 0.0682 | 0.0563 | 11.2 | 1.84 | 4.05 | 0.28 | 11.2 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 106 | 0.09 | 0.11 | 0.11 | 11.2 | 1.99 | 4.05 | 0.28 | 11.2 | 0 |
|  |  |  |  | Total Δ9-THC (%AR) |  |  |  |  |  |  |  |  |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 189 | 0.196 | 0.1973 | 0.1974 | 0.4 | 0.08 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 116 | 0.1786 | 0.1805 | 0.1787 | 0.6 | 0.12 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 148 | 0.172 | 0.17 | 0.17 | 0.68 | 0.13 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 169 | 0.1508 | 0.1526 | 0.15 | 0.88 | 0.17 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 199 | 0.1854 | 0.1884 | 0.1844 | 1.12 | 0.22 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 183 | 0.192 | 0.197 | 0.193 | 1.36 | 0.27 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 164 | 0.1566 | 0.1527 | 0.1527 | 1.46 | 0.28 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 154 | 0.1408 | 0.1442 | 0.1453 | 1.64 | 0.31 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 138 | 0.1986 | 0.2031 | 0.2056 | 1.75 | 0.35 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 144 | 0.1598 | 0.1577 | 0.1636 | 1.86 | 0.36 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 101 | 0.186 | 0.181 | 0.188 | 1.95 | 0.38 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 152 | 0.3073 | 0.3075 | 0.3182 | 2.01 | 0.42 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 142 | 0.1601 | 0.1657 | 0.159 | 2.22 | 0.42 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 173 | 0.16 | 0.167 | 0.158 | 2.92 | 0.56 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 163 | 0.214 | 0.218 | 0.206 | 2.87 | 0.57 | 4.19 | 0.4 | 137 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 116 | 0.1745 | 0.1649 | 0.1673 | 2.96 | 0.57 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.50 | GC-MS | 231 | 0.1525 | 0.1475 | 0.1569 | 3.09 | 0.58 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 149 | 0.152 | 0.144 | 0.153 | 3.3 | 0.62 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 190 | 0.1503 | 0.1405 | 0.1506 | 3.91 | 0.73 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 175 | 0.1758 | 0.1708 | 0.1605 | 4.62 | 0.89 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 165 | 0.1685 | 0.158 | 0.1741 | 4.9 | 0.94 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 192 | 0.2329 | 0.2194 | 0.2424 | 4.99 | 1 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 102 | 0.1831 | 0.159 | 0.164 | 7.54 | 1.45 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 103 | 0.1497 | 0.1725 | 0.154 | 7.63 | 1.45 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 114 | 0.1711 | 0.1976 | 0.1658 | 9.56 | 1.85 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 170 | 0.1846 | 0.1523 | 0.1643 | 9.77 | 1.87 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 115 | 0.1548 | 0.1495 | 0.1273 | 10.1 | 1.9 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.03 | AOAC 2018.11, M | 139 | 0.1626 | 0.1927 | 0.1494 | 13.2 | 2.53 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 204 | 0.274 | 0.1724 | 0.1678 | 29.3 | 5.79 | 4.19 | 0.4 | 137 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 153 | 2.055 | 0.1555 | 0.1724 | 137 | 33.28 | 4.19 | 0.4 | 137 | 0 |
|  |  |  |  | Total CBD (%AR) |  |  |  |  |  |  |  |  |
| Total CBD (%AR) | 007.10 | Other LC-UV | 189 | 7.4594 | 7.4628 | 7.4648 | 0.04 | 0.01 | 2.85 | 0.04 | 11.3 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 173 | 7.07 | 7.048 | 7.11 | 0.44 | 0.15 | 2.85 | 0.04 | 11.3 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 204 | 7.7342 | 7.8094 | 7.7415 | 0.53 | 0.18 | 2.85 | 0.04 | 11.3 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 199 | 7.4786 | 7.4229 | 7.338 | 0.96 | 0.32 | 2.85 | 0.04 | 11.3 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 149 | 7.102 | 7.147 | 7.014 | 0.95 | 0.32 | 2.85 | 0.04 | 11.3 | 0 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%AR) | 007.30 | Other LC-MS | 164 | 7.3121 | 7.1808 | 7.1446 | 1.22 | 0.41 | 2.85 | 0.04 | 11.3 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 163 | 7.877 | 8.143 | 8.046 | 1.68 | 0.57 | 2.85 | 0.04 | 11.3 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 161 | 6.5247 | 6.7499 | 6.632 | 1.7 | 0.57 | 2.85 | 0.04 | 11.3 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 144 | 7.7003 | 7.4393 | 7.7093 | 2.01 | 0.68 | 2.85 | 0.04 | 11.3 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 142 | 6.7763 | 7.1281 | 6.8113 | 2.81 | 0.94 | 2.85 | 0.04 | 11.3 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 154 | 6.1339 | 6.1837 | 5.8494 | 2.98 | 0.98 | 2.85 | 0.04 | 11.3 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 116 | 7.5962 | 7.8869 | 8.0503 | 2.93 | 1 | 2.85 | 0.04 | 11.3 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 183 | 7.8 | 8.311 | 8.133 | 3.21 | 1.1 | 2.85 | 0.04 | 11.3 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 116 | 8.1192 | 7.8095 | 8.3782 | 3.51 | 1.21 | 2.85 | 0.04 | 11.3 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 169 | 7.701 | 8.2 | 7.626 | 3.98 | 1.36 | 2.85 | 0.04 | 11.3 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 138 | 7.876 | 7.93 | 8.496 | 4.24 | 1.45 | 2.85 | 0.04 | 11.3 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 152 | 7.1649 | 6.6421 | 6.5147 | 5.09 | 1.7 | 2.85 | 0.04 | 11.3 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 192 | 9.0677 | 7.8674 | 8.5273 | 7.08 | 2.45 | 2.85 | 0.04 | 11.3 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 170 | 7.4795 | 6.4449 | 6.9298 | 7.45 | 2.5 | 2.85 | 0.04 | 11.3 | 0 |
| Total CBD (%AR) | 007.03 | AOAC 2018.11, M | 139 | 6.9336 | 7.9193 | 6.341 | 11.3 | 3.79 | 2.85 | 0.04 | 11.3 | 0 |

Lab

Population of Lab RSDr

CBDV (%AR)

| CBDV (%AR) | 008.99 | Other | 189 | 0.032 | 0.0324 | 0.0324 | 0.72 | 0.11 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDV (%AR) | 008.99 | Other | 173 | 0.096 | 0.095 | 0.093 | 1.61 | 0.28 | 0 |
| CBDV (%AR) | 008.02 | AOAC 2018.11, U | 183 | 0.059 | 0.063 | 0.06 | 3.43 | 0.56 | 0 |
| CBDV (%AR) | 008.99 | Other | 142 | 0.0388 | 0.0429 | 0.041 | 5.02 | 0.78 | 0 |
| CBDV (%AR) | 008.99 | Other | 106 | 0.05 | 0.07 | 0.05 | 20.4 | 3.32 | 0 |

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

RSDr	HorRat(r)

Population of Lab RSDr Rob Mean	min	max

Flag

| CBDVA (%AR) |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDVA (%AR) | 009.02 | AOAC 2018.11, U | 183 | 0.038 | 0.038 | 0.039 | 1.51 | 0.23 | 0 |
| CBDVA (%AR) | 009.99 | Other | 142 | 0.0179 | 0.0192 | 0.0185 | 3.51 | 0.48 | 0 |
| CBDVA (%AR) | 009.01 | AOAC 2018.10 | 149 | 0.028 | 0.026 | 0.026 | 4.33 | 0.63 | 0 |

CBG (%AR)

| CBG (%AR) | 010.99 | Other | 106 | 0.25 | 0.25 | 0.25 | 0 | 0 | 1.71 | 0 | 4.01 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%AR) | 010.99 | Other | 189 | 0.3065 | 0.3032 | 0.3042 | 0.56 | 0.12 | 1.71 | 0 | 4.01 | 0 |
| CBG (%AR) | 010.02 | AOAC 2018.11, U | 183 | 0.263 | 0.262 | 0.259 | 0.8 | 0.16 | 1.71 | 0 | 4.01 | 0 |
| CBG (%AR) | 010.99 | Other | 204 | 0.2385 | 0.243 | 0.2443 | 1.26 | 0.25 | 1.71 | 0 | 4.01 | 0 |
| CBG (%AR) | 010.99 | Other | 173 | 0.217 | 0.215 | 0.21 | 1.68 | 0.34 | 1.71 | 0 | 4.01 | 0 |
| CBG (%AR) | 010.99 | Other | 142 | 0.1986 | 0.1879 | 0.2015 | 3.65 | 0.72 | 1.71 | 0 | 4.01 | 0 |
| CBG (%AR) | 010.01 | AOAC 2018.10 | 149 | 0.207 | 0.209 | 0.194 | 4.01 | 0.79 | 1.71 | 0 | 4.01 | 0 |

CBGA (%AR)

| CBGA (%AR) | 011.99 | Other | 106 | 0.11 | 0.11 | 0.11 | 0 | 0 | 2.14 | 0 | 5.36 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%AR) | 011.99 | Other | 189 | 0.1581 | 0.1581 | 0.1584 | 0.11 | 0.02 | 2.14 | 0 | 5.36 | 0 |
| CBGA (%AR) | 011.02 | AOAC 2018.11, U | 183 | 0.155 | 0.155 | 0.153 | 0.75 | 0.14 | 2.14 | 0 | 5.36 | 0 |
| CBGA (%AR) | 011.99 | Other | 173 | 0.131 | 0.131 | 0.125 | 2.69 | 0.49 | 2.14 | 0 | 5.36 | 0 |
| CBGA (%AR) | 011.01 | AOAC 2018.10 | 149 | 0.129 | 0.13 | 0.123 | 2.97 | 0.55 | 2.14 | 0 | 5.36 | 0 |
| CBGA (%AR) | 011.99 | Other | 142 | 0.1054 | 0.1114 | 0.1112 | 3.12 | 0.56 | 2.14 | 0 | 5.36 | 0 |

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%AR) | 011.99 | Other | 204 | 0.1386 | 0.1534 | 0.142 | 5.36 | 1 | 2.14 | 0 | 5.36 | 0 |

Lab

Population of Lab RSDr

THCV (%AR)

| THCV (%AR) | 012.99 | Other | 173 | 0 | 0 | 0 | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- |

Δ8-THC (%AR)

| Δ8-THC (%AR) | 013.99 | Other | 173 | 0 | 0 | 0 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%AR) | 013.99 | Other | 213 | 0.086 | 0.0851 | 0.0893 | 2.55 | 0.44 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 138 | 0.0758 | 0.0811 | 0.0794 | 3.44 | 0.59 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 189 | 0.0131 | 0.0136 | 0.0122 | 5.47 | 0.71 | 0 |

CBC (%AR)

| CBC (%AR) | 014.99 | Other | 106 | 0.24 | 0.24 | 0.24 | 0 | 0 | 1.34 | 0 | 13.1 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%AR) | 014.99 | Other | 189 | 0.279 | 0.2793 | 0.2791 | 0.05 | 0.01 | 1.34 | 0 | 13.1 | 0 |
| CBC (%AR) | 014.99 | Other | 173 | 0.215 | 0.213 | 0.213 | 0.54 | 0.11 | 1.34 | 0 | 13.1 | 0 |
| CBC (%AR) | 014.01 | AOAC 2018.10 | 149 | 0.226 | 0.224 | 0.229 | 1.11 | 0.22 | 1.34 | 0 | 13.1 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, U | 183 | 0.281 | 0.29 | 0.284 | 1.61 | 0.33 | 1.34 | 0 | 13.1 | 0 |
| CBC (%AR) | 014.99 | Other | 142 | 0.2175 | 0.2273 | 0.2182 | 2.47 | 0.49 | 1.34 | 0 | 13.1 | 0 |
| CBC (%AR) | 014.99 | Other | 204 | 0.3049 | 0.2512 | 0.2401 | 13.1 | 2.68 | 1.34 | 0 | 13.1 | 0 |

|  |  |  | Cd (ug/g AR) |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cd (ug/g AR) | 101.04 | ICP-MS | 142 | 0.1918 | 0.1915 | 0.1932 | 0.47 | 0.09 | 5.66 | 0.47 | 13.5 | 0 |

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cd (ug/g AR) | 101.04 | ICP-MS | 106 | 0.172 | 0.169 | 0.173 | 1.21 | 0.23 | 5.66 | 0.47 | 13.5 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 183 | 0.156 | 0.154 | 0.151 | 1.64 | 0.31 | 5.66 | 0.47 | 13.5 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 102 | 0.1821 | 0.1886 | 0.2044 | 5.98 | 1.17 | 5.66 | 0.47 | 13.5 | 0 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 164 | 0.179 | 0.188 | 0.151 | 11.2 | 2.15 | 5.66 | 0.47 | 13.5 | 0 |
| Cd (ug/g AR) | 101.03 | ICP-OES | 192 | 0.13 | 0.11 | 0.1 | 13.5 | 2.44 | 5.66 | 0.47 | 13.5 | 0 |

Lab

Population of Lab RSDr

As (ug/g AR)

| As (ug/g AR) | 102.04 | ICP-MS | 183 | 0.079 | 0.079 | 0.077 | 1.47 | 0.25 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| As (ug/g AR) | 102.04 | ICP-MS | 102 | 0.0722 | 0.0736 | 0.0814 | 6.55 | 1.11 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 142 | 0.0779 | 0.0698 | 0.0649 | 9.26 | 1.56 | 0 |
| As (ug/g AR) | 102.04 | ICP-MS | 164 | 0.074 | 0.061 | 0.076 | 11.6 | 1.95 | 0 |

Hg (ug/g AR)

| Hg (ug/g AR) | 103.03 | ICP-OES | 192 | 0 | 0 | 0 |  |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Hg (ug/g AR) | 103.04 | ICP-MS | 142 | 0.0088 | 0.0091 | 0.0086 | 2.85 | 0.35 | 0 |
| Hg (ug/g AR) | 103.99 | Other | 102 | 0.0086 | 0.0092 | 0.0091 | 3.59 | 0.44 | 0 |

|  |  |  | Pb (ug/g AR) |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Pb (ug/g AR) | 104.04 | ICP-MS | 102 | 0.357 | 0.3572 | 0.3541 | 0.49 | 0.1 | 6.9 | 0.49 | 24 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 183 | 0.283 | 0.287 | 0.287 | 0.81 | 0.17 | 6.9 | 0.49 | 24 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 164 | 0.271 | 0.298 | 0.31 | 6.82 | 1.42 | 6.9 | 0.49 | 24 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 162 | 0.3425 | 0.3965 | 0.364 | 7.39 | 1.59 | 6.9 | 0.49 | 24 | 0 |

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Pb (ug/g AR) | 104.04 | ICP-MS | 142 | 0.2475 | 0.21 | 0.2281 | 8.21 | 1.65 | 6.9 | 0.49 | 24 | 0 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 106 | 0.357 | 0.335 | 0.299 | 8.86 | 1.88 | 6.9 | 0.49 | 24 | 0 |
| Pb (ug/g AR) | 104.03 | ICP-OES | 192 | 0.54 | 0.56 | 0.35 | 24 | 5.39 | 6.9 | 0.49 | 24 | 0 |

Lab

Population of Lab RSDr

Cu (ug/g AR)

Cu (ug/g AR)

105.04

ICP-MS

142

13.5016	13.3583	14.1284	3	1.11	0

Ni (ug/g AR)

| Ni (ug/g AR) | 106.04 | ICP-MS | 142 | 1.3414 | 1.3369 | 1.3084 | 1.35 | 0.35 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

Cr (ug/g AR)

| Cr (ug/g AR) | 107.04 | ICP-MS | 142 | 0.4008 | 0.45 | 0.4343 | 5.87 | 1.29 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

Se (ug/g AR)

| Se (ug/g AR) | 108.04 | ICP-MS | 142 | 0.0883 | 0.0782 | 0.0801 | 6.53 | 1.12 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

Ag (ug/g AR)

| Ag (ug/g AR) | 109.04 | ICP-MS | 142 | 0.0124 | 0.0131 | 0.0121 | 4.09 | 0.53 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

Zn (ug/g AR)

Zn (ug/g AR)

111.04

ICP-MS

142

58.324	59.6313	60.0322	1.51	0.7	0

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | (-)-alpha-Bisabolol (%AR) |  |  |  |  |  |  |  |  |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 183 | 0.11 | 0.111 | 0.111 | 0.52 | 0.09 |  |  |  | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 162 | 0.0314 | 0.0321 | 0.0272 | 8.77 | 1.3 |  |  |  | 0 |
| (-)-alpha-Bisabolol (%A | 201.01 | GC,FID | 142 | 1500 | 1400 | 1360 | 5.08 | 3.78 |  |  |  | 0 |
|  |  |  |  | beta-Caryophyllene (%AR) |  |  |  |  |  |  |  |  |
| beta-Caryophyllene (% | 202.02 | GC,MS | 183 | 0.042 | 0.043 | 0.042 | 1.36 | 0.21 |  |  |  | 0 |
| beta-Caryophyllene (% | 202.02 | GC,MS | 162 | 0.039 | 0.0392 | 0.0349 | 6.44 | 0.99 |  |  |  | 0 |
| beta-Caryophyllene (% | 202.01 | GC,FID | 142 | 665 | 644 | 643 | 1.91 | 1.26 |  |  |  | 0 |
|  |  |  | (-)-Caryophyllene oxide (%AR) |  |  |  |  |  |  |  |  |  |
| (-)-Caryophyllene oxide | 203.02 | GC,MS | 183 | 0.019 | 0.019 | 0.02 | 2.99 | 0.41 |  |  |  | 0 |
| (-)-Caryophyllene oxide | 203.01 | GC,FID | 142 | 289 | 284 | 279 | 1.76 | 1.03 |  |  |  | 0 |
|  |  |  |  | alpha-Humulene (%AR) |  |  |  |  |  |  |  |  |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 183 | 0.019 | 0.019 | 0.019 | 0 | 0 |  |  |  | 0 |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 162 | 0.018 | 0.0181 | 0.0149 | 10.7 | 1.46 |  |  |  | 0 |
| alpha-Humulene (%AR) | 205.01 | GC,FID | 142 | 302 | 306 | 317 | 2.52 | 1.49 |  |  |  | 0 |
|  |  |  |  | Moisture (%AR) |  |  |  |  |  |  |  |  |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 189 | 5.015 | 5.015 | 5.015 | 0 | 0 | 2.67 | 0 | 8.31 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 199 | 9.7345 | 9.715 | 9.729 | 0.1 | 0.04 | 2.67 | 0 | 8.31 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Moisture (%AR) | 500.99 | Other | 142 | 4.41 | 4.37 | 4.33 | 0.92 | 0.29 | 2.67 | 0 | 8.31 | 0 |
| Moisture (%AR) | 500.99 | Other | 165 | 7.1101 | 7.1938 | 7.0728 | 0.87 | 0.29 | 2.67 | 0 | 8.31 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 106 | 5.96 | 5.97 | 6.15 | 1.77 | 0.58 | 2.67 | 0 | 8.31 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 192 | 9.8627 | 9.6611 | 9.9905 | 1.69 | 0.6 | 2.67 | 0 | 8.31 | 0 |
| Moisture (%AR) | 500.99 | Other | 102 | 6 | 6 | 5.8 | 1.95 | 0.64 | 2.67 | 0 | 8.31 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 116 | 5.75 | 5.67 | 5.5 | 2.26 | 0.74 | 2.67 | 0 | 8.31 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 183 | 7.59 | 7.68 | 7.94 | 2.35 | 0.8 | 2.67 | 0 | 8.31 | 0 |
| Moisture (%AR) | 500.99 | Other | 103 | 6.08 | 6.16 | 6.43 | 2.95 | 0.97 | 2.67 | 0 | 8.31 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 153 | 3.46 | 3.44 | 3.19 | 4.47 | 1.34 | 2.67 | 0 | 8.31 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 154 | 4.86 | 4.83 | 5.31 | 5.38 | 1.72 | 2.67 | 0 | 8.31 | 0 |
| Moisture (%AR) | 500.70 | Commerc. Anal. | 162 | 7.77 | 8.82 | 8.82 | 7.16 | 2.47 | 2.67 | 0 | 8.31 | 0 |
| Moisture (%AR) | 500.50 | AOAC 930.15 | 210 | 8.93 | 7.74 | 9.01 | 8.31 | 2.87 | 2.67 | 0 | 8.31 | 0 |
|  |  |  |  | Δ9-THC (%DW) |  |  |  |  |  |  |  |  |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 189 | 0.1776 | 0.1776 | 0.1775 | 0.03 | 0.01 | 1.64 | 0.03 | 5.1 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 116 | 0.1749 | 0.1773 | 0.1751 | 0.76 | 0.15 | 1.64 | 0.03 | 5.1 | 0 |
| Δ9-THC (%DW) | 501.01 | AOAC 2018.10 | 210 | 0.171 | 0.17 | 0.167 | 1.23 | 0.24 | 1.64 | 0.03 | 5.1 | 0 |
| Δ9-THC (%DW) | 501.02 | AOAC 2018.11, U | 199 | 0.1811 | 0.1844 | 0.18 | 1.26 | 0.24 | 1.64 | 0.03 | 5.1 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 162 | 0.1808 | 0.1805 | 0.1864 | 1.82 | 0.35 | 1.64 | 0.03 | 5.1 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 154 | 0.148 | 0.1515 | 0.1535 | 1.84 | 0.35 | 1.64 | 0.03 | 5.1 | 0 |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 142 | 0.1599 | 0.1651 | 0.1582 | 2.23 | 0.43 | 1.64 | 0.03 | 5.1 | 0 |
| Δ9-THC (%DW) | 501.30 | Other LC-MS | 102 | 0.177 | 0.169 | 0.174 | 2.33 | 0.45 | 1.64 | 0.03 | 5.1 | 0 |

Population of Lab RSDr

Lab

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%DW) | 501.10 | Other LC-UV | 165 | 0.1629 | 0.1518 | 0.1678 | 5.1 | 0.97 | 1.64 | 0.03 | 5.1 | 0 |

Lab

Population of Lab RSDr

Δ9-THCA (%DW)

| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 116 | 0.0163 | 0.016 | 0.0163 | 1.07 | 0.14 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 189 | 0.0328 | 0.0343 | 0.0345 | 2.74 | 0.41 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 165 | 0.0211 | 0.021 | 0.0223 | 3.37 | 0.47 | 0 |
| Δ9-THCA (%DW) | 502.01 | AOAC 2018.10 | 210 | 0.014 | 0.015 | 0.014 | 4.03 | 0.53 | 0 |
| Δ9-THCA (%DW) | 502.10 | Other LC-UV | 142 | 0.0086 | 0.0094 | 0.0091 | 4.47 | 0.55 | 0 |

| CBD (%DW) |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%DW) | 503.10 | Other LC-UV | 189 | 6.0165 | 6.0214 | 6.0223 | 0.05 | 0.02 | 1.83 | 0.05 | 3.87 | 0 |
| CBD (%DW) | 503.01 | AOAC 2018.10 | 210 | 0.183 | 0.183 | 0.18 | 0.95 | 0.18 | 1.83 | 0.05 | 3.87 | 0 |
| CBD (%DW) | 503.02 | AOAC 2018.11, U | 199 | 6.2554 | 6.1754 | 6.1027 | 1.24 | 0.41 | 1.83 | 0.05 | 3.87 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 162 | 7.0036 | 6.7519 | 6.8016 | 1.95 | 0.65 | 1.83 | 0.05 | 3.87 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 142 | 5.5409 | 5.779 | 5.5676 | 2.32 | 0.75 | 1.83 | 0.05 | 3.87 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 154 | 5.1364 | 5.1865 | 4.9523 | 2.42 | 0.77 | 1.83 | 0.05 | 3.87 | 0 |
| CBD (%DW) | 503.10 | Other LC-UV | 116 | 6.0396 | 6.5251 | 6.268 | 3.87 | 1.28 | 1.83 | 0.05 | 3.87 | 0 |

CBDA (%DW)

| CBDA (%DW) | 504.10 | Other LC-UV | 189 | 2.0943 | 2.0929 | 2.0942 | 0.04 | 0.01 | 2.3 | 0.04 | 4.53 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDA (%DW) | 504.02 | AOAC 2018.11, U | 199 | 2.3143 | 2.3332 | 2.3103 | 0.53 | 0.15 | 2.3 | 0.04 | 4.53 | 0 |
| CBDA (%DW) | 504.01 | AOAC 2018.10 | 210 | 6.527 | 6.519 | 6.449 | 0.66 | 0.22 | 2.3 | 0.04 | 4.53 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDA (%DW) | 504.10 | Other LC-UV | 116 | 1.9758 | 2.0903 | 2.0322 | 2.82 | 0.79 | 2.3 | 0.04 | 4.53 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 154 | 1.4946 | 1.4949 | 1.397 | 3.86 | 1.02 | 2.3 | 0.04 | 4.53 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 162 | 2.3832 | 2.2447 | 2.2301 | 3.7 | 1.05 | 2.3 | 0.04 | 4.53 | 0 |
| CBDA (%DW) | 504.10 | Other LC-UV | 142 | 1.7652 | 1.9097 | 1.7697 | 4.53 | 1.24 | 2.3 | 0.04 | 4.53 | 0 |
|  |  |  |  | CBN (%DW) |  |  |  |  |  |  |  |  |
| CBN (%DW) | 505.10 | Other LC-UV | 189 | 0.0876 | 0.0872 | 0.0872 | 0.26 | 0.05 | 3.37 | 0.26 | 9.32 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 162 | 0.0878 | 0.0884 | 0.0863 | 1.24 | 0.21 | 3.37 | 0.26 | 9.32 | 0 |
| CBN (%DW) | 505.02 | AOAC 2018.11, U | 199 | 0.0853 | 0.0858 | 0.0836 | 1.36 | 0.24 | 3.37 | 0.26 | 9.32 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 142 | 0.0722 | 0.0741 | 0.0721 | 1.55 | 0.26 | 3.37 | 0.26 | 9.32 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 116 | 0.1008 | 0.0971 | 0.0951 | 2.96 | 0.52 | 3.37 | 0.26 | 9.32 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 154 | 0.0705 | 0.0768 | 0.0726 | 4.38 | 0.74 | 3.37 | 0.26 | 9.32 | 0 |
| CBN (%DW) | 505.10 | Other LC-UV | 192 | 0.1214 | 0.1058 | 0.1151 | 6.88 | 1.24 | 3.37 | 0.26 | 9.32 | 0 |
| CBN (%DW) | 505.30 | GC-FID | 116 | 0.0864 | 0.0722 | 0.0836 | 9.32 | 1.6 | 3.37 | 0.26 | 9.32 | 0 |
|  |  |  |  | Total Δ9-THC (%DW) |  |  |  |  |  |  |  |  |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 189 | 0.2064 | 0.2077 | 0.2078 | 0.38 | 0.07 | 3.16 | 0.38 | 14.4 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 116 | 0.1893 | 0.1913 | 0.1894 | 0.59 | 0.12 | 3.16 | 0.38 | 14.4 | 0 |
| Total Δ9-THC (%DW) | 506.01 | AOAC 2018.10 | 210 | 2.442 | 2.414 | 2.449 | 0.76 | 0.22 | 3.16 | 0.38 | 14.4 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 154 | 0.148 | 0.1515 | 0.1535 | 1.84 | 0.35 | 3.16 | 0.38 | 14.4 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 162 | 0.1808 | 0.1805 | 0.1864 | 1.82 | 0.35 | 3.16 | 0.38 | 14.4 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 142 | 0.1675 | 0.1733 | 0.1662 | 2.24 | 0.43 | 3.16 | 0.38 | 14.4 | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 116 | 0.1849 | 0.1748 | 0.1773 | 2.94 | 0.57 | 3.16 | 0.38 | 14.4 | 0 |
| Total Δ9-THC (%DW) | 506.02 | AOAC 2018.11, U | 199 | 0.1943 | 0.2063 | 0.2019 | 3.02 | 0.6 | 3.16 | 0.38 | 14.4 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 165 | 0.1814 | 0.1702 | 0.1873 | 4.84 | 0.94 | 3.16 | 0.38 | 14.4 | 0 |
| Total Δ9-THC (%DW) | 506.10 | Other LC-UV | 192 | 0.2578 | 0.2437 | 0.269 | 4.94 | 1.01 | 3.16 | 0.38 | 14.4 | 0 |
| Total Δ9-THC (%DW) | 506.30 | Other LC-MS | 102 | 0.1952 | 0.169 | 0.174 | 7.75 | 1.5 | 3.16 | 0.38 | 14.4 | 0 |
| Total Δ9-THC (%DW) | 506.40 | GC-FID | 153 | 0.2129 | 0.161 | 0.1781 | 14.4 | 2.79 | 3.16 | 0.38 | 14.4 | 0 |
|  |  |  |  | Total CBD (%DW) |  |  |  |  |  |  |  |  |
| Total CBD (%DW) | 507.10 | Other LC-UV | 189 | 7.8532 | 7.8569 | 7.859 | 0.04 | 0.01 | 2.46 | 0.04 | 6.87 | 0 |
| Total CBD (%DW) | 507.01 | AOAC 2018.10 | 210 | 8.669 | 8.636 | 8.597 | 0.42 | 0.14 | 2.46 | 0.04 | 6.87 | 0 |
| Total CBD (%DW) | 507.02 | AOAC 2018.11, U | 199 | 8.285 | 8.2216 | 8.1288 | 0.96 | 0.33 | 2.46 | 0.04 | 6.87 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 162 | 9.0937 | 8.7206 | 8.7574 | 2.32 | 0.81 | 2.46 | 0.04 | 6.87 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 154 | 6.4472 | 6.4975 | 6.1775 | 2.7 | 0.89 | 2.46 | 0.04 | 6.87 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 142 | 7.089 | 7.4539 | 7.1196 | 2.8 | 0.94 | 2.46 | 0.04 | 6.87 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 116 | 7.7723 | 8.3583 | 8.0502 | 3.64 | 1.25 | 2.46 | 0.04 | 6.87 | 0 |
| Total CBD (%DW) | 507.40 | GC-FID | 116 | 8.6045 | 8.2763 | 8.9388 | 3.85 | 1.33 | 2.46 | 0.04 | 6.87 | 0 |
| Total CBD (%DW) | 507.10 | Other LC-UV | 192 | 10.0374 | 8.746 | 9.4603 | 6.87 | 2.41 | 2.46 | 0.04 | 6.87 | 0 |
|  |  |  |  | CBDV (%DW) |  |  |  |  |  |  |  |  |
| CBDV (%DW) | 508.99 | Other | 189 | 0.0337 | 0.0342 | 0.0341 | 0.78 | 0.12 |  |  |  | 0 |
| CBDV (%DW) | 508.99 | Other | 162 | 0.0509 | 0.053 | 0.0504 | 2.68 | 0.43 |  |  |  | 0 |
| CBDV (%DW) | 508.02 | AOAC 2018.11, U | 210 | 0.062 | 0.058 | 0.059 | 3.49 | 0.57 |  |  |  | 0 |

Population of Lab RSDr

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab RSDr | HorRat(r) | Rob Mean | min | max | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDV (%DW) | 508.99 | Other | 142 | 0.0406 | 0.0449 | 0.0429 | 5.03 | 0.79 |  |  |  | 0 |
|  |  |  |  | CBDVA (%DW) |  |  |  |  |  |  |  |  |
| CBDVA (%DW) | 509.99 | Other | 142 | 0.0188 | 0.0201 | 0.0194 | 3.35 | 0.46 |  |  |  | 0 |
| CBDVA (%DW) | 509.01 | AOAC 2018.10 | 210 | 0.033 | 0.031 | 0.031 | 3.65 | 0.54 |  |  |  | 0 |
|  |  |  |  | CBG (%DW) |  |  |  |  |  |  |  |  |
| CBG (%DW) | 510.99 | Other | 189 | 0.3227 | 0.3192 | 0.3203 | 0.56 | 0.12 |  |  |  | 0 |
| CBG (%DW) | 510.01 | AOAC 2018.10 | 210 | 0.239 | 0.244 | 0.235 | 1.88 | 0.38 |  |  |  | 0 |
| CBG (%DW) | 510.99 | Other | 162 | 0.2148 | 0.2296 | 0.225 | 3.39 | 0.68 |  |  |  | 0 |
| CBG (%DW) | 510.99 | Other | 142 | 0.2077 | 0.1965 | 0.2107 | 3.65 | 0.72 |  |  |  | 0 |
|  |  |  |  | CBGA (%DW) |  |  |  |  |  |  |  |  |
| CBGA (%DW) | 511.99 | Other | 189 | 0.1664 | 0.1665 | 0.1668 | 0.12 | 0.02 |  |  |  | 0 |
| CBGA (%DW) | 511.99 | Other | 162 | 0.161 | 0.1577 | 0.156 | 1.61 | 0.31 |  |  |  | 0 |
| CBGA (%DW) | 511.99 | Other | 142 | 0.1103 | 0.1165 | 0.1163 | 3.08 | 0.56 |  |  |  | 0 |
| CBGA (%DW) | 511.01 | AOAC 2018.10 | 210 | 0.284 | 0.278 | 0.343 | 11.9 | 2.49 |  |  |  | 0 |
|  |  |  |  | THCV (%DW) |  |  |  |  |  |  |  |  |
| THCV (%DW) | 512.02 | AOAC 2018.11, U | 210 | 0.013 | 0.013 | 0.013 | 0 | 0 |  |  |  | 0 |

Population of Lab RSDr

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

RSDr	HorRat(r)

Population of Lab RSDr Rob Mean	min	max

Flag

Δ8-THC (%DW)

| Δ8-THC (%DW) | 513.01 | AOAC 2018.10 | 210 | 0.013 | 0.013 | 0.013 | 0 | 0 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%DW) | 513.99 | Other | 189 | 0.0138 | 0.0143 | 0.0128 | 5.6 | 0.74 | 0 |

CBC (%DW)

| CBC (%DW) | 514.99 | Other | 189 | 0.2938 | 0.294 | 0.2939 | 0.03 | 0.01 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%DW) | 514.01 | AOAC 2018.10 | 210 | 0.256 | 0.259 | 0.259 | 0.67 | 0.14 | 0 |
| CBC (%DW) | 514.99 | Other | 162 | 0.2653 | 0.265 | 0.2573 | 1.73 | 0.35 | 0 |
| CBC (%DW) | 514.99 | Other | 142 | 0.2275 | 0.2377 | 0.2281 | 2.48 | 0.5 | 0 |

- Flag definitions: 1 (Lab RSDr not calculated because there were less than 3 numeric lab results)

- Statistical parameters of robust mean, minimum, and maximum for RSDr only shown only if there were 6 or more lab RSDr values.

- AR refers to concentration on as-received basis. DW refers to concentration on dry weight basis.c