| Sample # | Replicate 1 | Replicate 2 |
| --- | --- | --- |
| 1 | 1.8536 | 1.8737 |
| 2 | 1.7189 | 1.7590 |
| 3 | 1.8496 | 1.8172 |
| 4 | 1.6875 | 1.7733 |
| 5 | 1.7800 | 1.7748 |
| 6 | 1.8734 | 1.8202 |
| 7 | 1.8420 | 1.8233 |
| 8 | 1.8843 | 1.9380 |
| 9 | 1.8698 | 1.8664 |
| 10 | 1.9050 | 1.8827 |

Homogeneity and Stability

Sample ID . . . . . . . . . . . . . . . . . . . . . . . .

Samples shipped . . . . . . . . . . . . . . . . . . .

Results due date . . . . . . . . . . . . . . . . . . . .

OL21NOV-1

8-Nov-21

15-Dec-21

## Homogeneity

10 sample packets randomly chosen for analysis.

Duplicate test portions analyzed for analyte and method shown in table caption.

Total delta-9 THC (%AR) via GC-FID (method 006.40) 	%RSD	HorRat-r

| Sample # | Replicate 1 | Replicate 2 |
| --- | --- | --- |
| 1 | 0.0775 | 0.0787 |
| 2 | 0.0737 | 0.0754 |
| 3 | 0.0803 | 0.0769 |
| 4 | 0.0723 | 0.0744 |
| 5 | 0.0746 | 0.0757 |
| 6 | 0.0800 | 0.0785 |
| 7 | 0.0779 | 0.0786 |
| 8 | 0.0798 | 0.0822 |
| 9 | 0.0790 | 0.0776 |
| 10 | 0.0812 | 0.0805 |

Analysis Date: 12/6/21	Overall average:	0.0778

SD of sample avgs:	0.002575	3.31

repeatability SD:	0.001278	1.64	0.3

between-sample SD:	0.00241	3.10

reproducibilty SD:	0.002729	3.51

SD used for Analyte z scores:	0.0084

Check value:	0.00252

Is between-sample SD less than check value? YES Homogeneity test passed

| Sample # | Replicate 1 | Replicate 2 |
| --- | --- | --- |
| 1 | 0.2764 | 0.2818 |
| 2 | 0.2547 | 0.2606 |
| 3 | 0.2722 | 0.2714 |
| 4 | 0.2555 | 0.2643 |
| 5 | 0.2639 | 0.2635 |
| 6 | 0.2725 | 0.2666 |
| 7 | 0.2681 | 0.2676 |
| 8 | 0.2752 | 0.2815 |
| 9 | 0.2719 | 0.2687 |
| 10 | 0.2766 | 0.2728 |

Total delta-9 THC (%AR) via GC-FID (method 006.40)

| delta-8 THC (%AR) via GC-FID (method 0013.99) Analysis Date: 12/6/21 | Overall average: | 0.2693 | %RSD |  |
| --- | --- | --- | --- | --- |
|  | SD of sample avgs: | 0.007249 | 2.69 |  |
|  | repeatability SD: | 0.003469 | 1.29 | 0.3 |
|  | between-sample SD: | 0.00682 | 2.53 |  |
|  | reproducibilty SD: | 0.007653 | 2.84 |  |
|  | SD used for Analyte z scores: | 0.0177 |  |  |
|  | Check value: | 0.00531 |  |  |
| delta-8 THC (%AR) via GC-FID (method 0013.99) Is between-sample SD less than check value? NO Homogeneitty test did not pass |  |  |  |  |
| Total CBD (%AR) via GC-FID (method 007.40) Analysis Date: 12/6/21 | Overall average: | 1.8296 | %RSD |  |
|  | SD of sample avgs: | 0.06194 | 3.39 |  |
|  | repeatability SD: | 0.02917 | 1.59 | 0.4 |
|  | between-sample SD: | 0.0584 | 3.19 |  |
|  | reproducibilty SD: | 0.06528 | 3.57 |  |
|  | SD used for Analyte z scores: | 0.1042 |  |  |
|  | Check value: | 0.03126 |  |  |
| Total CBD (%AR) via GC-FID (method 007.40) Is between-sample SD less than check value? NO Homogeneitty test did not pass |  |  |  |  |

## Stability

Sample tested on different days from date sample shipped to results due date. Sample stored at room temperature during period of testing.

| delta-8 THC (%AR) via GC-FID (method 006.40) |  |  |
| --- | --- | --- |
| Date of Analysis | Days | Conc. |
| 11/10/21 | 3 | 0.2815 |
| 11/17/21 | 10 | 0.2777 |
| 11/23/21 | 16 | 0.2662 |
| 12/5/21 | 28 | 0.2597 |
| 12/6/21 | 29 | 0.3024 |
| 12/12/21 | 35 | 0.2991 |
| 12/14/21 | 37 | 0.3009 |
| 12/19/21 | 42 | 0.2872 |

Days = x variable, Conc = y variable SUMMARY OUTPUT

Regression Statistics

Multiple R  0.4594417 R Square	0.2110867 Adjusted R Squ 0.0796011 Standard Error  0.0154472

Observations	8

0.34

0.32

0.30

0.28

0.26

0.24

0.22

y = 0.0005x + 0.271

R² = 0.2111

0.20

0

10

20

30

40

ANOVA

df	SS	MS	F	Significance F

| Regression | 1 | 0.00038308 0.000383076 | 1.605398 | 0.252097467 |  |
| --- | --- | --- | --- | --- | --- |
| Residual | 6 | 0.0014317 0.000238617 |  |  |  |
| Total | 7 | 0.00181478 |  |  |  |
| Coefficients tandard Error t Stat P-value Lower 95% |  |  |  |  | Upper 95% |

Intercept	0.2710103 0.01185167 22.86685807  4.58E-07  0.242010368  0.300010329

X Variable 1	0.0005331 0.00042073 1.267043156 0.252097 -0.000496409  0.001562581

Coefficient for x variable is not statistically significantly different from 0. No signficant change in result during testing period

Stability test passed.

%RSD = 5.7

| Total CBD (%AR) via GC-FID (method 006.40) |  |  |
| --- | --- | --- |
| Date of Analysis | Days | Conc. |
| 11/10/21 | 3 | 1.7956 |
| 11/17/21 | 10 | 1.8251 |
| 11/23/21 | 16 | 1.8406 |
| 12/5/21 | 28 | 1.6726 |
| 12/6/21 | 29 | 1.9914 |
| 12/12/21 | 35 | 1.8468 |
| 12/14/21 | 37 | 1.8183 |
| 12/19/21 | 42 | 1.7530 |

Days = x variable, Conc = y variable SUMMARY OUTPUT

Regression Statistics

Multiple R  0.0424428 R Square	0.0018014 Adjusted R Squ -0.164565 Standard Error  0.0976583

Observations	8

ANOVA

y = -0.0003x + 1.8248 R² = 0.0018

2.20

2.10

2.00

1.90

1.80

1.70

1.60

1.50

1.40

0

10

20

30

40

df	SS	MS	F	Significance F

| Regression | 1 | 0.00010327 0.000103267 | 0.010828 | 0.920515279 |  |
| --- | --- | --- | --- | --- | --- |
| Residual | 6 | 0.05722287 0.009537145 |  |  |  |
| Total | 7 | 0.05732614 |  |  |  |
| Coefficients tandard Error t Stat P-value Lower 95% |  |  |  |  | Upper 95% |

Intercept	1.8248445 0.07492689 24.35500146  3.15E-07  1.641505016  2.008184005

X Variable 1	-0.0002768 0.00265989 -0.10405696 0.920515 -0.006785305  0.006231744

Coefficient for x variable is not statistically significantly different from 0. No signficant change in result during testing period

Stability test passed.

%RSD = 5