# Homogeneity and Stability

Sample ID . . . . . . . . . . . . . . . . . . . . . . . .

Samples shipped . . . . . . . . . . . . . . . . . . .

Results due date . . . . . . . . . . . . . . . . . . . .

HM21SEP-1

7-Sep-21

15-Oct-21

## Homogeneity

10 sample packets randomly chosen for analysis.

| Sample # | Replicate 1 | Replicate 2 |
| --- | --- | --- |
| 1 | 0.1036 | 0.1146 |
| 2 | 0.1070 | 0.1026 |
| 3 | 0.1063 | 0.1152 |
| 4 | 0.1126 | 0.1062 |
| 5 | 0.1033 | 0.1089 |
| 6 | 0.1060 | 0.1096 |
| 7 | 0.1058 | 0.1066 |
| 8 | 0.1069 | 0.0993 |
| 9 | 0.0996 | 0.1079 |
| 10 | 0.1091 | 0.1047 |

Duplicate test portions analyzed for analyte and method shown in table caption.

| Total delta-9 THC (%AR) via GC-FID (method 006.40) Analysis Date: 10/4/21 | Overall average: | 0.1068 | %RSD | HorRat-r |
| --- | --- | --- | --- | --- |
|  | SD of sample avgs: | 0.002496 | 2.34 |  |
|  | repeatability SD: | 0.004747 | 4.45 | 0.8 |
|  | between-sample SD: | 0 | 0.00 |  |
|  | reproducibilty SD: | 0.004747 | 4.45 |  |
|  | SD used for Analyte z scores: | 0.008 |  |  |
|  | Check value: | 0.0024 |  |  |
| Total delta-9 THC (%AR) via GC-FID (method 006.40) Is between-sample SD less than check value? YES Homogeneity test passed |  |  |  |  |
| delta-8 THC (%AR) via GC-FID (method 0013.99) Analysis Date: 10/4/21 | Overall average: | 0.3897 | %RSD |  |
|  | SD of sample avgs: | 0.003166 | 0.81 |  |
|  | repeatability SD: | 0.004996 | 1.28 | 0.3 |
|  | between-sample SD: | 0 | 0.00 |  |
|  | reproducibilty SD: | 0.004996 | 1.28 |  |
|  | SD used for Analyte z scores: | 0.0205 |  |  |
|  | Check value: | 0.00615 |  |  |
| delta-8 THC (%AR) via GC-FID (method 0013.99) Is between-sample SD less than check value? YES Homogeneity test passed |  |  |  |  |

| Sample # | Replicate 1 | Replicate 2 |
| --- | --- | --- |
| 1 | 0.4036 | 0.3864 |
| 2 | 0.3921 | 0.3915 |
| 3 | 0.3902 | 0.3879 |
| 4 | 0.3921 | 0.3886 |
| 5 | 0.3863 | 0.3910 |
| 6 | 0.3814 | 0.3844 |
| 7 | 0.3924 | 0.3863 |
| 8 | 0.3890 | 0.3915 |
| 9 | 0.3883 | 0.3867 |
| 10 | 0.3865 | 0.3970 |

Stability

Sample tested on different days from date sample shipped to results due date. Sample stored at room temperature during period of testing.

| Date of Analysis | Days | Conc. |
| --- | --- | --- |
| 9/13/21 | 0 | 0.0906 |
| 9/15/21 | 2 | 0.0850 |
| 9/18/21 | 5 | 0.1206 |
| 9/26/21 | 13 | 0.0812 |
| 9/29/21 | 16 | 0.1071 |
| 10/2/21 | 19 | 0.0861 |
| 10/5/21 | 22 | 0.1262 |
| 10/9/21 | 26 | 0.0912 |
| 10/13/21 | 30 | 0.0954 |

Total THC (%AR) via GC-FID (method 006.40)		Days = x variable, Conc = y variable

SUMMARY OUTPUT

Regression Statistics

Multiple R  0.0964902 R Square	0.0093104 Adjusted R Squ -0.1322167 Standard Error  0.0171953

Observations	9

ANOVA

df	SS	MS	F	Significance F

y = 0.0001x + 0.096 R² = 0.0093

0.14

0.12

0.10

0.08

0.06

0.04

0.02

0.00

0

5

10	15	20	25	30

Regression	1  1.9451E-05 1.94512E-05 0.065785  0.804950559

Residual	7 0.00206975 0.000295679

Total	8	0.0020892

Coefficients tandard Error	t Stat	P-value	Lower 95%	Upper 95%

Intercept	0.0959945 0.01019044 9.420050306  3.17E-05  0.071897921	0.12009106

X Variable 1	0.0001462 0.00057016 0.256485769 0.804951 -0.001201973  0.001494448

Coefficient for x variable is not statistically significantly different from 0. No signficant change in result during testing period

Stability test passed.

%RSD = 16.5

| Total CBD (%AR) via GC-FID (method 006.40) |  |  |
| --- | --- | --- |
| Date of Analysis | Days | Conc. |
| 9/13/21 | 0 | 1.6570 |
| 9/15/21 | 2 | 1.7407 |
| 9/18/21 | 5 | 1.7474 |
| 9/22/21 | 9 | 1.6966 |
| 9/26/21 | 13 | 1.6280 |
| 9/29/21 | 16 | 1.7155 |
| 10/2/21 | 19 | 1.6966 |
| 10/5/21 | 22 | 1.6558 |
| 10/9/21 | 26 | 1.7023 |
| 10/13/21 | 30 | 1.6428 |

Days = x variable, Conc = y variable SUMMARY OUTPUT

Regression Statistics

Multiple R  0.3796164 R Square	0.1441086 Adjusted R Squ 0.0371222 Standard Error  0.0401326

Observations	10

ANOVA

df	SS	MS	F	Significance F

1.80

1.75

1.70

1.65

1.60

y = -0.0015x + 1.7098

R² = 0.1441

1.55

1.50

0

5

10	15	20	25	30

Regression	1 0.00216948 0.002169478 1.346981	0.27926622

Residual	8 0.01288498 0.001610623

Total	9 0.01505446

Coefficients tandard Error	t Stat	P-value	Lower 95%	Upper 95%

Intercept	1.7098472 0.02251011 75.95907213  1.01E-12  1.657938759  1.761755579

X Variable 1	-0.0015195 0.00130926 -1.16059499 0.279266 -0.004538675  0.001499637

Coefficient for x variable is not statistically significantly different from 0. No signficant change in result during testing period

Stability test passed.

%RSD = 2.4

| delta-8 THC (%AR) via GC-FID (method 006.40) |  |  |
| --- | --- | --- |
| Date of Analysis | Days | Conc. |
| 9/13/21 | 0 | 0.3499 |
| 9/15/21 | 2 | 0.3357 |
| 9/18/21 | 5 | 0.4131 |
| 9/22/21 | 9 | 0.3681 |
| 9/26/21 | 13 | 0.3478 |
| 9/29/21 | 16 | 0.4088 |
| 10/2/21 | 19 | 0.3463 |
| 10/5/21 | 22 | 0.3771 |
| 10/9/21 | 26 | 0.3568 |
| 10/13/21 | 30 | 0.3531 |

Days = x variable, Conc = y variable SUMMARY OUTPUT

Regression Statistics

Multiple R  0.0276327 R Square	0.0007636 Adjusted R Squ -0.124141 Standard Error  0.0281066

Observations	10

ANOVA

df	SS	MS	F	Significance F

Regression	1  4.8293E-06 4.82933E-06 0.006113  0.939599528

y = -7E-05x + 0.3667 R² = 0.0008

0.43

0.41

0.39

0.37

0.35

0.33

0.31

0.29

0.27

0.25

0

5

10	15	20	25	30

Residual	8 0.00631986 0.000789983

Total	9 0.00632469

Coefficients tandard Error	t Stat	P-value	Lower 95%	Upper 95%

Intercept	0.366689 0.01576484 23.25993294  1.24E-08  0.330335252  0.403042806

X Variable 1	-7.169E-05 0.00091693 -0.07818703	0.9396 -0.002186142  0.002042757

Coefficient for x variable is not statistically significantly different from 0. No signficant change in result during testing period

Stability test passed.

%RSD = 7.2