Hemp Proficiency Testing

All Labs Trueness OL22SEP-1 (oil)

Sample: OL22SEP-1 (oil)

All Labs Trueness

Issue Date:

ovember 18, 2022

| Analyte | Code | Method | Lab Num |  | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Δ9-THC (%AR) |  |  |  |  |  |  |  |  |  |  |
| Δ9-THC (%AR) | 001.60 | NIR |  | 187 | 0.02 | 0.03 | 0.03 | 0.0267 | -2.95 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 184 | 0.0554 | 0.0548 | 0.056 | 0.0554 | -1.68 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.99 | Other |  | 183 | 0.0578 | 0.0636 | 0.0717 | 0.0644 | -1.28 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 204 | 0.0673 | 0.07 | 0.0698 | 0.069 | -1.07 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.03 | AOAC 2018.11, M |  | 139 | 0.0818 | 0.0621 | 0.07 | 0.0713 | -0.97 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 191 | 0.0807 | 0.0689 | 0.0805 | 0.0767 | -0.73 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 197 | 0.0849 | 0.081 | 0.0853 | 0.0837 | -0.42 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 192 | 0.0944 | 0.088 | 0.0748 | 0.0857 | -0.33 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 216 | 0.0855 | 0.0871 | 0.0868 | 0.0865 | -0.3 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS |  | 102 | 0.1032 | 0.0694 | 0.0957 | 0.0894 | -0.17 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 |  | 183 | 0.0886 | 0.09 | 0.0895 | 0.0894 | -0.17 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 200 | 0.089 | 0.089 | 0.089 | 0.089 | -0.19 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U |  | 125 | 0.0913 | 0.0883 | 0.0911 | 0.0902 | -0.13 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS |  | 149 | 0.097 | 0.085 | 0.092 | 0.0913 | -0.08 | 0.0932 | 26 | 0.0225 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 194 | 0.0947 | 0.0952 | 0.0828 | 0.0909 | -0.1 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 106 | 0.09 | 0.1 | 0.09 | 0.0933 | 0 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 149 | 0.096 | 0.096 | 0.095 | 0.0957 | 0.11 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 196 | 0.1004 | 0.1024 | 0.1009 | 0.101 | 0.35 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 190 | 0.101 |  |  | 0.101 |  |  |  |  | 1 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 189 | 0.1041 | 0.1081 | 0.109 | 0.107 | 0.61 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 212 | 0.1057 | 0.1145 | 0.1029 | 0.108 | 0.64 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 155 | 0.111 | 0.121 | 0.105 | 0.112 | 0.85 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 120 | 0.1118 | 0.1152 | 0.1193 | 0.115 | 0.98 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 162 | 0.1131 | 0.1143 | 0.1187 | 0.115 | 0.98 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 199 | 0.1265 | 0.1175 | 0.133 | 0.126 | 1.44 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 | 170 | 0.1243 | 0.1135 | 0.1437 | 0.127 | 1.51 | 0.0932 | 26 | 0.0225 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 122 | 0.1274 | 0.1275 | 0.1269 | 0.127 | 1.51 | 0.0932 | 26 | 0.0225 | 0 |

Population of Lab Values

| Δ9-THCA (%AR) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 190 | <0.0010 |  |  | <0.0010 | 1 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 120 | <0.0025 | <0.0025 | <0.0025 | <0.0025 | 1 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 149 | <0.0040 | <0.0040 | <0.0040 | <0.0040 | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 197 | <0.0094 | <0.0098 | <0.0092 | <0.0094 | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 216 | <0.0099 | <0.0091 | <0.0103 | <0.0099 | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 106 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |

Lab

Lab

Population of Lab Values

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 191 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.03 | AOAC 2018.11, M | 139 | <0.0125 | <0.0125 | <0.0125 | <0.0125 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 102 | <0.0194 | <0.0188 | <0.0187 | <0.0194 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 196 | <0.0210 | <0.0210 | <0.0210 | <0.0210 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 194 | <0.0270 | <0.0270 | <0.0270 | <0.0270 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 149 | <0.0370 | <0.0370 | <0.0400 | <0.0370 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 199 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 122 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 184 | <0.1000 | <0.1000 | <0.1000 | <0.1000 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 170 | <0.1300 | <0.1300 | <0.1300 | <0.1300 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 183 | 0 | 0 | 0 | 0 |  |  |  |  | 2 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 125 | 0.002 | 0.0021 | 0.0025 | 0.0022 | -1.05 | 0.0309 | 6 | 0.0274 | 0 |
| Δ9-THCA (%AR) | 002.60 | NIR | 187 | 0.01 | 0.02 | 0.01 | 0.0133 | -0.64 | 0.0309 | 6 | 0.0274 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 200 | 0.022 | 0.022 | 0.022 | 0.022 | -0.32 | 0.0309 | 6 | 0.0274 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 189 | 0.0341 | 0.0377 | 0.0387 | 0.0368 | 0.22 | 0.0309 | 6 | 0.0274 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 155 | 0.05 | 0.03 | 0 | 0.04 | 0.33 | 0.0309 | 6 | 0.0274 | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 192 | 0.0729 | 0.0647 | 0.0748 | 0.0708 | 1.46 | 0.0309 | 6 | 0.0274 | 0 |

| CBD (%AR) |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%AR) | 003.10 | Other LC-UV | 191 | 4.6798 | 3.1997 | 3.3241 | 3.73 | -3.67 | 5.7535 | 24 | 0.5499 | 0 |
| CBD (%AR) | 003.03 | AOAC 2018.11, M | 139 | 5.1067 | 4.0449 | 4.3818 | 4.51 | -2.26 | 5.7535 | 24 | 0.5499 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 200 | 5.128 | 4.946 | 4.977 | 5.02 | -1.34 | 5.7535 | 24 | 0.5499 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%AR) | 003.10 | Other LC-UV | 189 | 5.0957 | 5.1238 | 5.0945 | 5.1 | -1.18 | 5.7535 | 24 | 0.5499 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 192 | 5.4912 | 5.2414 | 4.7246 | 5.15 | -1.09 | 5.7535 | 24 | 0.5499 | 0 |
| CBD (%AR) | 003.99 | Other | 212 | 5.4674 | 5.4764 | 5.4158 | 5.45 | -0.55 | 5.7535 | 24 | 0.5499 | 0 |
| CBD (%AR) | 003.01 | AOAC 2018.10 | 183 | 5.5916 | 5.5525 | 5.5681 | 5.57 | -0.33 | 5.7535 | 24 | 0.5499 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 194 | 5.5474 | 5.5454 | 5.6459 | 5.58 | -0.32 | 5.7535 | 24 | 0.5499 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 184 | 5.736 | 5.481 | 5.631 | 5.62 | -0.25 | 5.7535 | 24 | 0.5499 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 149 | 5.655 | 5.521 | 5.816 | 5.66 | -0.16 | 5.7535 | 24 | 0.5499 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 155 | 5.67 | 5.74 | 5.71 | 5.71 | -0.09 | 5.7535 | 24 | 0.5499 | 0 |
| CBD (%AR) | 003.01 | AOAC 2018.10 | 170 | 5.6788 | 5.7715 | 5.6892 | 5.71 | -0.07 | 5.7535 | 24 | 0.5499 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 125 | 5.83 | 5.63 | 5.69 | 5.72 | -0.07 | 5.7535 | 24 | 0.5499 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 190 | 5.916 |  |  | 5.92 |  |  |  |  | 1 |
| CBD (%AR) | 003.10 | Other LC-UV | 204 | 5.9638 | 5.9321 | 5.9633 | 5.95 | 0.36 | 5.7535 | 24 | 0.5499 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 196 | 6.0616 | 6.0048 | 5.7996 | 5.96 | 0.37 | 5.7535 | 24 | 0.5499 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 106 | 6.04 | 5.96 | 6.01 | 6 | 0.45 | 5.7535 | 24 | 0.5499 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 216 | 5.8788 | 6.1045 | 6.0279 | 6 | 0.45 | 5.7535 | 24 | 0.5499 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 102 | 5.9804 | 6.2777 | 5.7592 | 6.01 | 0.46 | 5.7535 | 24 | 0.5499 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 199 | 5.954 | 5.959 | 6.135 | 6.02 | 0.48 | 5.7535 | 24 | 0.5499 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 162 | 6.1554 | 6.3674 | 6.1204 | 6.21 | 0.84 | 5.7535 | 24 | 0.5499 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 197 | 6.23 | 6.2888 | 6.3578 | 6.29 | 0.98 | 5.7535 | 24 | 0.5499 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 149 | 6.427 | 6.31 | 6.262 | 6.33 | 1.05 | 5.7535 | 24 | 0.5499 | 0 |
| CBD (%AR) | 003.60 | NIR | 187 | 6.62 | 6.73 | 6.5 | 6.62 | 1.57 | 5.7535 | 24 | 0.5499 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 122 | 6.8498 | 6.8495 | 6.8501 | 6.85 | 1.99 | 5.7535 | 24 | 0.5499 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num |  | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | CBDA (%AR) |  |  |  |  |  |  |  |  |  |  |
| CBDA (%AR) | 004.30 | Other LC-MS |  | 149 | <0.0040 | <0.0040 | <0.0040 | <0.0040 |  |  |  |  | 1 |
| CBDA (%AR) | 004.10 | Other LC-UV |  | 197 | <0.0094 | <0.0098 | <0.0092 | <0.0094 |  |  |  |  | 1 |
| CBDA (%AR) | 004.10 | Other LC-UV |  | 216 | <0.0099 | <0.0091 | <0.0103 | <0.0099 |  |  |  |  | 1 |
| CBDA (%AR) | 004.10 | Other LC-UV |  | 106 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| CBDA (%AR) | 004.10 | Other LC-UV |  | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| CBDA (%AR) | 004.03 | AOAC 2018.11, M |  | 139 | <0.0125 | <0.0125 | <0.0125 | <0.0125 |  |  |  |  | 1 |
| CBDA (%AR) | 004.10 | Other LC-UV |  | 194 | <0.0270 | <0.0270 | <0.0270 | <0.0270 |  |  |  |  | 1 |
| CBDA (%AR) | 004.10 | Other LC-UV |  | 149 | <0.0370 | <0.0370 | <0.0400 | <0.0370 |  |  |  |  | 1 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U |  | 199 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| CBDA (%AR) | 004.30 | Other LC-MS |  | 122 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| CBDA (%AR) | 004.01 | AOAC 2018.10 |  | 170 | <0.1300 | <0.1300 | <0.1300 | <0.1300 |  |  |  |  | 1 |
| CBDA (%AR) | 004.01 | AOAC 2018.10 |  | 183 | 0 | 0 | 0 | 0 |  |  |  |  | 2 |
| CBDA (%AR) | 004.10 | Other LC-UV |  | 200 | 0.004 | 0.004 | 0.004 | 0.004 | -1.05 | 0.0316 | 11 | 0.0263 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U |  | 125 | 0.005 | 0.0055 | 0.0062 | 0.0056 | -0.99 | 0.0316 | 11 | 0.0263 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV |  | 189 | 0.0044 | 0.0098 | 0.0106 | 0.0083 | -0.89 | 0.0316 | 11 | 0.0263 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV |  | 192 | 0.0236 | 0.0252 | 0.0218 | 0.0235 | -0.31 | 0.0316 | 11 | 0.0263 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV |  | 191 | 0.0251 | 0.0232 | 0.0267 | 0.025 | -0.25 | 0.0316 | 11 | 0.0263 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV |  | 212 | 0.0266 | 0.0294 | 0 | 0.028 | -0.14 | 0.0316 | 11 | 0.0263 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV |  | 196 | 0.0325 | 0.032 | 0.0322 | 0.0322 | 0.02 | 0.0316 | 11 | 0.0263 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV |  | 204 | 0.0351 | 0.0328 | 0.0405 | 0.0361 | 0.17 | 0.0316 | 11 | 0.0263 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV |  | 184 | 0.0438 | 0.0426 | 0.0431 | 0.0432 | 0.44 | 0.0316 | 11 | 0.0263 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDA (%AR) | 004.10 | Other LC-UV | 155 | 0.079 | 0.075 | 0.077 | 0.077 | 1.73 | 0.0316 | 11 | 0.0263 | 0 |
| CBDA (%AR) | 004.60 | NIR | 187 | 0.15 | 0.19 | 0.23 | 0.19 | 6.02 | 0.0316 | 11 | 0.0263 | 0 |

Population of Lab Values

| CBN (%AR) |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%AR) | 005.20 | Other LC-MS | 102 | <0.0484 | <0.0470 | <0.0467 | <0.0484 |  |  |  |  | 1 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 199 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| CBN (%AR) | 005.03 | AOAC 2018.11, M | 139 | 0.0338 | 0.0304 | 0.0275 | 0.0306 | -1.58 | 0.0424 | 17 | 0.0074 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 107 | 0.032 | 0.029 | 0.045 | 0.0353 | -0.94 | 0.0424 | 17 | 0.0074 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 194 | 0.0368 | 0.0368 | 0.0309 | 0.0348 | -1.01 | 0.0424 | 17 | 0.0074 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 197 | 0.0362 | 0.0363 | 0.0371 | 0.0365 | -0.78 | 0.0424 | 17 | 0.0074 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 216 | 0.0372 | 0.0382 | 0.0384 | 0.0379 | -0.6 | 0.0424 | 17 | 0.0074 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 149 | 0.039 | <0.0370 | <0.0400 | 0.0387 |  |  |  |  | 1 |
| CBN (%AR) | 005.01 | AOAC 2018.10 | 183 | 0.0397 | 0.039 | 0.0396 | 0.0394 | -0.39 | 0.0424 | 17 | 0.0074 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 200 | 0.04 | 0.04 | 0.04 | 0.04 | -0.32 | 0.0424 | 17 | 0.0074 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 189 | 0.0404 | 0.0411 | 0.0412 | 0.0409 | -0.2 | 0.0424 | 17 | 0.0074 | 0 |
| CBN (%AR) | 005.20 | Other LC-MS | 149 | 0.044 | 0.04 | 0.043 | 0.0423 | 0 | 0.0424 | 17 | 0.0074 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 204 | 0.0402 | 0.0435 | 0.0427 | 0.0421 | -0.03 | 0.0424 | 17 | 0.0074 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 196 | 0.0427 | 0.043 | 0.0423 | 0.0427 | 0.04 | 0.0424 | 17 | 0.0074 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 125 | 0.0444 | 0.0418 | 0.047 | 0.0444 | 0.27 | 0.0424 | 17 | 0.0074 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 162 | 0.0442 | 0.0462 | 0.0448 | 0.0451 | 0.36 | 0.0424 | 17 | 0.0074 | 0 |
| CBN (%AR) | 005.20 | Other LC-MS | 190 | 0.048 |  |  | 0.048 |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 212 | 0.0501 | 0.0472 | 0.0484 | 0.0486 | 0.83 | 0.0424 | 17 | 0.0074 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%AR) | 005.10 | Other LC-UV | 192 | 0.0589 | 0.0538 | 0.0435 | 0.0521 | 1.3 | 0.0424 | 17 | 0.0074 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 106 | 0.05 | 0.06 | 0.05 | 0.0533 | 1.47 | 0.0424 | 17 | 0.0074 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 138 | 0.0554 | 0.0526 | 0.0618 | 0.0566 | 1.91 | 0.0424 | 17 | 0.0074 | 0 |
|  |  |  |  | Total Δ9-THC (%AR) |  |  |  |  |  |  |  |  |
| Total Δ9-THC (%AR) | 006.60 | NIR | 187 | 0.029 | 0.048 | 0.039 | 0.0387 | -2.38 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.99 | Other | 183 | 0.0578 | 0.0636 | 0.0717 | 0.0644 | -1.37 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 204 | 0.0673 | 0.07 | 0.0698 | 0.069 | -1.19 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.03 | AOAC 2018.11, M | 139 | 0.0818 | 0.0621 | 0.07 | 0.0713 | -1.1 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 191 | 0.0807 | 0.0689 | 0.0805 | 0.0767 | -0.89 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 197 | 0.0849 | 0.081 | 0.0853 | 0.0837 | -0.61 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 216 | 0.0855 | 0.0871 | 0.0868 | 0.0865 | -0.5 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 107 | 0.078 | 0.083 | 0.102 | 0.0877 | -0.46 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 183 | 0.0886 | 0.09 | 0.0895 | 0.0894 | -0.39 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 102 | 0.1032 | 0.0694 | 0.0957 | 0.0894 | -0.39 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 149 | 0.097 | 0.085 | 0.092 | 0.0913 | -0.31 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 194 | 0.0947 | 0.0952 | 0.0826 | 0.0908 | -0.33 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 125 | 0.0931 | 0.0902 | 0.0933 | 0.0922 | -0.28 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 149 | 0.096 | 0.096 | 0.095 | 0.0957 | -0.14 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 190 | 0.101 |  |  | 0.101 |  |  |  |  | 1 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 196 | 0.1004 | 0.1024 | 0.1009 | 0.101 | 0.08 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 138 | 0.1018 | 0.0992 | 0.1045 | 0.102 | 0.1 | 0.0993 | 26 | 0.0255 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 212 | 0.1057 | 0.1145 | 0.1029 | 0.108 | 0.33 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 200 | 0.108 | 0.111 | 0.108 | 0.109 | 0.38 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 120 | 0.1118 | 0.1152 | 0.1193 | 0.115 | 0.63 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 162 | 0.1131 | 0.1143 | 0.1187 | 0.115 | 0.63 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 169 | 0.1193 | 0.111 | 0.1175 | 0.116 | 0.65 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 122 | 0.1274 | 0.1275 | 0.1269 | 0.127 | 1.1 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 199 | 0.1309 | 0.1175 | 0.133 | 0.127 | 1.09 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 155 | 0.156 | 0.148 | 0.105 | 0.136 | 1.45 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 189 | 0.134 | 0.1411 | 0.1429 | 0.139 | 1.57 | 0.0993 | 26 | 0.0255 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 192 | 0.1583 | 0.1447 | 0.1396 | 0.148 | 1.89 | 0.0993 | 26 | 0.0255 | 0 |
|  |  |  |  | Total CBD (%AR) |  |  |  |  |  |  |  |  |
| Total CBD (%AR) | 007.40 | GC-FID | 169 | 3.6661 | 3.337 | 3.3931 | 3.47 | -2.54 | 5.6662 | 23 | 0.8652 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 191 | 4.7018 | 3.22 | 3.3475 | 3.76 | -2.21 | 5.6662 | 23 | 0.8652 | 0 |
| Total CBD (%AR) | 007.03 | AOAC 2018.11, M | 139 | 5.1067 | 4.0449 | 4.3818 | 4.51 | -1.34 | 5.6662 | 23 | 0.8652 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 107 | 4.491 | 4.237 | 5.03 | 4.59 | -1.25 | 5.6662 | 23 | 0.8652 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 200 | 5.132 | 4.95 | 4.981 | 5.02 | -0.75 | 5.6662 | 23 | 0.8652 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 189 | 5.0995 | 5.1323 | 5.1038 | 5.11 | -0.64 | 5.6662 | 23 | 0.8652 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 192 | 5.5119 | 5.2635 | 4.7437 | 5.17 | -0.57 | 5.6662 | 23 | 0.8652 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 212 | 5.4907 | 5.5022 | 5.4158 | 5.47 | -0.23 | 5.6662 | 23 | 0.8652 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 194 | 5.5474 | 5.5454 | 5.5459 | 5.55 | -0.14 | 5.6662 | 23 | 0.8652 | 0 |
| Total CBD (%AR) | 007.01 | AOAC 2018.10 | 183 | 5.5916 | 5.5525 | 5.5681 | 5.57 | -0.11 | 5.6662 | 23 | 0.8652 | 0 |

Population of Lab Values

Lab

Lab

Population of Lab Values

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%AR) | 007.10 | Other LC-UV | 149 | 5.655 | 5.521 | 5.816 | 5.66 | 0 | 5.6662 | 23 | 0.8652 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 125 | 5.84 | 5.63 | 5.69 | 5.72 | 0.06 | 5.6662 | 23 | 0.8652 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 155 | 5.74 | 5.8 | 5.78 | 5.77 | 0.12 | 5.6662 | 23 | 0.8652 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 196 | 6.09 | 6.0329 | 5.8278 | 5.98 | 0.37 | 5.6662 | 23 | 0.8652 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 204 | 5.9946 | 5.9921 | 5.9676 | 5.98 | 0.37 | 5.6662 | 23 | 0.8652 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 216 | 5.8788 | 6.1045 | 6.0279 | 6 | 0.39 | 5.6662 | 23 | 0.8652 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 199 | 5.9851 | 5.9888 | 6.1666 | 6.05 | 0.44 | 5.6662 | 23 | 0.8652 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 162 | 6.1554 | 6.3674 | 6.1204 | 6.21 | 0.63 | 5.6662 | 23 | 0.8652 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 197 | 6.23 | 6.2888 | 6.3578 | 6.29 | 0.72 | 5.6662 | 23 | 0.8652 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 149 | 6.427 | 6.31 | 6.262 | 6.33 | 0.77 | 5.6662 | 23 | 0.8652 | 0 |
| Total CBD (%AR) | 007.60 | NIR | 187 | 6.7 | 6.9 | 6.7 | 6.77 | 1.27 | 5.6662 | 23 | 0.8652 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 122 | 6.8498 | 6.8495 | 6.8501 | 6.85 | 1.37 | 5.6662 | 23 | 0.8652 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 138 | 6.8387 | 7.382 | 7.5153 | 7.25 | 1.83 | 5.6662 | 23 | 0.8652 | 0 |

| CBDV (%AR) |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDV (%AR) | 008.99 | Other | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| CBDV (%AR) | 008.99 | Other | 197 | <0.0142 | <0.0146 | <0.0138 | <0.0142 |  |  |  |  | 1 |
| CBDV (%AR) | 008.99 | Other | 216 | <0.0148 | <0.0136 | <0.0154 | <0.0148 |  |  |  |  | 1 |
| CBDV (%AR) | 008.99 | Other | 196 | <0.0210 | <0.0210 | <0.0210 | <0.0210 |  |  |  |  | 1 |
| CBDV (%AR) | 008.99 | Other | 189 | 0.018 | 0.017 | 0.0182 | 0.0177 | -0.89 | 0.0248 | 7 | 0.0079 | 0 |
| CBDV (%AR) | 008.02 | AOAC 2018.11, U | 183 | 0.0208 | 0.0213 | 0.0206 | 0.0209 | -0.49 | 0.0248 | 7 | 0.0079 | 0 |
| CBDV (%AR) | 008.99 | Other | 149 | 0.021 | 0.02 | 0.021 | 0.0207 | -0.52 | 0.0248 | 7 | 0.0079 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDV (%AR) | 008.99 | Other | 200 | 0.023 | 0.022 | 0.022 | 0.0223 | -0.31 | 0.0248 | 7 | 0.0079 | 0 |
| CBDV (%AR) | 008.99 | Other | 194 | 0.0238 | 0.0246 | 0.0191 | 0.0225 | -0.29 | 0.0248 | 7 | 0.0079 | 0 |
| CBDV (%AR) | 008.99 | Other | 106 | 0.02 | 0.04 | 0.04 | 0.0333 | 1.09 | 0.0248 | 7 | 0.0079 | 0 |
| CBDV (%AR) | 008.99 | Other | 212 | 0.0363 | 0.0358 | 0.0355 | 0.0359 | 1.41 | 0.0248 | 7 | 0.0079 | 0 |

Population of Lab Values

| CBDVA (%AR) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDVA (%AR) | 009.99 | Other | 197 | <0.0094 | <0.0098 | <0.0092 | <0.0094 | 1 |
| CBDVA (%AR) | 009.99 | Other | 216 | <0.0099 | <0.0091 | <0.0103 | <0.0099 | 1 |
| CBDVA (%AR) | 009.99 | Other | 196 | <0.0210 | <0.0210 | <0.0210 | <0.0210 | 1 |
| CBDVA (%AR) | 009.99 | Other | 149 | <0.0370 | <0.0370 | <0.0400 | <0.0370 | 1 |
| CBDVA (%AR) | 009.01 | AOAC 2018.10 | 183 | 0 | 0 | 0 | 0 | 2 |

| CBG (%AR) |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%AR) | 010.99 | Other | 149 | 0.067 | 0.07 | 0.067 | 0.068 | -1.05 | 0.1046 | 14 | 0.0348 | 0 |
| CBG (%AR) | 010.99 | Other | 197 | 0.0715 | 0.0714 | 0.071 | 0.0713 | -0.96 | 0.1046 | 14 | 0.0348 | 0 |
| CBG (%AR) | 010.99 | Other | 216 | 0.0721 | 0.0741 | 0.0725 | 0.0729 | -0.91 | 0.1046 | 14 | 0.0348 | 0 |
| CBG (%AR) | 010.99 | Other | 194 | 0.0917 | 0.0891 | 0.0792 | 0.0867 | -0.52 | 0.1046 | 14 | 0.0348 | 0 |
| CBG (%AR) | 010.99 | Other | 162 | 0.0855 | 0.0883 | 0.1004 | 0.0914 | -0.38 | 0.1046 | 14 | 0.0348 | 0 |
| CBG (%AR) | 010.99 | Other | 196 | 0.09 | 0.0914 | 0.0906 | 0.0907 | -0.4 | 0.1046 | 14 | 0.0348 | 0 |
| CBG (%AR) | 010.01 | AOAC 2018.10 | 183 | 0.0933 | 0.094 | 0.0939 | 0.0937 | -0.31 | 0.1046 | 14 | 0.0348 | 0 |
| CBG (%AR) | 010.99 | Other | 200 | 0.095 | 0.097 | 0.096 | 0.096 | -0.25 | 0.1046 | 14 | 0.0348 | 0 |
| CBG (%AR) | 010.99 | Other | 212 | 0.0971 | 0.1052 | 0.1094 | 0.104 | -0.02 | 0.1046 | 14 | 0.0348 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%AR) | 010.02 | AOAC 2018.11, U | 125 | 0.107 | 0.108 | 0.102 | 0.106 | 0.03 | 0.1046 | 14 | 0.0348 | 0 |
| CBG (%AR) | 010.99 | Other | 189 | 0.1146 | 0.1204 | 0.1198 | 0.118 | 0.39 | 0.1046 | 14 | 0.0348 | 0 |
| CBG (%AR) | 010.99 | Other | 122 | 0.1523 | 0.152 | 0.1521 | 0.152 | 1.37 | 0.1046 | 14 | 0.0348 | 0 |
| CBG (%AR) | 010.99 | Other | 155 | 0.214 | 0.19 | 0.201 | 0.202 | 2.79 | 0.1046 | 14 | 0.0348 | 0 |
| CBG (%AR) | 010.99 | Other | 106 | 0.25 | 0.25 | 0.26 | 0.253 | 4.28 | 0.1046 | 14 | 0.0348 | 0 |

Population of Lab Values

| CBGA (%AR) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%AR) | 011.02 | AOAC 2018.11, U | 125 | <0.0018 | <0.0016 | <0.0020 | <0.0018 | 1 |
| CBGA (%AR) | 011.99 | Other | 197 | <0.0094 | <0.0098 | <0.0092 | <0.0094 | 1 |
| CBGA (%AR) | 011.99 | Other | 216 | <0.0099 | <0.0091 | <0.0103 | <0.0099 | 1 |
| CBGA (%AR) | 011.99 | Other | 106 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| CBGA (%AR) | 011.99 | Other | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| CBGA (%AR) | 011.99 | Other | 196 | <0.0210 | <0.0210 | <0.0210 | <0.0210 | 1 |
| CBGA (%AR) | 011.99 | Other | 194 | <0.0270 | <0.0270 | <0.0270 | <0.0270 | 1 |
| CBGA (%AR) | 011.99 | Other | 149 | <0.0370 | <0.0370 | <0.0400 | <0.0370 | 1 |
| CBGA (%AR) | 011.99 | Other | 122 | <0.0500 | <0.0500 | <0.0500 | <0.0500 | 1 |
| CBGA (%AR) | 011.01 | AOAC 2018.10 | 183 | 0 | 0 | 0 | 0 | 2 |
| CBGA (%AR) | 011.99 | Other | 189 | 0.0245 | 0.0235 | 0.0236 | 0.0239 | 0 |

|  |  |  | THCV (%AR) |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| THCV (%AR) | 012.99 | Other | 197 | <0.0094 <0.0098 <0.0092 | <0.0094 | 1 |
| THCV (%AR) | 012.99 | Other | 216 | <0.0099 <0.0091 <0.0103 | <0.0099 | 1 |

Lab

Lab

Population of Lab Values

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| THCV (%AR) | 012.99 | Other | 106 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 196 | <0.0210 | <0.0210 | <0.0210 | <0.0210 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 194 | <0.0270 | <0.0270 | <0.0270 | <0.0270 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 149 | <0.0370 | <0.0370 | <0.0400 | <0.0370 |  |  |  |  | 1 |
| THCV (%AR) | 012.02 | AOAC 2018.11, U | 183 | 0 | 0 | 0 | 0 |  |  |  |  | 2 |
| THCV (%AR) | 012.99 | Other | 200 | 0.009 | 0.009 | 0.009 | 0.009 |  |  |  |  | 0 |
| THCV (%AR) | 012.99 | Other | 189 | 0.0142 | 0.0153 | 0.0157 | 0.0151 |  |  |  |  | 0 |

| Δ8-THC (%AR) |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%AR) | 013.99 | Other | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 200 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 196 | <0.0210 | <0.0210 | <0.0210 | <0.0210 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 194 | <0.0270 | <0.0270 | <0.0270 | <0.0270 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 149 | <0.0370 | <0.0370 | <0.0400 | <0.0370 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 102 | <0.0484 | <0.0188 | <0.0187 | <0.0484 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 183 | 0 | 0 | 0 | 0 |  |  |  |  | 2 |
| Δ8-THC (%AR) | 013.01 | AOAC 2018.10 | 183 | 0.003 | 0.003 | 0.0031 | 0.003 | -1.27 | 0.0209 | 6 | 0.0141 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 216 | 0.0187 | 0.0171 | 0.0189 | 0.0182 | -0.19 | 0.0209 | 6 | 0.0141 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 106 | 0.02 | 0.02 | 0.02 | 0.02 | -0.07 | 0.0209 | 6 | 0.0141 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 197 | 0.0194 | 0.0193 | 0.0206 | 0.0198 | -0.08 | 0.0209 | 6 | 0.0141 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 189 | 0.0187 | 0.0251 | 0.0239 | 0.0226 | 0.12 | 0.0209 | 6 | 0.0141 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 138 | 0.0757 | 0.0687 | 0.0766 | 0.0737 | 3.74 | 0.0209 | 6 | 0.0141 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | CBC (%AR) |  |  |  |  |  |  |  |
| CBC (%AR) | 014.99 | Other | 212 | 0.2289 | 0.2364 | 0.2337 | 0.233 | -1.21 | 0.2501 | 14 | 0.0142 | 0 |
| CBC (%AR) | 014.99 | Other | 194 | 0.2436 | 0.2418 | 0.2174 | 0.234 | -1.12 | 0.2501 | 14 | 0.0142 | 0 |
| CBC (%AR) | 014.01 | AOAC 2018.10 | 183 | 0.2402 | 0.2354 | 0.2436 | 0.24 | -0.73 | 0.2501 | 14 | 0.0142 | 0 |
| CBC (%AR) | 014.99 | Other | 106 | 0.25 | 0.25 | 0.22 | 0.24 | -0.71 | 0.2501 | 14 | 0.0142 | 0 |
| CBC (%AR) | 014.99 | Other | 162 | 0.2395 | 0.2443 | 0.2405 | 0.241 | -0.61 | 0.2501 | 14 | 0.0142 | 0 |
| CBC (%AR) | 014.99 | Other | 216 | 0.2404 | 0.246 | 0.2466 | 0.244 | -0.41 | 0.2501 | 14 | 0.0142 | 0 |
| CBC (%AR) | 014.99 | Other | 149 | 0.243 | 0.245 | 0.248 | 0.245 | -0.33 | 0.2501 | 14 | 0.0142 | 0 |
| CBC (%AR) | 014.99 | Other | 197 | 0.2479 | 0.2476 | 0.2566 | 0.251 | 0.04 | 0.2501 | 14 | 0.0142 | 0 |
| CBC (%AR) | 014.99 | Other | 196 | 0.2509 | 0.2532 | 0.2511 | 0.252 | 0.12 | 0.2501 | 14 | 0.0142 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, U | 125 | 0.251 | 0.256 | 0.252 | 0.253 | 0.21 | 0.2501 | 14 | 0.0142 | 0 |
| CBC (%AR) | 014.99 | Other | 155 | 0.265 | 0.274 | 0.257 | 0.265 | 1.08 | 0.2501 | 14 | 0.0142 | 0 |
| CBC (%AR) | 014.99 | Other | 189 | 0.2619 | 0.2673 | 0.2681 | 0.266 | 1.11 | 0.2501 | 14 | 0.0142 | 0 |
| CBC (%AR) | 014.99 | Other | 200 | 0.269 | 0.268 | 0.267 | 0.268 | 1.27 | 0.2501 | 14 | 0.0142 | 0 |
| CBC (%AR) | 014.99 | Other | 204 | 0.2667 | 0.2712 | 0.267 | 0.268 | 1.29 | 0.2501 | 14 | 0.0142 | 0 |
|  |  |  |  | Cd (ug/g AR) |  |  |  |  |  |  |  |  |
| Cd (ug/g AR) | 101.04 | ICP-MS | 102 | <0.008 | <0.008 | <0.008 | <0.008 |  |  |  |  | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 194 | <0.0200 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 196 | <0.0241 | <0.0241 | <0.0241 | <0.0241 |  |  |  |  | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 106 | <0.0750 | <0.0750 | <0.0750 | <0.0750 |  |  |  |  | 1 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cd (ug/g AR) | 101.04 | ICP-MS | 197 | <0.0963 | <0.1002 | <0.0994 | <0.0963 |  |  |  |  | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 | <0.2000 |  |  |  |  | 1 |
|  |  |  |  | As (ug/g AR) |  |  |  |  |  |  |  |  |
| As (ug/g AR) | 102.04 | ICP-MS | 102 | <0.008 | <0.008 | <0.008 | <0.008 |  |  |  |  | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 194 | <0.0200 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 196 | <0.0482 | <0.0482 | <0.0482 | <0.0482 |  |  |  |  | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 106 | <0.0750 | <0.0750 | <0.0750 | <0.0750 |  |  |  |  | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 197 | <0.0963 | <0.1002 | <0.0994 | <0.0963 |  |  |  |  | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 | <0.2000 |  |  |  |  | 1 |
|  |  |  |  | Hg (ug/g AR) |  |  |  |  |  |  |  |  |
| Hg (ug/g AR) | 103.04 | ICP-MS | 196 | <0.0072 | <0.0072 | <0.0072 | <0.0072 |  |  |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 106 | <0.0080 | <0.0080 | <0.0080 | <0.0080 |  |  |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 194 | <0.0200 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 197 | <0.0963 | <0.1002 | <0.0994 | <0.0963 |  |  |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 | <0.2000 |  |  |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 102 | 0.0013 | 0.0014 | 0.0012 | 0.0013 |  |  |  |  | 0 |
|  |  |  |  | Pb (ug/g AR) |  |  |  |  |  |  |  |  |
| Pb (ug/g AR) | 104.04 | ICP-MS | 102 | <0.008 <0.008 <0.008 |  |  | <0.008 |  |  |  |  | 1 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 194 | <0.0200 <0.0200 <0.0200 |  |  | <0.0200 |  |  |  |  | 1 |

Population of Lab Values

Lab

Lab

Population of Lab Values

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Pb (ug/g AR) | 104.04 | ICP-MS | 196 | <0.0241 | <0.0241 | <0.0241 | <0.0241 |  |  |  |  | 1 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 106 | <0.0750 | <0.0750 | <0.0750 | <0.0750 |  |  |  |  | 1 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 197 | <0.0963 | <0.1002 | <0.0994 | <0.0963 |  |  |  |  | 1 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 | <0.2000 |  |  |  |  | 1 |

Cu (ug/g AR)

Cu (ug/g AR)

105.04	ICP-MS

196

<0.4893  <0.4893	<0.4893	<0.4893	1

Ni (ug/g AR)

Ni (ug/g AR)

106.04	ICP-MS

196

<0.0722  <0.0722	<0.0722	<0.0722	1

|  |  |  | Cr (ug/g AR) |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Cr (ug/g AR) | 107.04 | ICP-MS | 196 | <0.0722 <0.0722 <0.0722 | <0.0722 | 1 |
| Cr (ug/g AR) | 107.04 | ICP-MS | 197 | <0.0963 <0.1002 <0.0994 | <0.0963 | 1 |

|  |  |  | Se (ug/g AR) |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Se (ug/g AR) | 108.04 | ICP-MS | 197 | <0.0963 <0.1002 <0.0994 | <0.0963 | 1 |
| Se (ug/g AR) | 108.04 | ICP-MS | 196 | <0.0963 <0.0963 <0.0963 | <0.0963 | 1 |

Zn (ug/g AR)

Zn (ug/g AR)

111.04	ICP-MS

196

1.84	1.66	1.89	1.8	0

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Sb (ug/g AR) |  |  |  |  |  |  |  |  |
| Sb (ug/g AR) | 112.04 | ICP-MS | 196 | <0.0241 <0.0241 <0.0241 |  |  | <0.0241 |  |  |  |  | 1 |
|  |  |  |  | (-)-alpha-Bisabolol (%AR) |  |  |  |  |  |  |  |  |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 162 | 0.021 | 0.0165 | 0.0191 | 0.0189 |  |  |  |  | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 196 | 0.1736 | 0.1655 | 0.16 | 0.166 |  |  |  |  | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 194 | 0.1686 | 0.1852 | 0.1844 | 0.179 |  |  |  |  | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 212 | 0.1906 | 0.2261 | 0.1899 | 0.202 |  |  |  |  | 0 |
|  |  |  |  | beta-Caryophyllene (%AR) |  |  |  |  |  |  |  |  |
| beta-Caryophyllene (% | 202.02 | GC,MS | 162 | 0.0819 | 0.0709 | 0.0768 | 0.0765 |  |  |  |  | 0 |
| beta-Caryophyllene (% | 202.02 | GC,MS | 194 | 0.1333 | 0.1533 | 0.1302 | 0.139 |  |  |  |  | 0 |
| beta-Caryophyllene (% | 202.02 | GC,MS | 212 | 0.1739 | 0.1965 | 0.1465 | 0.172 |  |  |  |  | 0 |
| beta-Caryophyllene (% | 202.02 | GC,MS | 196 | 0.1817 | 0.1854 | 0.1718 | 0.18 |  |  |  |  | 0 |
|  |  |  | (-)-Caryophyllene oxide (%AR) |  |  |  |  |  |  |  |  |  |
| (-)-Caryophyllene oxide | 203.02 | GC,MS | 194 | 0.0436 | 0.0419 | 0.0432 | 0.0429 |  |  |  |  | 0 |
| (-)-Caryophyllene oxide | 203.02 | GC,MS | 212 | 0.0487 | 0.0503 | 0.0449 | 0.048 |  |  |  |  | 0 |
| (-)-Caryophyllene oxide | 203.02 | GC,MS | 196 | 0.0689 | 0.0639 | 0.0635 | 0.0654 |  |  |  |  | 0 |
|  |  |  |  | Geraniol (%AR) |  |  |  |  |  |  |  |  |
| Geraniol (%AR) | 204.02 | GC,MS | 194 | <0.0004 <0.0004 <0.0004 |  |  | <0.0004 |  |  |  |  | 1 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Geraniol (%AR) | 204.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 |  |  |  |  | 1 |
| Geraniol (%AR) | 204.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
|  |  |  |  | alpha-Humulene (%AR) |  |  |  |  |  |  |  |  |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 162 | 0.0225 | 0.0189 | 0.0202 | 0.0205 |  |  |  |  | 0 |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 194 | 0.0495 | 0.0553 | 0.0489 | 0.0512 |  |  |  |  | 0 |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 196 | 0.055 | 0.0549 | 0.0517 | 0.0539 |  |  |  |  | 0 |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 212 | 0.0536 | 0.0653 | 0.0468 | 0.0552 |  |  |  |  | 0 |
|  |  |  |  | d-Limonene (%AR) |  |  |  |  |  |  |  |  |
| d-Limonene (%AR) | 206.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 |  |  |  |  | 1 |
| d-Limonene (%AR) | 206.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| d-Limonene (%AR) | 206.02 | GC,MS | 194 | 0.0025 | 0.0027 | 0.0025 | 0.0026 |  |  |  |  | 0 |
|  |  |  |  | Linalool (%AR) |  |  |  |  |  |  |  |  |
| Linalool (%AR) | 207.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Linalool (%AR) | 207.02 | GC,MS | 212 | 0.0006 | 0.0013 | 0.0017 | 0.0012 |  |  |  |  | 0 |
| Linalool (%AR) | 207.02 | GC,MS | 196 | 0.0017 | 0.0016 | 0.0017 | 0.0017 |  |  |  |  | 0 |
| Linalool (%AR) | 207.02 | GC,MS | 194 | 0.0058 | 0.0058 | 0.0059 | 0.0058 |  |  |  |  | 0 |
|  |  |  |  | beta-Myrcene (%AR) |  |  |  |  |  |  |  |  |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 196 | <0.0013 <0.0013 <0.0013 |  |  | <0.0013 |  |  |  |  | 1 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 194 | 0.0023 | 0.0023 | 0.0024 | 0.0023 |  |  |  |  | 0 |
|  |  |  |  | Ocimene (%AR) |  |  |  |  |  |  |  |  |
| Ocimene (%AR) | 210.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 | <0.0004 |  |  |  |  | 1 |
| Ocimene (%AR) | 210.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 |  |  |  |  | 1 |
| Ocimene (%AR) | 210.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
|  |  |  |  | alpha-Pinene (%AR) |  |  |  |  |  |  |  |  |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 | <0.0004 |  |  |  |  | 1 |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 |  |  |  |  | 1 |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
|  |  |  |  | (-)-beta-Pinene (%AR) |  |  |  |  |  |  |  |  |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 | <0.0004 |  |  |  |  | 1 |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 |  |  |  |  | 1 |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
|  |  |  |  | alpha-Terpinene (%AR) |  |  |  |  |  |  |  |  |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 | <0.0004 |  |  |  |  | 1 |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 |  |  |  |  | 1 |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |

Population of Lab Values

| Analyte Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

Population of Lab Values

|  |  |  | gamma-Terpinene (%AR) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| gamma-Terpinene (%A | 214.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 | <0.0004 | 1 |
| gamma-Terpinene (%A | 214.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 | 1 |
| gamma-Terpinene (%A | 214.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |

|  |  |  | alpha-Terpinolene (%AR) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 194 | <0.0004 | <0.0004 | <0.0004 | <0.0004 | 1 |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 | 1 |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |

| Δ9-THC (%DW) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%DW) | 501.01 | AOAC 2018.10 | 210 | 0.091 | 0.091 | 0.089 | 0.0903 | 0 |

Δ9-THCA (%DW)

Δ9-THCA (%DW)

502.01	AOAC 2018.10

210

<0.0460  <0.0430	<0.0470	<0.0460	1

| CBD (%DW) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%DW) | 503.01 | AOAC 2018.10 | 210 | 5.644 | 5.733 | 5.694 | 5.69 | 0 |

CBDA (%DW)

CBDA (%DW)

504.01	AOAC 2018.10

210

<0.0460  <0.0430	<0.0470	<0.0460	1

| Analyte Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

Population of Lab Values

Total Δ9-THC (%DW)

Total Δ9-THC (%DW)

506.01	AOAC 2018.10

210

0.131	0.129	0.13	0.13	0

Total CBD (%DW)

Total CBD (%DW)

507.01	AOAC 2018.10

210

5.684	5.771	5.735	5.73	0

CBDV (%DW)

CBDV (%DW)

508.02	AOAC 2018.11, U

210

<0.0430  <0.0400	<0.0430	<0.0430	1

CBDVA (%DW)

CBDVA (%DW)

509.01	AOAC 2018.10

210

<0.0460  <0.0430	<0.0460	<0.0460	1

CBG (%DW)

CBG (%DW)

510.01	AOAC 2018.10

210

0.109	0.102	0.107	0.106	0

CBGA (%DW)

CBGA (%DW)

511.01	AOAC 2018.10

210

<0.0460  <0.0430	<0.0460	<0.0460	1

THCV (%DW)

THCV (%DW)

512.02	AOAC 2018.11, U

210

<0.0430  <0.0400	<0.0430	<0.0430	1

Δ8-THC (%DW)

Analyte

Code

Method

Lab Num

Result1	Result2	Result3

Lab

Value	Z score

Population of Lab Values	 Rob Mean	n	Rob StDev

Flag

Δ8-THC (%DW)

513.01	AOAC 2018.10

210

0.043	0.041	0.043	0.0423	0

- Flag definitions: 1 (Lab Value not used because there were less than 2 numeric lab results), 2 (Lab Value not used because there were less than 2 numeric lab results greater than 0).

- Robust statistics performed to determine robust mean and Z scores only if there were 6 or more numeric lab values.

- AR refers to concentration on as-received basis. DW refers to concentration on dry weight basis.