| Sample # | Replicate 1 | Replicate 2 |
| --- | --- | --- |
| 1 | 0.0552 | 0.0558 |
| 2 | 0.0537 | 0.0533 |
| 3 | 0.0553 | 0.0553 |
| 4 | 0.0541 | 0.0547 |
| 5 | 0.0535 | 0.0562 |
| 6 | 0.0729 | 0.0723 |
| 7 | 0.0692 | 0.0716 |
| 8 | 0.0718 | 0.0697 |
| 9 | 0.0734 | 0.0702 |
| 10 | 0.0701 | 0.0688 |

| Sample # | Replicate 1 | Replicate 2 |
| --- | --- | --- |
| 1 | 7.1385 | 7.1180 |
| 2 | 7.1068 | 7.1332 |
| 3 | 7.1549 | 7.1530 |
| 4 | 7.1842 | 7.1490 |
| 5 | 7.1217 | 7.1391 |
| 6 | 7.3166 | 7.3528 |
| 7 | 7.3656 | 7.2878 |
| 8 | 7.3873 | 7.3515 |
| 9 | 7.3941 | 7.3542 |
| 10 | 7.3888 | 7.4133 |

Homogeneity and Stability

Sample ID . . . . . . . . . . . . . . . . . . . . . . . .

Samples shipped . . . . . . . . . . . . . . . . . . .

Results due date . . . . . . . . . . . . . . . . . . . .

OL22SEP-1

20-Sep-22

1-Nov-22

## Homogeneity

10 sample packets randomly chosen for analysis.

| Sample # | Replicate 1 | Replicate 2 |
| --- | --- | --- |
| 1 | 0.0846 | 0.0951 |
| 2 | 0.0949 | 0.0953 |
| 3 | 0.0949 | 0.0943 |
| 4 | 0.0946 | 0.0938 |
| 5 | 0.0935 | 0.0958 |
| 6 | 0.0969 | 0.0965 |
| 7 | 0.0951 | 0.0964 |
| 8 | 0.0955 | 0.0966 |
| 9 | 0.0976 | 0.0964 |
| 10 | 0.0963 | 0.0967 |

Duplicate test portions analyzed for analyte and method shown in table caption.

| Total delta-9 THC (%AR) via GC-FID (method 006.40) |  |  | %RSD | HorRat-r |
| --- | --- | --- | --- | --- |
| Analysis Date: 9/22/22 | Overall average: | 0.0950 |  |  |
|  | SD of sample avgs: | 0.002056 | 2.16 |  |
|  | repeatability SD: | 0.002472 | 2.60 | 0.5 |
|  | between-sample SD: | 0.00108 | 1.14 |  |
|  | reproducibilty SD: | 0.002699 | 2.84 |  |
|  | SD used for Analyte z scores: | 0.0262 |  |  |
|  | Check value: | 0.00786 |  |  |
| Total delta-9 THC (%AR) via GC-FID (method 006.40) Is between-sample SD less than check value? YES Homogeneity test passed |  |  |  |  |
| delta-8 THC (%AR) via GC-FID (method 0013.99) |  |  | %RSD | HorRat-r |
| Analysis Date: 9/22/22 | Overall average: | 0.0629 |  |  |
|  | SD of sample avgs: | 0.008646 | 13.75 |  |
|  | repeatability SD: | 0.001234 | 1.96 | 0.3 |
|  | between-sample SD: | 0.0086 | 13.68 |  |
|  | reproducibilty SD: | 0.00869 | 13.82 |  |
|  | SD used for Analyte z scores: | 0.0197 |  |  |
|  | Check value: | 0.00591 |  |  |
| delta-8 THC (%AR) via GC-FID (method 0013.99) Is between-sample SD less than check value? NO Homogeneity test did NOT pass |  |  |  |  |
| Total CBD (%AR) via GC-FID (method 007.40) |  |  | %RSD | HorRat-r |
| Analysis Date: 9/22/22 | Overall average: | 7.2505 |  |  |
|  | SD of sample avgs: | 0.1191 | 1.64 |  |
|  | repeatability SD: | 0.02598 | 0.36 | 0.1 |
|  | between-sample SD: | 0.1177 | 1.62 |  |
|  | reproducibilty SD: | 0.1205 | 1.66 |  |
|  | SD used for Analyte z scores: | 0.8339 |  |  |
|  | Check value: | 0.25017 |  |  |
| Total CBD (%AR) via GC-FID (method 007.40) Is between-sample SD less than check value? YES Homogeneity test passed |  |  |  |  |

## Stability

Sample tested on different days from date sample shipped to results due date. Sample stored at room temperature during period of testing.

| Total THC (%AR) via GC-FID (method 006.40) |  |  |
| --- | --- | --- |
| Date of Analysis | Days | Conc. |
| 9/13/22 | 0 | 0.1025 |
| 9/19/22 | 6 | 0.1143 |
| 9/23/22 | 10 | 0.0884 |
| 9/30/22 | 17 | 0.0923 |
| 10/7/22 | 24 | 0.1173 |
| 10/14/22 | 31 | 0.0947 |
| 10/19/22 | 36 | 0.0962 |
| 10/24/22 | 41 | 0.1015 |
| 10/31/22 | 48 | 0.1015 |

Days = x variable, Conc = y variable SUMMARY OUTPUT

Regression Statistics

Multiple R  0.0933786 R Square	0.0087196 Adjusted R Squ -0.1328919 Standard Error  0.0102541

Observations	9

ANOVA

0.14

0.12

0.10

0.08

0.06

y = -5E-05x + 0.1022

R² = 0.0087

0.04

0.02

0.00

0

10

20

30

40

50

df	SS	MS	F	Significance F

Regression	1  6.4743E-06 6.47428E-06 0.061574  0.811147489

Residual	7 0.00073603 0.000105147

Total	8	0.0007425

Coefficients tandard Error	t Stat	P-value	Lower 95%	Upper 95%

Intercept	0.10225 0.00619907 16.49439221  7.35E-07  0.087591474  0.116908435

X Variable 1	-5.422E-05 0.00021852 -0.24814081 0.811147 -0.000570938  0.000462491

Coefficient for x variable is not statistically significantly different from 0. No signficant change in result during testing period

Stability test passed.

%RSD = 9.5	HorRat-R = 1.7

| delta-8 THC (%AR) via GC-FID (method 006.40) |  |  |
| --- | --- | --- |
| Date of Analysis | Days | Conc. |
| 9/13/22 | 0 | 0.0786 |
| 9/19/22 | 6 | 0.0974 |
| 9/23/22 | 10 | 0.0685 |
| 9/30/22 | 17 | 0.0680 |
| 10/7/22 | 24 | 0.0836 |
| 10/14/22 | 31 | 0.0666 |
| 10/19/22 | 36 | 0.0785 |
| 10/24/22 | 41 | 0.0814 |
| 10/31/22 | 48 | 0.0801 |
| 0.12 0.10 |  |  |

Days = x variable, Conc = y variable SUMMARY OUTPUT

Regression Statistics

Multiple R  0.0983847 R Square	0.0096796 Adjusted R Squ -0.1317948 Standard Error  0.0102626

Observations	9

ANOVA

df	SS	MS	F	Significance F

Regression	1	7.206E-06	7.206E-06 0.068419  0.801181997

y = -6E-05x + 0.0794 R² = 0.0097

Residual	7 0.00073725 0.000105321

Total	8 0.00074446

0.08

0.06

0.04

0.02

0.00

0

10

20

30

40

50

Coefficients tandard Error	t Stat	P-value	Lower 95%	Upper 95%

| Intercept 0.0794316 0.00620423 12.80282985 4.11E-06 | 0.064760981 | 0.094102305 |
| --- | --- | --- |
| X Variable 1 -5.721E-05 0.0002187 -0.26157049 0.801182 | -0.00057435 | 0.000459939 |

Coefficient for x variable is not statistically significantly different from 0. No signficant change in result during testing period

Stability test passed.

%RSD = 12.4	HorRat-R = 2.1

| Total CBD (%AR) via GC-FID (method 006.40) |  |  |
| --- | --- | --- |
| Date of Analysis | Days | Conc. |
| 9/13/22 | 0 | 6.8387 |
| 9/19/22 | 6 | 7.3820 |
| 9/23/22 | 10 | 7.5153 |
| 9/30/22 | 17 | 7.5745 |
| 10/7/22 | 24 | 7.4163 |
| 10/14/22 | 31 | 6.3443 |
| 10/19/22 | 36 | 7.0135 |
| 10/24/22 | 41 | 7.2540 |
| 10/31/22 | 48 | 6.5998 |

Days = x variable, Conc = y variable SUMMARY OUTPUT

Regression Statistics

Multiple R		0.411678 R Square	0.1694788 Adjusted R Squ 0.0508329 Standard Error  0.4215525

Observations	9

ANOVA

df	SS	MS	F	Significance F

Regression	1 0.25384336  0.25384336 1.428442  0.270934912

y = -0.0107x + 7.3584 R² = 0.1695

8.50

7.50

6.50

5.50

4.50

3.50

2.50

1.50

0

10

20

30

40

50

Residual	7	1.2439453 0.177706471

Total	8 1.49778866

Coefficients tandard Error	t Stat	P-value	Lower 95%	Upper 95%

Intercept	7.3583706 0.25484784 28.87358453  1.54E-08  6.755751196  7.960989949

X Variable 1	-0.0107368 0.00898345  -1.1951744 0.270935 -0.031979259	0.01050569

Coefficient for x variable is not statistically significantly different from 0. No signficant change in result during testing period

Stability test passed.

%RSD = 6.1	HorRat-R = 2