Hemp Proficiency Testing

All Labs Trueness OL23NOV-1 (oil)

Sample: OL23NOV-1 (oil)

All Labs Trueness

Issue Date:

January 7, 2024

| Analyte | Code | Method | Lab Num |  | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Δ9-THC (%AR) |  |  |  |  |  |  |  |  |  |  |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 |  | 210 | <0.0670 | <0.0690 | <0.0700 | <0.0670 |  |  |  |  | 1 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 192 | 0.0033 | 0.0031 | 0.0028 | 0.0031 | -5.98 | 0.0693 | 16 | 0.0111 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 197 | 0.058 | 0.0521 | 0.0561 | 0.0554 | -1.25 | 0.0693 | 16 | 0.0111 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U |  | 103 | 0.058 | 0.057 | 0.055 | 0.0567 | -1.14 | 0.0693 | 16 | 0.0111 | 0 |
| Δ9-THC (%AR) | 001.99 | Other |  | 215 | 0.0628 | 0.0597 | 0.0624 | 0.0616 | -0.69 | 0.0693 | 16 | 0.0111 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 106 | 0.06 | 0.06 | 0.07 | 0.0633 | -0.54 | 0.0693 | 16 | 0.0111 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U |  | 125 | 0.0676 | 0.0659 | 0.0664 | 0.0666 | -0.24 | 0.0693 | 16 | 0.0111 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 144 | 0.0694 | 0.0666 | 0.0667 | 0.0676 | -0.16 | 0.0693 | 16 | 0.0111 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS |  | 190 | 0.068 | 0.067 |  | 0.0675 | -0.16 | 0.0693 | 16 | 0.0111 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 196 | 0.0702 | 0.0674 | 0.0692 | 0.0689 | -0.03 | 0.0693 | 16 | 0.0111 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 189 | 0.0742 | 0.0732 | 0.0716 | 0.073 | 0.34 | 0.0693 | 16 | 0.0111 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS |  | 102 | 0.076 | 0.065 | 0.08 | 0.0737 | 0.4 | 0.0693 | 16 | 0.0111 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 184 | 0.076 | 0.08 | 0.073 | 0.0763 | 0.64 | 0.0693 | 16 | 0.0111 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U |  | 199 | 0.0775 | 0.0785 | 0.0765 | 0.0775 | 0.74 | 0.0693 | 16 | 0.0111 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 162 | 0.0726 | 0.0869 | 0.0817 | 0.0804 | 1 | 0.0693 | 16 | 0.0111 | 0 |
| Δ9-THC (%AR) | 001.03 | AOAC 2018.11, M | 139 | 0.0891 | 0.0802 | 0.0754 | 0.0816 | 1.11 | 0.0693 | 16 | 0.0111 | 0 |
| Δ9-THC (%AR) | 001.99 | Other | 183 | 0.0858 | 0.0854 | 0.086 | 0.0857 | 1.49 | 0.0693 | 16 | 0.0111 | 0 |
|  |  |  |  | Δ9-THCA (%AR) |  |  |  |  |  |  |  |  |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 190 | <0.0010 | <0.0010 |  | <0.0010 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.99 | Other | 183 | <0.0076 | <0.0076 | <0.0076 | <0.0076 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 144 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 197 | <0.0101 | <0.0101 | <0.0090 | <0.0101 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.03 | AOAC 2018.11, M | 139 | <0.0125 | <0.0125 | <0.0125 | <0.0125 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 199 | <0.0200 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 196 | <0.0210 | <0.0210 | <0.0210 | <0.0210 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 210 | <0.0230 | <0.0250 | <0.0230 | <0.0230 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 106 | <0.0300 | <0.0300 | <0.0300 | <0.0300 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 102 | <0.0480 | <0.0480 | <0.0480 | <0.0480 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.99 | Other | 215 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 192 | 0 | 0 | 0.0024 | 0.0024 |  |  |  |  | 2 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 189 | 0.0087 | 0.0121 | 0.0071 | 0.0093 |  |  |  |  | 0 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 125 | 0.0135 | 0.0142 | 0.0136 | 0.0138 |  |  |  |  | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 184 | 0.343 | 0.356 | 0.324 | 0.341 |  |  |  |  | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | CBD (%AR) |  |  |  |  |  |  |  |  |
| CBD (%AR) | 003.10 | Other LC-UV | 192 | 0.1969 | 0.2032 | 0.1967 | 0.199 | -7.15 | 3.8467 | 17 | 0.5104 | 0 |
| CBD (%AR) | 003.03 | AOAC 2018.11, M | 139 | 2.428 | 2.333 | 2.4367 | 2.4 | -2.84 | 3.8467 | 17 | 0.5104 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 170 | 3.3348 | 3.3116 | 3.3355 | 3.33 | -1.02 | 3.8467 | 17 | 0.5104 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 190 | 3.343 | 3.43 |  | 3.39 | -0.9 | 3.8467 | 17 | 0.5104 | 0 |
| CBD (%AR) | 003.99 | Other | 215 | 3.4309 | 3.4388 | 3.5034 | 3.46 | -0.76 | 3.8467 | 17 | 0.5104 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 189 | 3.5271 | 3.5207 | 3.532 | 3.53 | -0.63 | 3.8467 | 17 | 0.5104 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 196 | 3.8675 | 3.7859 | 3.895 | 3.85 | 0.01 | 3.8467 | 17 | 0.5104 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 162 | 4.0117 | 3.9202 | 3.9423 | 3.96 | 0.22 | 3.8467 | 17 | 0.5104 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 199 | 4.0095 | 4.0085 | 4.0395 | 4.02 | 0.34 | 3.8467 | 17 | 0.5104 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 103 | 4.113 | 4.115 | 3.905 | 4.04 | 0.39 | 3.8467 | 17 | 0.5104 | 0 |
| CBD (%AR) | 003.01 | AOAC 2018.10 | 210 | 4.055 | 4.023 | 4.045 | 4.04 | 0.38 | 3.8467 | 17 | 0.5104 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 106 | 4.1 | 4.04 | 4.05 | 4.06 | 0.42 | 3.8467 | 17 | 0.5104 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 144 | 4.003 | 4.0879 | 4.2342 | 4.11 | 0.51 | 3.8467 | 17 | 0.5104 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 125 | 4.23 | 4.24 | 4.19 | 4.22 | 0.73 | 3.8467 | 17 | 0.5104 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 197 | 4.2921 | 4.3234 | 4.4531 | 4.36 | 1 | 3.8467 | 17 | 0.5104 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 184 | 4.296 | 4.612 | 4.273 | 4.39 | 1.07 | 3.8467 | 17 | 0.5104 | 0 |
| CBD (%AR) | 003.99 | Other | 183 | 4.48 | 4.48 | 4.48 | 4.48 | 1.24 | 3.8467 | 17 | 0.5104 | 0 |
|  |  |  |  | CBDA (%AR) |  |  |  |  |  |  |  |  |
| CBDA (%AR) | 004.99 | Other | 183 | <0.0043 <0.0043 <0.0043 |  |  | <0.0043 |  |  |  |  | 1 |
| CBDA (%AR) | 004.10 | Other LC-UV | 162 | <0.0100 <0.0100 <0.0100 |  |  | <0.0100 |  |  |  |  | 1 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDA (%AR) | 004.10 | Other LC-UV | 197 | <0.0152 | <0.0152 | <0.0136 | <0.0152 |  |  |  |  | 1 |
| CBDA (%AR) | 004.10 | Other LC-UV | 192 | 0.0003 | 0.0006 | 0.0004 | 0.0004 | -1.54 | 0.0187 | 12 | 0.0119 | 0 |
| CBDA (%AR) | 004.03 | AOAC 2018.11, M | 139 | 0.004 | 0.0038 | 0.0035 | 0.0038 | -1.26 | 0.0187 | 12 | 0.0119 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 189 | 0.0132 | 0.0148 | 0.0133 | 0.0138 | -0.42 | 0.0187 | 12 | 0.0119 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 144 | 0.0148 | 0.0166 | 0.0164 | 0.0159 | -0.24 | 0.0187 | 12 | 0.0119 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 190 | 0.015 | 0.016 |  | 0.0155 | -0.27 | 0.0187 | 12 | 0.0119 | 0 |
| CBDA (%AR) | 004.99 | Other | 215 | 0.0166 | 0.0175 | 0.0162 | 0.0168 | -0.17 | 0.0187 | 12 | 0.0119 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 103 | 0.018 | 0.018 | 0.017 | 0.0177 | -0.09 | 0.0187 | 12 | 0.0119 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 125 | 0.0197 | 0.019 | 0.0206 | 0.0198 | 0.09 | 0.0187 | 12 | 0.0119 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 199 | 0.023 | 0.024 | 0.024 | 0.0237 | 0.42 | 0.0187 | 12 | 0.0119 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 196 | 0.0282 | 0.0278 | 0.0279 | 0.028 | 0.78 | 0.0187 | 12 | 0.0119 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 106 | <0.0300 | <0.0300 | 0.03 | 0.03 |  |  |  |  | 1 |
| CBDA (%AR) | 004.10 | Other LC-UV | 184 | 0.033 | 0.033 | 0.032 | 0.0327 | 1.18 | 0.0187 | 12 | 0.0119 | 0 |
| CBDA (%AR) | 004.01 | AOAC 2018.10 | 210 | 0.037 | 0.036 | 0.036 | 0.0363 | 1.48 | 0.0187 | 12 | 0.0119 | 0 |
|  |  |  |  | CBN (%AR) |  |  |  |  |  |  |  |  |
| CBN (%AR) | 005.10 | Other LC-UV | 192 | 0.002 | 0.0025 | 0.0019 | 0.0021 | -5.32 | 0.0351 | 14 | 0.0062 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 107 | 0.0341 | 0.0288 | 0.0239 | 0.0289 | -1 | 0.0351 | 14 | 0.0062 | 0 |
| CBN (%AR) | 005.03 | AOAC 2018.11, M | 139 | 0.0286 | 0.0317 | 0.0343 | 0.0315 | -0.58 | 0.0351 | 14 | 0.0062 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 144 | 0.0308 | 0.0293 | 0.0315 | 0.0305 | -0.74 | 0.0351 | 14 | 0.0062 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 197 | 0.0329 | 0.0321 | 0.0312 | 0.0321 | -0.49 | 0.0351 | 14 | 0.0062 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 199 | 0.0335 | 0.033 | 0.0335 | 0.0333 | -0.29 | 0.0351 | 14 | 0.0062 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%AR) | 005.10 | Other LC-UV | 106 | 0.03 | 0.04 | 0.03 | 0.0333 | -0.29 | 0.0351 | 14 | 0.0062 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 196 | 0.0347 | 0.0353 | 0.0355 | 0.0352 | 0.01 | 0.0351 | 14 | 0.0062 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 189 | 0.0349 | 0.0349 | 0.0345 | 0.0348 | -0.05 | 0.0351 | 14 | 0.0062 | 0 |
| CBN (%AR) | 005.20 | Other LC-MS | 190 | 0.037 | 0.038 |  | 0.0375 | 0.39 | 0.0351 | 14 | 0.0062 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 162 | 0.0393 | 0.0395 | 0.0392 | 0.0393 | 0.68 | 0.0351 | 14 | 0.0062 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 125 | 0.0446 | 0.0386 | 0.0441 | 0.0424 | 1.18 | 0.0351 | 14 | 0.0062 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 138 | 0.0422 | 0.0448 | 0.0415 | 0.0428 | 1.25 | 0.0351 | 14 | 0.0062 | 0 |
| CBN (%AR) | 005.99 | Other | 183 | 0.0437 | 0.0437 | 0.0443 | 0.0439 | 1.42 | 0.0351 | 14 | 0.0062 | 0 |
|  |  |  |  | Total Δ9-THC (%AR) |  |  |  |  |  |  |  |  |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 210 | <0.0870 | <0.0910 | <0.0900 | <0.0870 |  |  |  |  | 1 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 192 | 0.0033 | 0.0031 | 0.0048 | 0.0037 | -6.02 | 0.0707 | 16 | 0.0111 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 197 | 0.058 | 0.0521 | 0.0561 | 0.0554 | -1.37 | 0.0707 | 16 | 0.0111 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 103 | 0.058 | 0.057 | 0.055 | 0.0567 | -1.26 | 0.0707 | 16 | 0.0111 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 107 | 0.0779 | 0.0616 | 0.0591 | 0.0662 | -0.4 | 0.0707 | 16 | 0.0111 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 138 | 0.0677 | 0.066 | 0.0667 | 0.0668 | -0.35 | 0.0707 | 16 | 0.0111 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 144 | 0.0694 | 0.0638 | 0.0668 | 0.0667 | -0.36 | 0.0707 | 16 | 0.0111 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 190 | 0.068 | 0.067 |  | 0.0675 | -0.28 | 0.0707 | 16 | 0.0111 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 196 | 0.0702 | 0.0674 | 0.0692 | 0.0689 | -0.16 | 0.0707 | 16 | 0.0111 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 169 | 0.0674 | 0.07 | 0.0721 | 0.0698 | -0.08 | 0.0707 | 16 | 0.0111 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 102 | 0.076 | 0.065 | 0.08 | 0.0737 | 0.27 | 0.0707 | 16 | 0.0111 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 199 | 0.0775 | 0.0785 | 0.0765 | 0.0775 | 0.61 | 0.0707 | 16 | 0.0111 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 125 | 0.0795 | 0.0783 | 0.0783 | 0.0787 | 0.72 | 0.0707 | 16 | 0.0111 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 162 | 0.0726 | 0.0869 | 0.0817 | 0.0804 | 0.88 | 0.0707 | 16 | 0.0111 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 189 | 0.0818 | 0.0838 | 0.0778 | 0.0811 | 0.94 | 0.0707 | 16 | 0.0111 | 0 |
| Total Δ9-THC (%AR) | 006.03 | AOAC 2018.11, M | 139 | 0.0891 | 0.0802 | 0.0754 | 0.0816 | 0.98 | 0.0707 | 16 | 0.0111 | 0 |
| Total Δ9-THC (%AR) | 006.99 | Other | 183 | 0.0858 | 0.0854 | 0.086 | 0.0857 | 1.35 | 0.0707 | 16 | 0.0111 | 0 |
|  |  |  |  | Total CBD (%AR) |  |  |  |  |  |  |  |  |
| Total CBD (%AR) | 007.10 | Other LC-UV | 192 | 0.1971 | 0.2038 | 0.1971 | 0.199 | -7.33 | 3.978 | 17 | 0.5153 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 190 | 3.356 | 3.444 |  | 3.4 | -1.12 | 3.978 | 17 | 0.5153 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 215 | 3.4454 | 3.4541 | 3.5176 | 3.47 | -0.98 | 3.978 | 17 | 0.5153 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 189 | 3.5387 | 3.5337 | 3.5437 | 3.54 | -0.85 | 3.978 | 17 | 0.5153 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 107 | 3.1515 | 3.7001 | 3.9641 | 3.61 | -0.72 | 3.978 | 17 | 0.5153 | 0 |
| Total CBD (%AR) | 007.03 | AOAC 2018.11, M | 139 | 3.6581 | 3.615 | 3.8357 | 3.7 | -0.53 | 3.978 | 17 | 0.5153 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 196 | 3.8922 | 3.8103 | 3.9194 | 3.87 | -0.2 | 3.978 | 17 | 0.5153 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 162 | 4.0117 | 3.9202 | 3.9423 | 3.96 | -0.04 | 3.978 | 17 | 0.5153 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 199 | 4.0297 | 4.0295 | 4.0605 | 4.04 | 0.12 | 3.978 | 17 | 0.5153 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 103 | 4.1288 | 4.1308 | 3.9199 | 4.06 | 0.16 | 3.978 | 17 | 0.5153 | 0 |
| Total CBD (%AR) | 007.01 | AOAC 2018.10 | 210 | 4.087 | 4.054 | 4.076 | 4.07 | 0.18 | 3.978 | 17 | 0.5153 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 144 | 4.016 | 4.1025 | 4.2486 | 4.12 | 0.28 | 3.978 | 17 | 0.5153 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 125 | 4.24 | 4.26 | 4.21 | 4.24 | 0.5 | 3.978 | 17 | 0.5153 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 197 | 4.2921 | 4.3234 | 4.4531 | 4.36 | 0.73 | 3.978 | 17 | 0.5153 | 0 |
| Total CBD (%AR) | 007.99 | Other | 183 | 4.48 | 4.48 | 4.48 | 4.48 | 0.97 | 3.978 | 17 | 0.5153 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%AR) | 007.40 | GC-FID | 169 | 4.726 | 4.7474 | 4.849 | 4.77 | 1.55 | 3.978 | 17 | 0.5153 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 138 | 4.9075 | 4.9177 | 4.8977 | 4.91 | 1.8 | 3.978 | 17 | 0.5153 | 0 |
|  |  |  |  | CBDV (%AR) |  |  |  |  |  |  |  |  |
| CBDV (%AR) | 008.99 | Other | 197 | <0.0101 | <0.0101 | <0.0090 | <0.0101 |  |  |  |  | 1 |
| CBDV (%AR) | 008.99 | Other | 196 | <0.0210 | <0.0210 | <0.0210 | <0.0210 |  |  |  |  | 1 |
| CBDV (%AR) | 008.02 | AOAC 2018.11, U | 210 | <0.0210 | <0.0230 | <0.0220 | <0.0210 |  |  |  |  | 1 |
| CBDV (%AR) | 008.99 | Other | 189 | 0.0137 | 0.0134 | 0.0136 | 0.0136 |  |  |  |  | 0 |
| CBDV (%AR) | 008.99 | Other | 162 | 0.0232 | 0.0239 | 0.0173 | 0.0215 |  |  |  |  | 0 |
| CBDV (%AR) | 008.99 | Other | 106 | 0.03 | 0.03 | 0.03 | 0.03 |  |  |  |  | 0 |
| CBDV (%AR) | 008.99 | Other | 183 | 0.0348 | 0.0388 | 0.037 | 0.0369 |  |  |  |  | 0 |
| CBDV (%AR) | 008.02 | AOAC 2018.11, U | 199 | 0.0555 | 0.0565 | 0.0555 | 0.0558 |  |  |  |  | 0 |
|  |  |  |  | CBDVA (%AR) |  |  |  |  |  |  |  |  |
| CBDVA (%AR) | 009.99 | Other | 183 | <0.0021 | <0.0021 | <0.0021 | <0.0021 |  |  |  |  | 1 |
| CBDVA (%AR) | 009.99 | Other | 197 | <0.0101 | <0.0101 | <0.0090 | <0.0101 |  |  |  |  | 1 |
| CBDVA (%AR) | 009.99 | Other | 196 | <0.0210 | <0.0210 | <0.0210 | <0.0210 |  |  |  |  | 1 |
| CBDVA (%AR) | 009.01 | AOAC 2018.10 | 210 | 0.023 | 0.025 | 0.023 | 0.0237 |  |  |  |  | 0 |
|  |  |  |  | CBG (%AR) |  |  |  |  |  |  |  |  |
| CBG (%AR) | 010.99 | Other | 197 | 0.0542 | 0.0564 | 0.0534 | 0.0547 | -1.41 | 0.0768 | 10 | 0.0157 | 0 |
| CBG (%AR) | 010.99 | Other | 106 | 0.06 | 0.06 | 0.07 | 0.0633 | -0.86 | 0.0768 | 10 | 0.0157 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%AR) | 010.01 | AOAC 2018.10 | 210 | 0.071 | 0.064 | 0.063 | 0.066 | -0.69 | 0.0768 | 10 | 0.0157 | 0 |
| CBG (%AR) | 010.99 | Other | 196 | 0.0669 | 0.0666 | 0.0687 | 0.0674 | -0.6 | 0.0768 | 10 | 0.0157 | 0 |
| CBG (%AR) | 010.99 | Other | 189 | 0.0766 | 0.0789 | 0.0792 | 0.0782 | 0.09 | 0.0768 | 10 | 0.0157 | 0 |
| CBG (%AR) | 010.02 | AOAC 2018.11, U | 199 | 0.081 | 0.0795 | 0.081 | 0.0805 | 0.23 | 0.0768 | 10 | 0.0157 | 0 |
| CBG (%AR) | 010.02 | AOAC 2018.11, U | 103 | 0.084 | 0.085 | 0.078 | 0.0823 | 0.35 | 0.0768 | 10 | 0.0157 | 0 |
| CBG (%AR) | 010.99 | Other | 162 | 0.0884 | 0.0856 | 0.0872 | 0.0871 | 0.65 | 0.0768 | 10 | 0.0157 | 0 |
| CBG (%AR) | 010.02 | AOAC 2018.11, U | 125 | 0.0882 | 0.0897 | 0.0875 | 0.0885 | 0.74 | 0.0768 | 10 | 0.0157 | 0 |
| CBG (%AR) | 010.99 | Other | 183 | 0.104 | 0.107 | 0.104 | 0.105 | 1.8 | 0.0768 | 10 | 0.0157 | 0 |

Population of Lab Values

| CBGA (%AR) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%AR) | 011.99 | Other | 183 | <0.0049 | <0.0049 | <0.0049 | <0.0049 | 1 |
| CBGA (%AR) | 011.99 | Other | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| CBGA (%AR) | 011.99 | Other | 197 | <0.0101 | <0.0101 | <0.0090 | <0.0101 | 1 |
| CBGA (%AR) | 011.99 | Other | 196 | <0.0210 | <0.0210 | <0.0210 | <0.0210 | 1 |
| CBGA (%AR) | 011.01 | AOAC 2018.10 | 210 | <0.0230 | <0.0250 | <0.0230 | <0.0230 | 1 |
| CBGA (%AR) | 011.99 | Other | 106 | <0.0300 | <0.0300 | <0.0300 | <0.0300 | 1 |
| CBGA (%AR) | 011.02 | AOAC 2018.11, U | 125 | 0.0133 | 0.0137 | 0.0148 | 0.0139 | 0 |
| CBGA (%AR) | 011.99 | Other | 189 | 0.0152 | 0.0151 | 0.0149 | 0.0151 | 0 |
| CBGA (%AR) | 011.02 | AOAC 2018.11, U | 103 | 0.03 | 0.03 | 0.029 | 0.0297 | 0 |

THCV (%AR)

## THCV (%AR)

012.99	Other

183

<0.0069  <0.0069	<0.0069	<0.0069	1

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score Rob Mean |  | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| THCV (%AR) | 012.99 | Other | 197 | <0.0101 | <0.0101 | <0.0090 | <0.0101 |  |  |  |  | 1 |
| THCV (%AR) | 012.02 | AOAC 2018.11, U | 210 | <0.0210 | <0.0230 | <0.0220 | <0.0210 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 196 | <0.0210 | <0.0210 | <0.0210 | <0.0210 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 106 | <0.0300 | <0.0300 | <0.0300 | <0.0300 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 189 | 0.0091 | 0.0093 | 0.0094 | 0.0093 |  |  |  |  | 0 |

Population of Lab Values

| Δ8-THC (%AR) |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%AR) | 013.99 | Other | 196 | 0.0246 | 0.0235 | 0.0241 | 0.0241 | -1.46 | 0.0338 | 10 | 0.0066 | 0 |
| Δ8-THC (%AR) | 013.02 | AOAC 2018.11, U | 199 | 0.0275 | 0.0255 | 0.0255 | 0.0262 | -1.14 | 0.0338 | 10 | 0.0066 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 162 | 0.0301 | 0.0321 | 0.0318 | 0.0313 | -0.37 | 0.0338 | 10 | 0.0066 | 0 |
| Δ8-THC (%AR) | 013.03 | AOAC 2018.11, M | 139 | 0.0288 | 0.0322 | 0.0366 | 0.0325 | -0.19 | 0.0338 | 10 | 0.0066 | 0 |
| Δ8-THC (%AR) | 013.01 | AOAC 2018.10 | 210 | 0.029 | 0.035 | 0.037 | 0.0337 | -0.01 | 0.0338 | 10 | 0.0066 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 183 | 0.0342 | 0.0356 | 0.0334 | 0.0344 | 0.1 | 0.0338 | 10 | 0.0066 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 197 | 0.0372 | 0.0352 | 0.032 | 0.0348 | 0.16 | 0.0338 | 10 | 0.0066 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 189 | 0.0377 | 0.0362 | 0.0369 | 0.0369 | 0.48 | 0.0338 | 10 | 0.0066 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 106 | 0.04 | 0.04 | 0.04 | 0.04 | 0.94 | 0.0338 | 10 | 0.0066 | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 138 | 0.0814 | 0.082 | 0.0822 | 0.0819 | 7.24 | 0.0338 | 10 | 0.0066 | 0 |

|  |  |  | CBC (%AR) |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%AR) | 014.99 | Other | 106 | 0.15 | 0.15 | 0.16 | 0.153 | -1.41 | 0.1832 | 8 | 0.0213 | 0 |
| CBC (%AR) | 014.99 | Other | 197 | 0.1792 | 0.1731 | 0.1636 | 0.172 | -0.53 | 0.1832 | 8 | 0.0213 | 0 |
| CBC (%AR) | 014.99 | Other | 196 | 0.1757 | 0.1715 | 0.1756 | 0.174 | -0.42 | 0.1832 | 8 | 0.0213 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%AR) | 014.99 | Other | 189 | 0.1777 | 0.1772 | 0.1774 | 0.177 | -0.27 | 0.1832 | 8 | 0.0213 | 0 |
| CBC (%AR) | 014.99 | Other | 162 | 0.1799 | 0.1943 | 0.1727 | 0.182 | -0.04 | 0.1832 | 8 | 0.0213 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, U | 199 | 0.193 | 0.1945 | 0.1955 | 0.194 | 0.52 | 0.1832 | 8 | 0.0213 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, U | 125 | 0.2 | 0.202 | 0.19 | 0.197 | 0.66 | 0.1832 | 8 | 0.0213 | 0 |
| CBC (%AR) | 014.99 | Other | 183 | 0.214 | 0.213 | 0.218 | 0.215 | 1.49 | 0.1832 | 8 | 0.0213 | 0 |

Population of Lab Values

|  |  |  | Cd (ug/g AR) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cd (ug/g AR) | 101.04 | ICP-MS | 183 | <0.0010 | <0.0010 | <0.0010 | <0.0010 | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 102 | <0.0032 | <0.0032 | <0.0032 | <0.0032 | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 196 | <0.0241 | <0.0241 | <0.0241 | <0.0241 | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 106 | <0.0800 | <0.0800 | <0.0800 | <0.0800 | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 197 | <0.0984 | <0.0998 | <0.0986 | <0.0984 | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 | <0.2000 | 1 |

|  |  |  | As (ug/g AR) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| As (ug/g AR) | 102.04 | ICP-MS | 183 | <0.0020 | <0.0020 | <0.0020 | <0.0020 | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 102 | <0.0032 | <0.0032 | <0.0032 | <0.0032 | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 196 | <0.0482 | <0.0482 | <0.0482 | <0.0482 | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 106 | <0.0800 | <0.0800 | <0.0800 | <0.0800 | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 197 | <0.0984 | <0.0998 | <0.0986 | <0.0984 | 1 |
| As (ug/g AR) | 102.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 | <0.2000 | 1 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Hg (ug/g AR) |  |  |  |  |  |  |  |  |
| Hg (ug/g AR) | 103.04 | ICP-MS | 102 | <0.0032 | <0.0032 | <0.0032 | <0.0032 |  |  |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 196 | <0.0072 | <0.0072 | <0.0072 | <0.0072 |  |  |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 106 | <0.0080 | <0.0080 | <0.0080 | <0.0080 |  |  |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 183 | <0.0120 | <0.0120 | <0.0120 | <0.0120 |  |  |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 197 | <0.0984 | <0.0998 | <0.0986 | <0.0984 |  |  |  |  | 1 |
| Hg (ug/g AR) | 103.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 | <0.2000 |  |  |  |  | 1 |
|  |  |  |  | Pb (ug/g AR) |  |  |  |  |  |  |  |  |
| Pb (ug/g AR) | 104.04 | ICP-MS | 183 | <0.0200 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  | 1 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 196 | <0.0241 | <0.0241 | <0.0241 | <0.0241 |  |  |  |  | 1 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 106 | <0.0800 | <0.0800 | <0.0800 | <0.0800 |  |  |  |  | 1 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 197 | <0.0984 | <0.0998 | <0.0986 | <0.0984 |  |  |  |  | 1 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 | <0.2000 |  |  |  |  | 1 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 102 | 0.0069 | 0.0065 | 0.0059 | 0.0064 |  |  |  |  | 0 |
|  |  |  |  | Cu (ug/g AR) |  |  |  |  |  |  |  |  |
| Cu (ug/g AR) | 105.04 | ICP-MS | 196 | <0.4893 <0.4893 <0.4893 |  |  | <0.4893 |  |  |  |  | 1 |
|  |  |  |  | Ni (ug/g AR) |  |  |  |  |  |  |  |  |
| Ni (ug/g AR) | 106.04 | ICP-MS | 196 | <0.0722 <0.0722 <0.0722 |  |  | <0.0722 |  |  |  |  | 1 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Cr (ug/g AR) |  |  |  |  |  |  |  |  |  |
| Cr (ug/g AR) | 107.04 | ICP-MS | 196 | <0.0722 <0.0722 <0.0722 |  |  | <0.0722 | 1 |  |  |  |  |
| Cr (ug/g AR) | 107.04 | ICP-MS | 197 | 10.0000 10.0000 <10.0000 |  |  | <10.0000 | 1 |  |  |  |  |

Population of Lab Values

|  |  |  | Se (ug/g AR) |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Se (ug/g AR) | 108.04 | ICP-MS | 196 | <0.0963 <0.0963 <0.0963 | <0.0963 | 1 |
| Se (ug/g AR) | 108.04 | ICP-MS | 197 | <0.0984 <0.0998 <0.0986 | <0.0984 | 1 |

Zn (ug/g AR)

Zn (ug/g AR)

## 111.04	ICP-MS

196

<5.0000  <5.0000	<5.0000	<5.0000	1

Sb (ug/g AR)

Sb (ug/g AR)

## 112.04	ICP-MS

196

<0.0241  <0.0241	<0.0241	<0.0241	1

|  |  |  | (-)-alpha-Bisabolol (%AR) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 196 | <0.1235 | <0.1235 | <0.1235 | <0.1235 | 1 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 162 | 0.0122 | 0.0106 | 0.0102 | 0.011 | 0 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 183 | 0.19 | 0.191 | 0.191 | 0.191 | 0 |

|  |  |  | beta-Caryophyllene (%AR) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| beta-Caryophyllene (% | 202.02 | GC,MS | 196 | <0.1235 <0.1235 <0.1235 |  |  | <0.1235 | 1 |
| beta-Caryophyllene (% | 202.02 | GC,MS | 162 | 0.0443 | 0.0416 | 0.0392 | 0.0417 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| beta-Caryophyllene (% | 202.02 | GC,MS | 183 | 0.131 | 0.128 | 0.137 | 0.132 |  |  |  |  | 0 |
|  |  |  | (-)-Caryophyllene oxide (%AR) |  |  |  |  |  |  |  |  |  |
| (-)-Caryophyllene oxide | 203.02 | GC,MS | 196 | <0.1235 <0.1235 <0.1235 |  |  | <0.1235 |  |  |  |  | 1 |
| (-)-Caryophyllene oxide | 203.02 | GC,MS | 183 | 0.0605 | 0.0642 | 0.0636 | 0.0628 |  |  |  |  | 0 |
|  |  |  |  | Geraniol (%AR) |  |  |  |  |  |  |  |  |
| Geraniol (%AR) | 204.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 |  |  |  |  | 1 |
| Geraniol (%AR) | 204.02 | GC,MS | 183 | <0.0020 | <0.0020 | <0.0020 | <0.0020 |  |  |  |  | 1 |
| Geraniol (%AR) | 204.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
|  |  |  |  | alpha-Humulene (%AR) |  |  |  |  |  |  |  |  |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 196 | <0.1235 | <0.1235 | <0.1235 | <0.1235 |  |  |  |  | 1 |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 162 | 0.0125 | 0.0112 | 0.0101 | 0.0113 |  |  |  |  | 0 |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 183 | 0.0458 | 0.0469 | 0.0475 | 0.0467 |  |  |  |  | 0 |
|  |  |  |  | d-Limonene (%AR) |  |  |  |  |  |  |  |  |
| d-Limonene (%AR) | 206.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 |  |  |  |  | 1 |
| d-Limonene (%AR) | 206.02 | GC,MS | 183 | <0.0020 | <0.0020 | <0.0020 | <0.0020 |  |  |  |  | 1 |
| d-Limonene (%AR) | 206.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Linalool (%AR) |  |  |  |  |  |  |  |  |
| Linalool (%AR) | 207.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 |  |  |  |  | 1 |
| Linalool (%AR) | 207.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Linalool (%AR) | 207.02 | GC,MS | 183 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
|  |  |  |  | beta-Myrcene (%AR) |  |  |  |  |  |  |  |  |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 |  |  |  |  | 1 |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 183 | <0.0020 | <0.0020 | <0.0020 | <0.0020 |  |  |  |  | 1 |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
|  |  |  |  | Ocimene (%AR) |  |  |  |  |  |  |  |  |
| Ocimene (%AR) | 210.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 |  |  |  |  | 1 |
| Ocimene (%AR) | 210.02 | GC,MS | 183 | <0.0020 | <0.0020 | <0.0020 | <0.0020 |  |  |  |  | 1 |
| Ocimene (%AR) | 210.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
|  |  |  |  | alpha-Pinene (%AR) |  |  |  |  |  |  |  |  |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 |  |  |  |  | 1 |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 183 | <0.0020 | <0.0020 | <0.0020 | <0.0020 |  |  |  |  | 1 |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
|  |  |  |  | (-)-beta-Pinene (%AR) |  |  |  |  |  |  |  |  |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 196 | <0.0013 <0.0013 <0.0013 |  |  | <0.0013 |  |  |  |  | 1 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 183 | <0.0020 | <0.0020 | <0.0020 | <0.0020 |  |  |  |  | 1 |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |

Population of Lab Values

|  |  |  | alpha-Terpinene (%AR) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 | 1 |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 183 | <0.0020 | <0.0020 | <0.0020 | <0.0020 | 1 |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |

|  |  |  | gamma-Terpinene (%AR) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| gamma-Terpinene (%A | 214.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 | 1 |
| gamma-Terpinene (%A | 214.02 | GC,MS | 183 | <0.0020 | <0.0020 | <0.0020 | <0.0020 | 1 |
| gamma-Terpinene (%A | 214.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |

|  |  |  | alpha-Terpinolene (%AR) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 196 | <0.0013 | <0.0013 | <0.0013 | <0.0013 | 1 |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 183 | <0.0020 | <0.0020 | <0.0020 | <0.0020 | 1 |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |

- Flag definitions: 1 (Lab Value not used because there were less than 2 numeric lab results), 2 (Lab Value not used because there were less than 2 numeric lab results greater than 0).

- Robust statistics performed to determine robust mean and Z scores only if there were 6 or more numeric lab values.

- AR refers to concentration on as-received basis. DW refers to concentration on dry weight basis.