| Homogeneity and Stability |  |
| --- | --- |
| Sample ID . . . . . . . . . . . . . . . . . . . . . . . . . | OL23NOV-1 |
| Samples shipped . . . . . . . . . . . . . . . . . . . | 6-Nov-23 |
| Results due date . . . . . . . . . . . . . . . . . . . . | 15-Dec-23 |

| Sample # | Replicate 1 | Replicate 2 |
| --- | --- | --- |
| 1 | 0.0826 | 0.0878 |
| 2 | 0.0993 | 0.0867 |
| 3 | 0.0914 | 0.0858 |
| 4 | 0.0862 | 0.0865 |
| 5 | 0.0854 | 0.0882 |
| 6 | 0.0878 | 0.0879 |
| 7 | 0.0883 | 0.0848 |
| 8 | 0.0868 | 0.0862 |
| 9 | 0.0876 | 0.0880 |
| 10 | 0.0894 | 0.0863 |

| Sample # | Replicate 1 | Replicate 2 |
| --- | --- | --- |
| 1 | 4.6044 | 4.6564 |
| 2 | 4.6018 | 4.6844 |
| 3 | 4.8508 | 4.6359 |
| 4 | 4.6453 | 4.6880 |
| 5 | 4.6225 | 4.6716 |
| 6 | 4.6522 | 4.6261 |
| 7 | 4.6103 | 4.6587 |
| 8 | 4.6410 | 4.6638 |
| 9 | 4.6047 | 4.6090 |
| 10 | 4.7409 | 4.6428 |

# Homogeneity

10 sample vials randomly chosen for analysis

| Sample # | Replicate 1 | Replicate 2 |
| --- | --- | --- |
| 1 | 0.0783 | 0.0774 |
| 2 | 0.0774 | 0.0767 |
| 3 | 0.0797 | 0.0771 |
| 4 | 0.0755 | 0.0772 |
| 5 | 0.0755 | 0.0779 |
| 6 | 0.0771 | 0.0757 |
| 7 | 0.0770 | 0.0760 |
| 8 | 0.0785 | 0.0782 |
| 9 | 0.0774 | 0.0752 |
| 10 | 0.0807 | 0.0775 |

Duplicate test portions analyzed for analyte and method shown in table caption

| Total delta-9 THC (%AR) via GC-FID (method 006.40) |  |  |  | %RSD | HorRat-r |
| --- | --- | --- | --- | --- | --- |
| Analysis Date: 11/14/23 |  | Overall average: | 0.0773 |  |  |
|  |  | SD of sample avgs: | 0.001023 | 1.32 |  |
|  |  | repeatability SD: | 0.001326 | 1.72 | 0.3 |
|  |  | between-sample SD: | 0.00041 | 0.53 |  |
|  |  | reproducibilty SD: | 0.001388 | 1.80 |  |
|  |  | SD used for Analyte z scores: | 0.0111 |  |  |
|  |  | Check value: | 0.00333 |  |  |
| Total delta-9 THC (%AR) via GC-FID (method 006.40) Is between-sample SD less than check value? YES Homogeneity test passed |  |  |  |  |  |
| delta-8 THC (%AR) via GC-FID (method 0013.99) |  |  |  | %RSD | HorRat-r |
| Analysis Date: 11/14/23 |  | Overall average: | 0.0877 |  |  |
|  |  | SD of sample avgs: | 0.002137 | 2.44 |  |
|  |  | repeatability SD: | 0.003519 | 4.01 | 0.7 |
|  |  | between-sample SD: | 0 | 0.00 |  |
|  |  | reproducibilty SD: | 0.003519 | 4.01 |  |
|  |  | SD used for Analyte z scores: | 0.0066 |  |  |
|  |  | Check value: | 0.00198 |  |  |
| delta-8 THC (%AR) via GC-FID (method 0013.99) Is between-sample SD less than check value? YES Homogeneity test passed |  |  |  |  |  |
| Total CBD (%AR) via GC-FID (method 007.40) |  |  |  | %RSD | HorRat-r |
| Analysis Date: 11/14/23 |  | Overall average: | 4.6555 |  |  |
|  |  | SD of sample avgs: | 0.03817 | 0.82 |  |
|  |  | repeatability SD: | 0.06046 | 1.30 | 0.4 |
|  |  | between-sample SD: | 0 | 0.00 |  |
|  |  | reproducibilty SD: | 0.06046 | 1.30 |  |
|  |  | SD used for Analyte z scores: | 0.5153 |  |  |
|  |  | Check value: | 0.15459 |  |  |
| Total CBD (%AR) via GC-FID (method 007.40) Is between-sample SD less than check value? YES Homogeneity test passed |  |  |  |  |  |

# Stability

Sample tested on different days from date sample shipped to results due date Sample stored at room temperature during period of testing

| Total THC (%AR) via GC-FID (method 006.40) |  |  |
| --- | --- | --- |
| Date of Analysis | Days | Conc. |
| 11/10/23 | 0 | 0.0705 |
| 11/14/23 | 4 | 0.0762 |
| 11/29/23 | 19 | 0.0664 |
| 12/4/23 | 24 | 0.0770 |
| 12/7/23 | 27 | 0.0645 |
| 12/12/23 | 32 | 0.0763 |

Days = x variable, Conc = y variable SUMMARY OUTPUT

Regression Statistics

Multiple R	0.0425034 R Square	 0.0018065 Adjusted R Squ -0.2477418 Standard Error  0.0061342

Observations	6

ANOVA

0.08

0.07

0.06

0.05

y = -2E-05x + 0.0721

R² = 0.0018

0.04

0.03

0.02

0

10

20

30

40

df	SS	MS	F	Significance F

Regression	1  2.7241E-07 2.72405E-07 0.007239  0.936283254

Residual	4 0.00015052	3.7629E-05

Total	5 0.00015079

Coefficients Standard Erro	t Stat	P-value	Lower 95%	Upper 95%

Intercept	0.0721361 0.00451274  15.9849915  8.95E-05	0.05960671	0.08466545

X Variable 1	-1.808E-05	0.0002125 -0.08508374 0.936283  -0.000608065  0.000571905

Coefficient for x variable is not statistically significantly different from 0. No signficant change in result during testing period.

Stability test passed.

%RSD = 7.6	HorRat-R = 1.3

| delta-8 THC (%AR) via GC-FID (method 006.40) |  |  |
| --- | --- | --- |
| Date of Analysis | Days | Conc. |
| 11/10/23 | 0 | 0.0836 |
| 11/14/23 | 4 | 0.0087 |
| 11/29/23 | 19 | 0.0858 |
| 12/4/23 | 24 | 0.0939 |
| 12/7/23 | 27 | 0.0848 |
| 12/12/23 | 32 | 0.0898 |

Days = x variable, Conc = y variable SUMMARY OUTPUT

Regression Statistics

Multiple R	0.5694605 R Square	 0.3242853 Adjusted R Squ 0.1553566 Standard Error  0.0297993

Observations	6

ANOVA

df	SS	MS	F	Significance F

Regression	1 0.00170466 0.001704655 1.919658  0.238143043

y = 0.0014x + 0.0492 R² = 0.3243

0.20

0.18

0.16

0.14

0.12

0.10

0.08

0.06

0.04

0.02

0.00

0

10

20

30

40

Residual	4	0.003552 0.000887999

Total	5 0.00525665

Coefficients Standard Erro	t Stat	P-value	Lower 95%	Upper 95%

| Intercept 0.0491658 0.02192226 2.242732648 0.088354 | -0.011700181 | 0.110031701 |
| --- | --- | --- |
| X Variable 1 0.0014302 0.00103228 1.385517256 0.238143 | -0.001435825 | 0.004296305 |

Coefficient for x variable is not statistically significantly different from 0. No signficant change in result during testing period.

Stability test passed.

%RSD = 43.6	HorRat-R = 7.4

| Total CBD (%AR) via GC-FID (method 006.40) |  |  |
| --- | --- | --- |
| Date of Analysis | Days | Conc. |
| 11/10/23 | 0 | 4.7833 |
| 11/14/23 | 4 | 4.6699 |
| 11/29/23 | 19 | 4.4257 |
| 12/4/23 | 24 | 4.3595 |
| 12/7/23 | 27 | 4.5114 |
| 12/12/23 | 32 | 4.3226 |

Days = x variable, Conc = y variable SUMMARY OUTPUT

Regression Statistics

Multiple R	0.9166348 R Square	 0.8402193 Adjusted R Squ 0.8002741 Standard Error  0.0812124

Observations	6

ANOVA

df	SS	MS	F	Significance F

Regression	1 0.13873091 0.138730906 21.03431  0.010134958

y = -0.0129x + 4.74 R² = 0.8402

5.0

4.5

4.0

3.5

3.0

2.5

2.0

0

10

20

30

40

Residual	4 0.02638183 0.006595457

Total	5 0.16511273

Coefficients Standard Erro	t Stat	P-value	Lower 95%	Upper 95%

Intercept	4.7400126 0.05974497 79.33743546  1.51E-07  4.574133974  4.905891226

X Variable 1	-0.0129026 0.00281328 -4.58631803 0.010135  -0.020713519  -0.005091681

Coefficient for x variable is statistically significantly different from 0. Signficant change in result occurred during testing period.

Stability test did not pass.

%RSD = 4	HorRat-R = 1.3