# Homogeneity and Stability

Sample ID . . . . . . . . . . . . . . . . . . . . . . . . .

Samples shipped . . . . . . . . . . . . . . . . . . .

Results due date . . . . . . . . . . . . . . . . . . . .

OL24NOV-1

11-Nov-24

20-Dec-24

## Homogeneity

Not deteremined

## Stability

Sample tested on different days from date sample shipped to results due date. Sample stored at room temperature during period of testing.

| Total THC (%AR) via GC-FID (method 006.40) |  |  |
| --- | --- | --- |
| Date of Analysis | Days | Conc. |
| 11/13/24 | 0 | 0.1573 |
| 12/2/24 | 19 | 0.1449 |
| 12/10/24 | 27 | 0.1783 |
| 12/16/24 | 33 | 0.1522 |
| 12/24/24 | 41 | 0.1578 |

Days = x variable, Conc = y variable SUMMARY OUTPUT

Regression Statistics

Multiple R	 0.1125797 R Square	0.0126742 Adjusted R Squ -0.3164344 Standard Error  0.0142594

Observations	5

ANOVA

0.20

0.15

0.10

y = 9E-05x + 0.156

R² = 0.0127

0.05

0

10

20

30

40

50

df	SS	MS	F	Significance F

Regression	1  7.8304E-06 7.83037E-06  0.038510658  0.856962441

Residual	3 0.00060999  0.00020333

Total	4 0.00061782

Coefficients tandard Error	t Stat	P-value	Lower 95%	Upper 95%

| Intercept | 0.1559547 0.01265599 12.32259681 | 0.001151231 | 0.115677677 | 0.196231711 |
| --- | --- | --- | --- | --- |
| X Variable 1 | 8.939E-05 0.0004555 0.196241326 | 0.856962441 | -0.001360214 | 0.001538989 |

Coefficient for x variable is not statistically significantly different from 0. No signficant change in result during testing period.

Stability test passed.

%RSD = 7.9	HorRat-R = 1.5

| delta-8 THC (%AR) via GC-FID (method 006.40) |  |  |
| --- | --- | --- |
| Date of Analysis | Days | Conc. |
| 11/13/24 | 0 | 0.4190 |
| 12/2/24 | 19 | 0.4283 |
| 12/10/24 | 27 | 0.4289 |
| 12/16/24 | 33 | 0.4382 |
| 12/24/24 | 41 | 0.4246 |

Days = x variable, Conc = y variable SUMMARY OUTPUT

Regression Statistics

Multiple R	 0.5712722 R Square	0.3263519 Adjusted R Squ 0.1018025 Standard Error  0.0066561

Observations	5

ANOVA

df	SS	MS	F	Significance F

Regression	1  6.4389E-05 6.43892E-05  1.453363527  0.314403336

y = 0.0003x + 0.4216 R² = 0.3264

0.50

0.48

0.46

0.44

0.42

0.40

0

10

20

30

40

50

Residual	3 0.00013291 4.43036E-05

Total	4	0.0001973

Coefficients tandard Error	t Stat	P-value	Lower 95%	Upper 95%

Intercept	0.4216482  0.00590765 71.37319395	6.0612E-06	0.40284737  0.440448956

X Variable 1	0.0002563 0.00021262 1.205555277  0.314403336  -0.000420329  0.000932982

Coefficient for x variable is not statistically significantly different from 0. No signficant change in result during testing period.

Stability test passed.

%RSD = 1.6	HorRat-R = 0.4

| Total CBD (%AR) via GC-FID (method 006.40) |  |  |
| --- | --- | --- |
| Date of Analysis | Days | Conc. |
| 11/13/24 | 0 | 2.3794 |
| 12/2/24 | 19 | 2.4317 |
| 12/10/24 | 27 | 2.3114 |
| 12/16/24 | 33 | 2.3828 |
| 12/24/24 | 41 | 2.5141 |

Days = x variable, Conc = y variable SUMMARY OUTPUT

Regression Statistics

Multiple R	 0.3948756 R Square	0.1559267 Adjusted R Squ -0.125431 Standard Error  0.0795976

Observations	5

ANOVA

df	SS	MS	F	Significance F

Regression	1 0.00351125  0.00351125  0.554193859  0.510618719

y = 0.0019x + 2.3585 R² = 0.1559

3.0

2.8

2.6

2.4

2.2

2.0

0

10

20

30

40

50

Residual	3 0.01900734 0.006335779

Total	4 0.02251859

Coefficients tandard Error	t Stat	P-value	Lower 95%	Upper 95%

Intercept	2.3584514 0.07064731 33.38345662	5.90848E-05  2.133620156  2.583282701

X Variable 1	0.0018929 0.00254265 0.744441978  0.510618719  -0.006198997  0.009984712

Coefficient for x variable is not statistically significantly different from 0. No signficant change in result during testing period.

Stability test passed.

%RSD = 3.1	HorRat-R = 0.9