Hemp Proficiency Testing

All Labs Trueness OL24SEP-1 (oil)

Sample: OL24SEP-1 (oil)

All Labs Trueness

Issue Date: November 11, 2024

| Analyte | Code | Method | Lab Num |  | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Δ9-THC (%AR) |  |  |  |  |  |  |  |  |  |  |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 |  | 228 | 0.0791 | 0.0802 | 0.0786 | 0.0793 | -4.61 | 0.1873 | 24 | 0.0234 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 213 | 0.0810 | 0.0760 | 0.0790 | 0.0787 | -4.64 | 0.1873 | 24 | 0.0234 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 149 | 0.1660 | 0.1640 | 0.1530 | 0.161 | -1.12 | 0.1873 | 24 | 0.0234 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS |  | 190 | 0.2140 | 0.1650 | 0.1320 | 0.17 | -0.72 | 0.1873 | 24 | 0.0234 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 106 | 0.1700 | 0.1700 | 0.1700 | 0.17 | -0.74 | 0.1873 | 24 | 0.0234 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U |  | 103 | 0.1756 | 0.1693 | 0.1667 | 0.171 | -0.71 | 0.1873 | 24 | 0.0234 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 197 | 0.1657 | 0.1778 | 0.1732 | 0.172 | -0.64 | 0.1873 | 24 | 0.0234 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS |  | 102 | 0.1690 | 0.1790 | 0.1750 | 0.174 | -0.55 | 0.1873 | 24 | 0.0234 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U |  | 125 | 0.1842 | 0.1721 | 0.1867 | 0.181 | -0.27 | 0.1873 | 24 | 0.0234 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 184 | 0.2120 | 0.1644 | 0.1709 | 0.182 | -0.21 | 0.1873 | 24 | 0.0234 | 0 |
| Δ9-THC (%AR) | 001.01 | AOAC 2018.10 |  | 210 | 0.1760 | 0.1860 | 0.1880 | 0.183 | -0.17 | 0.1873 | 24 | 0.0234 | 0 |
| Δ9-THC (%AR) | 001.99 | Other |  | 183 | 0.1899 | 0.1745 | 0.1862 | 0.184 | -0.16 | 0.1873 | 24 | 0.0234 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 142 | 0.1860 | 0.1860 | 0.1850 | 0.186 | -0.07 | 0.1873 | 24 | 0.0234 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV |  | 189 | 0.1917 | 0.1868 | 0.1895 | 0.189 | 0.09 | 0.1873 | 24 | 0.0234 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 196 | 0.1890 | 0.1828 | 0.1942 | 0.189 | 0.06 | 0.1873 | 24 | 0.0234 | 0 |
| Δ9-THC (%AR) | 001.30 | Other LC-MS | 149 | 0.1880 | 0.2000 | 0.1880 | 0.192 | 0.2 | 0.1873 | 24 | 0.0234 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 183 | 0.1947 | 0.1920 | 0.1931 | 0.193 | 0.26 | 0.1873 | 24 | 0.0234 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 173 | 0.2050 | 0.1980 | 0.1980 | 0.2 | 0.56 | 0.1873 | 24 | 0.0234 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 192 | 0.2102 | 0.2248 | 0.1783 | 0.204 | 0.73 | 0.1873 | 24 | 0.0234 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 163 | 0.2090 | 0.2100 | 0.2060 | 0.208 | 0.9 | 0.1873 | 24 | 0.0234 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 162 | 0.2081 | 0.2091 | 0.2193 | 0.212 | 1.06 | 0.1873 | 24 | 0.0234 | 0 |
| Δ9-THC (%AR) | 001.02 | AOAC 2018.11, U | 199 | 0.2265 | 0.2245 | 0.2280 | 0.226 | 1.67 | 0.1873 | 24 | 0.0234 | 0 |
| Δ9-THC (%AR) | 001.10 | Other LC-UV | 170 | 0.2440 | 0.2392 | 0.2401 | 0.241 | 2.3 | 0.1873 | 24 | 0.0234 | 0 |
| Δ9-THC (%AR) | 001.03 | AOAC 2018.11, M | 139 | 0.2661 | 0.2868 | 0.2968 | 0.283 | 4.1 | 0.1873 | 24 | 0.0234 | 0 |

Population of Lab Values

| Δ9-THCA (%AR) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 190 | <0.0003 | <0.0003 | <0.0003 | <0.0003 | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 183 | <0.0020 | <0.0020 | <0.0020 | <0.0020 | 1 |
| Δ9-THCA (%AR) | 002.99 | Other | 183 | <0.0020 | <0.0020 | <0.0020 | <0.0020 | 1 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 125 | <0.0020 | <0.0018 | <0.0018 | <0.0020 | 1 |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 149 | <0.0040 | <0.0040 | <0.0040 | <0.0040 | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 142 | <0.0040 | <0.0040 | <0.0040 | <0.0040 | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 196 | <0.0050 | <0.0050 | <0.0050 | <0.0050 | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 184 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| Δ9-THCA (%AR) | 002.03 | AOAC 2018.11, M | 139 | <0.0125 | <0.0125 | <0.0125 | <0.0125 | 1 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score Rob Mean |  | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THCA (%AR) | 002.30 | Other LC-MS | 102 | <0.0200 | <0.0200 | <0.0190 | <0.0200 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 163 | <0.0200 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 210 | <0.0240 | <0.0240 | <0.0240 | <0.0240 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 170 | <0.0300 | <0.0300 | <0.0300 | <0.0300 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 106 | <0.0300 | <0.0300 | <0.0300 | <0.0300 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 149 | <0.0400 | <0.0400 | <0.0400 | <0.0400 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.02 | AOAC 2018.11, U | 199 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 173 | 0.0000 | 0.0000 | 0.0000 | 0 |  |  |  |  | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 192 | 0.0000 | 0.0000 | 0.0043 | 0.0014 |  |  |  |  | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 189 | 0.0057 | 0.0031 | 0.0124 | 0.0071 |  |  |  |  | 0 |
| Δ9-THCA (%AR) | 002.10 | Other LC-UV | 197 | <0.0152 | <0.0150 | 0.0149 | 0.015 |  |  |  |  | 1 |
| Δ9-THCA (%AR) | 002.01 | AOAC 2018.10 | 228 | 0.0748 | 0.0757 | 0.0712 | 0.0739 |  |  |  |  | 0 |

Population of Lab Values

| CBD (%AR) |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%AR) | 003.01 | AOAC 2018.10 | 228 | 2.3601 | 2.3511 | 2.3625 | 2.36 | -8.85 | 7.0785 | 23 | 0.5334 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 149 | 6.2660 | 6.2250 | 6.1200 | 6.2 | -1.64 | 7.0785 | 23 | 0.5334 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 213 | 6.3750 | 6.0910 | 6.1590 | 6.21 | -1.63 | 7.0785 | 23 | 0.5334 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 170 | 6.6177 | 6.5285 | 6.5602 | 6.57 | -0.96 | 7.0785 | 23 | 0.5334 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 184 | 8.0260 | 5.2449 | 6.5825 | 6.62 | -0.86 | 7.0785 | 23 | 0.5334 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 190 | 6.6450 | 6.7340 | 6.6520 | 6.68 | -0.75 | 7.0785 | 23 | 0.5334 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 189 | 6.7264 | 6.7195 | 6.7363 | 6.73 | -0.66 | 7.0785 | 23 | 0.5334 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 103 | 7.0071 | 6.5985 | 7.0822 | 6.9 | -0.34 | 7.0785 | 23 | 0.5334 | 0 |

# Lab

# Lab

# Population of Lab Values

| Analyte | Code | Method | Num | Result1 | Result2 | Result3 | Value | Z score Rob Mean |  | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBD (%AR) | 003.10 | Other LC-UV | 173 | 6.9970 | 7.0440 | 6.9850 | 7.01 | -0.13 | 7.0785 | 23 | 0.5334 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 162 | 7.1044 | 6.9815 | 6.9581 | 7.01 | -0.12 | 7.0785 | 23 | 0.5334 | 0 |
| CBD (%AR) | 003.30 | Other LC-MS | 149 | 7.3170 | 7.6610 | 6.5130 | 7.16 | 0.16 | 7.0785 | 23 | 0.5334 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 125 | 6.7970 | 7.1249 | 7.5444 | 7.16 | 0.14 | 7.0785 | 23 | 0.5334 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 196 | 7.1661 | 7.2943 | 7.2344 | 7.23 | 0.29 | 7.0785 | 23 | 0.5334 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 183 | 7.2522 | 7.1677 | 7.2605 | 7.23 | 0.28 | 7.0785 | 23 | 0.5334 | 0 |
| CBD (%AR) | 003.01 | AOAC 2018.10 | 210 | 7.3160 | 7.3290 | 7.2170 | 7.29 | 0.39 | 7.0785 | 23 | 0.5334 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 106 | 7.4200 | 7.2700 | 7.2600 | 7.32 | 0.45 | 7.0785 | 23 | 0.5334 | 0 |
| CBD (%AR) | 003.99 | Other | 183 | 7.3792 | 7.2603 | 7.3259 | 7.32 | 0.46 | 7.0785 | 23 | 0.5334 | 0 |
| CBD (%AR) | 003.03 | AOAC 2018.11, M | 139 | 7.2748 | 7.4401 | 7.4452 | 7.39 | 0.58 | 7.0785 | 23 | 0.5334 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 197 | 7.5669 | 7.5898 | 7.5089 | 7.56 | 0.89 | 7.0785 | 23 | 0.5334 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 192 | 7.7542 | 8.1019 | 6.9697 | 7.61 | 0.99 | 7.0785 | 23 | 0.5334 | 0 |
| CBD (%AR) | 003.02 | AOAC 2018.11, U | 199 | 7.6165 | 7.6910 | 7.6165 | 7.64 | 1.06 | 7.0785 | 23 | 0.5334 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 142 | 7.6600 | 7.7000 | 7.7000 | 7.69 | 1.14 | 7.0785 | 23 | 0.5334 | 0 |
| CBD (%AR) | 003.10 | Other LC-UV | 163 | 8.9200 | 8.9700 | 9.1700 | 9.02 | 3.64 | 7.0785 | 23 | 0.5334 | 0 |

| CBDA (%AR) |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDA (%AR) | 004.10 | Other LC-UV | 170 | <0.0300 | <0.0300 | <0.0300 | <0.0300 |  |  |  |  | 1 |
| CBDA (%AR) | 004.10 | Other LC-UV | 149 | <0.0400 | <0.0400 | <0.0400 | <0.0400 |  |  |  |  | 1 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 199 | <0.0500 | <0.0500 | <0.0500 | <0.0500 |  |  |  |  | 1 |
| CBDA (%AR) | 004.10 | Other LC-UV | 173 | <0.1410 | <0.1530 | <0.1560 | <0.1410 |  |  |  |  | 1 |
| CBDA (%AR) | 004.10 | Other LC-UV | 192 | 0.0000 | 0.0000 | 0.0000 | 0 | -3.76 | 0.0335 | 17 | 0.0089 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score Rob Mean |  | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDA (%AR) | 004.10 | Other LC-UV | 196 | 0.0258 | 0.0248 | 0.0257 | 0.0254 | -0.91 | 0.0335 | 17 | 0.0089 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 149 | 0.0250 | 0.0260 | 0.0240 | 0.025 | -0.96 | 0.0335 | 17 | 0.0089 | 0 |
| CBDA (%AR) | 004.30 | Other LC-MS | 190 | 0.0350 | 0.0280 | 0.0210 | 0.028 | -0.62 | 0.0335 | 17 | 0.0089 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 197 | 0.0268 | 0.0288 | 0.0272 | 0.0276 | -0.67 | 0.0335 | 17 | 0.0089 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 106 | 0.0300 | 0.0300 | 0.0300 | 0.03 | -0.4 | 0.0335 | 17 | 0.0089 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 189 | 0.0369 | 0.0297 | 0.0279 | 0.0315 | -0.23 | 0.0335 | 17 | 0.0089 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 162 | 0.0335 | 0.0286 | 0.0308 | 0.031 | -0.29 | 0.0335 | 17 | 0.0089 | 0 |
| CBDA (%AR) | 004.01 | AOAC 2018.10 | 228 | 0.0322 | 0.0309 | 0.0341 | 0.0324 | -0.13 | 0.0335 | 17 | 0.0089 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 142 | 0.0323 | 0.0315 | 0.0313 | 0.0317 | -0.21 | 0.0335 | 17 | 0.0089 | 0 |
| CBDA (%AR) | 004.01 | AOAC 2018.10 | 210 | 0.0330 | 0.0350 | 0.0350 | 0.0343 | 0.09 | 0.0335 | 17 | 0.0089 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 125 | 0.0327 | 0.0324 | 0.0393 | 0.0348 | 0.14 | 0.0335 | 17 | 0.0089 | 0 |
| CBDA (%AR) | 004.02 | AOAC 2018.11, U | 103 | 0.0369 | 0.0373 | 0.0383 | 0.0375 | 0.44 | 0.0335 | 17 | 0.0089 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 183 | 0.0425 | 0.0412 | 0.0422 | 0.042 | 0.94 | 0.0335 | 17 | 0.0089 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 163 | 0.0450 | 0.0460 | 0.0440 | 0.045 | 1.28 | 0.0335 | 17 | 0.0089 | 0 |
| CBDA (%AR) | 004.03 | AOAC 2018.11, M | 139 | 0.0566 | 0.0442 | 0.0489 | 0.0499 | 1.83 | 0.0335 | 17 | 0.0089 | 0 |
| CBDA (%AR) | 004.10 | Other LC-UV | 184 | 0.0520 | 0.0514 | 0.0501 | 0.0512 | 1.98 | 0.0335 | 17 | 0.0089 | 0 |

Population of Lab Values

| CBN (%AR) |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%AR) | 005.30 | GC-FID | 163 | <0.0200 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 149 | <0.0400 | <0.0400 | <0.0400 | <0.0400 |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 173 | <0.1410 | <0.1530 | <0.1560 | <0.1410 |  |  |  |  | 1 |
| CBN (%AR) | 005.20 | Other LC-MS | 190 | 0.0380 | 0.0280 | 0.0280 | 0.0313 | -1.47 | 0.0406 | 15 | 0.0063 | 0 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBN (%AR) | 005.10 | Other LC-UV | 213 | 0.0340 | 0.0340 | 0.0360 | 0.0347 | -0.94 | 0.0406 | 15 | 0.0063 | 0 |
| CBN (%AR) | 005.20 | Other LC-MS | 149 | 0.0360 | 0.0370 | 0.0350 | 0.036 | -0.73 | 0.0406 | 15 | 0.0063 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 196 | 0.0376 | 0.0361 | 0.0376 | 0.0371 | -0.56 | 0.0406 | 15 | 0.0063 | 0 |
| CBN (%AR) | 005.03 | AOAC 2018.11, M | 139 | 0.0372 | 0.0377 | 0.0387 | 0.0379 | -0.43 | 0.0406 | 15 | 0.0063 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 189 | 0.0382 | 0.0385 | 0.0381 | 0.0383 | -0.37 | 0.0406 | 15 | 0.0063 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 197 | 0.0386 | 0.0392 | 0.0392 | 0.039 | -0.25 | 0.0406 | 15 | 0.0063 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 106 | 0.0400 | 0.0400 | 0.0400 | 0.04 | -0.1 | 0.0406 | 15 | 0.0063 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 162 | 0.0371 | 0.0407 | 0.0415 | 0.0398 | -0.13 | 0.0406 | 15 | 0.0063 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 142 | 0.0400 | 0.0406 | 0.0403 | 0.0403 | -0.05 | 0.0406 | 15 | 0.0063 | 0 |
| CBN (%AR) | 005.99 | Other | 183 | 0.0428 | 0.0394 | 0.0417 | 0.0413 | 0.11 | 0.0406 | 15 | 0.0063 | 0 |
| CBN (%AR) | 005.10 | Other LC-UV | 183 | 0.0445 | 0.0441 | 0.0449 | 0.0445 | 0.62 | 0.0406 | 15 | 0.0063 | 0 |
| CBN (%AR) | 005.30 | GC-FID | 138 | 0.0416 | 0.0577 | 0.0493 | 0.0495 | 1.42 | 0.0406 | 15 | 0.0063 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 125 | 0.0474 | 0.0486 | 0.0520 | 0.0493 | 1.39 | 0.0406 | 15 | 0.0063 | 0 |
| CBN (%AR) | 005.02 | AOAC 2018.11, U | 199 | <0.0500 | 0.0500 | <0.0500 | 0.05 |  |  |  |  | 1 |
| CBN (%AR) | 005.10 | Other LC-UV | 192 | 0.0554 | 0.0604 | 0.0476 | 0.0545 | 2.2 | 0.0406 | 15 | 0.0063 | 0 |
|  |  |  |  | Total Δ9-THC (%AR) |  |  |  |  |  |  |  |  |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 163 | <0.0200 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  | 1 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 149 | 0.1660 | 0.1640 | 0.1530 | 0.161 | -1.72 | 0.1905 | 21 | 0.0172 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 190 | 0.2140 | 0.1650 | 0.1320 | 0.17 | -1.17 | 0.1905 | 21 | 0.0172 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 103 | 0.1756 | 0.1693 | 0.1667 | 0.171 | -1.16 | 0.1905 | 21 | 0.0172 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 102 | 0.1690 | 0.1790 | 0.1750 | 0.174 | -0.94 | 0.1905 | 21 | 0.0172 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 197 | 0.1657 | 0.1778 | 0.1863 | 0.177 | -0.81 | 0.1905 | 21 | 0.0172 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 125 | 0.1842 | 0.1721 | 0.1867 | 0.181 | -0.55 | 0.1905 | 21 | 0.0172 | 0 |
| Total Δ9-THC (%AR) | 006.40 | GC-FID | 138 | 0.1813 | 0.1776 | 0.1881 | 0.182 | -0.48 | 0.1905 | 21 | 0.0172 | 0 |
| Total Δ9-THC (%AR) | 006.99 | Other | 183 | 0.1899 | 0.1745 | 0.1862 | 0.184 | -0.41 | 0.1905 | 21 | 0.0172 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 142 | 0.1860 | 0.1860 | 0.1850 | 0.186 | -0.28 | 0.1905 | 21 | 0.0172 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 107 | 0.1740 | 0.1920 | 0.1970 | 0.188 | -0.17 | 0.1905 | 21 | 0.0172 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 196 | 0.1890 | 0.1828 | 0.1942 | 0.189 | -0.11 | 0.1905 | 21 | 0.0172 | 0 |
| Total Δ9-THC (%AR) | 006.30 | Other LC-MS | 149 | 0.1880 | 0.2000 | 0.1880 | 0.192 | 0.09 | 0.1905 | 21 | 0.0172 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 183 | 0.1947 | 0.1920 | 0.1931 | 0.193 | 0.16 | 0.1905 | 21 | 0.0172 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 189 | 0.1967 | 0.1894 | 0.2004 | 0.196 | 0.29 | 0.1905 | 21 | 0.0172 | 0 |
| Total Δ9-THC (%AR) | 006.50 | GC-MS | 170 | 0.1965 | 0.2058 | 0.1970 | 0.2 | 0.54 | 0.1905 | 21 | 0.0172 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 173 | 0.2050 | 0.1980 | 0.1980 | 0.2 | 0.57 | 0.1905 | 21 | 0.0172 | 0 |
| Total Δ9-THC (%AR) | 006.01 | AOAC 2018.10 | 210 | 0.1970 | 0.2070 | 0.2090 | 0.204 | 0.8 | 0.1905 | 21 | 0.0172 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 192 | 0.2102 | 0.2248 | 0.1821 | 0.206 | 0.88 | 0.1905 | 21 | 0.0172 | 0 |
| Total Δ9-THC (%AR) | 006.10 | Other LC-UV | 162 | 0.2080 | 0.2090 | 0.2190 | 0.212 | 1.25 | 0.1905 | 21 | 0.0172 | 0 |
| Total Δ9-THC (%AR) | 006.02 | AOAC 2018.11, U | 199 | 0.2265 | 0.2245 | 0.2280 | 0.226 | 2.08 | 0.1905 | 21 | 0.0172 | 0 |
| Total Δ9-THC (%AR) | 006.03 | AOAC 2018.11, M | 139 | 0.2661 | 0.2868 | 0.2968 | 0.283 | 5.39 | 0.1905 | 21 | 0.0172 | 0 |
|  |  |  |  | Total CBD (%AR) |  |  |  |  |  |  |  |  |
| Total CBD (%AR) | 007.10 | Other LC-UV | 149 | 6.2660 | 6.2250 | 6.1200 | 6.2 | -2.39 | 7.2801 | 19 | 0.4502 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 190 | 6.6757 | 6.7586 | 6.6704 | 6.7 | -1.29 | 7.2801 | 19 | 0.4502 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 189 | 6.7587 | 6.7455 | 6.7608 | 6.76 | -1.17 | 7.2801 | 19 | 0.4502 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 103 | 7.0395 | 6.6312 | 7.1158 | 6.93 | -0.78 | 7.2801 | 19 | 0.4502 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 162 | 7.1340 | 7.0070 | 6.9850 | 7.04 | -0.53 | 7.2801 | 19 | 0.4502 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 163 | 6.9500 | 7.1400 | 7.0900 | 7.06 | -0.49 | 7.2801 | 19 | 0.4502 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 173 | 7.0496 | 7.1019 | 7.0446 | 7.07 | -0.48 | 7.2801 | 19 | 0.4502 | 0 |
| Total CBD (%AR) | 007.30 | Other LC-MS | 149 | 7.3390 | 7.6840 | 6.5340 | 7.19 | -0.21 | 7.2801 | 19 | 0.4502 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 125 | 6.8256 | 7.1533 | 7.5788 | 7.19 | -0.21 | 7.2801 | 19 | 0.4502 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 196 | 7.1887 | 7.3160 | 7.2570 | 7.25 | -0.06 | 7.2801 | 19 | 0.4502 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 183 | 7.2895 | 7.2038 | 7.2976 | 7.26 | -0.04 | 7.2801 | 19 | 0.4502 | 0 |
| Total CBD (%AR) | 007.01 | AOAC 2018.10 | 210 | 7.4890 | 7.3600 | 7.2480 | 7.37 | 0.19 | 7.2801 | 19 | 0.4502 | 0 |
| Total CBD (%AR) | 007.03 | AOAC 2018.11, M | 139 | 7.3244 | 7.4789 | 7.4881 | 7.43 | 0.33 | 7.2801 | 19 | 0.4502 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 197 | 7.5904 | 7.6151 | 7.5327 | 7.58 | 0.66 | 7.2801 | 19 | 0.4502 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 192 | 7.7542 | 8.1019 | 6.9697 | 7.61 | 0.73 | 7.2801 | 19 | 0.4502 | 0 |
| Total CBD (%AR) | 007.02 | AOAC 2018.11, U | 199 | 7.6384 | 7.7129 | 7.6384 | 7.66 | 0.85 | 7.2801 | 19 | 0.4502 | 0 |
| Total CBD (%AR) | 007.10 | Other LC-UV | 142 | 7.6900 | 7.7300 | 7.7300 | 7.72 | 0.97 | 7.2801 | 19 | 0.4502 | 0 |
| Total CBD (%AR) | 007.50 | GC-MS | 170 | 7.9093 | 8.8064 | 8.4193 | 8.38 | 2.44 | 7.2801 | 19 | 0.4502 | 0 |
| Total CBD (%AR) | 007.40 | GC-FID | 138 | 8.5155 | 8.8089 | 8.8845 | 8.74 | 3.23 | 7.2801 | 19 | 0.4502 | 0 |
|  |  |  |  | CBDV (%AR) |  |  |  |  |  |  |  |  |
| CBDV (%AR) | 008.99 | Other | 173 | <0.1410 | <0.1530 | <0.1560 | <0.1410 |  |  |  |  | 1 |
| CBDV (%AR) | 008.99 | Other | 197 | 0.0437 | 0.0416 | 0.0459 | 0.0437 | -2.55 | 0.0682 | 10 | 0.0096 | 0 |
| CBDV (%AR) | 008.99 | Other | 196 | 0.0611 | 0.0594 | 0.0611 | 0.0605 | -0.8 | 0.0682 | 10 | 0.0096 | 0 |
| CBDV (%AR) | 008.99 | Other | 183 | 0.0643 | 0.0615 | 0.0647 | 0.0635 | -0.49 | 0.0682 | 10 | 0.0096 | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDV (%AR) | 008.99 | Other | 149 | 0.0630 | 0.0660 | 0.0640 | 0.0643 | -0.4 | 0.0682 | 10 | 0.0096 | 0 |
| CBDV (%AR) | 008.02 | AOAC 2018.11, U | 210 | 0.0640 | 0.0700 | 0.0660 | 0.0667 | -0.16 | 0.0682 | 10 | 0.0096 | 0 |
| CBDV (%AR) | 008.99 | Other | 189 | 0.0692 | 0.0654 | 0.0709 | 0.0685 | 0.03 | 0.0682 | 10 | 0.0096 | 0 |
| CBDV (%AR) | 008.99 | Other | 162 | 0.0793 | 0.0744 | 0.0602 | 0.0713 | 0.32 | 0.0682 | 10 | 0.0096 | 0 |
| CBDV (%AR) | 008.02 | AOAC 2018.11, U | 183 | 0.0724 | 0.0717 | 0.0709 | 0.0717 | 0.36 | 0.0682 | 10 | 0.0096 | 0 |
| CBDV (%AR) | 008.99 | Other | 106 | 0.0800 | 0.0800 | 0.0800 | 0.08 | 1.23 | 0.0682 | 10 | 0.0096 | 0 |
| CBDV (%AR) | 008.99 | Other | 142 | 0.0823 | 0.0820 | 0.0804 | 0.0816 | 1.39 | 0.0682 | 10 | 0.0096 | 0 |

Population of Lab Values

| CBDVA (%AR) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDVA (%AR) | 009.02 | AOAC 2018.11, U | 183 | <0.0010 | <0.0010 | <0.0010 | <0.0010 | 1 |
| CBDVA (%AR) | 009.99 | Other | 142 | <0.0040 | <0.0040 | <0.0040 | <0.0040 | 1 |
| CBDVA (%AR) | 009.99 | Other | 196 | <0.0050 | <0.0050 | <0.0050 | <0.0050 | 1 |
| CBDVA (%AR) | 009.99 | Other | 197 | <0.0102 | <0.0100 | <0.0980 | <0.0102 | 1 |
| CBDVA (%AR) | 009.01 | AOAC 2018.10 | 210 | <0.0240 | <0.0240 | <0.0240 | <0.0240 | 1 |
| CBDVA (%AR) | 009.99 | Other | 173 | 0.0000 | 0.0000 | 0.0000 | 0 | 0 |

|  |  |  | CBG (%AR) |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%AR) | 010.99 | Other | 173 | <0.1410 | <0.1530 | <0.1560 | <0.1410 |  |  |  |  | 1 |
| CBG (%AR) | 010.99 | Other | 149 | 0.0700 | 0.0720 | 0.0700 | 0.0707 | -1.37 | 0.0885 | 11 | 0.013 | 0 |
| CBG (%AR) | 010.99 | Other | 162 | 0.0756 | 0.0712 | 0.0770 | 0.0746 | -1.07 | 0.0885 | 11 | 0.013 | 0 |
| CBG (%AR) | 010.99 | Other | 196 | 0.0810 | 0.0775 | 0.0781 | 0.0789 | -0.74 | 0.0885 | 11 | 0.013 | 0 |
| CBG (%AR) | 010.99 | Other | 106 | 0.0800 | 0.0800 | 0.0800 | 0.08 | -0.65 | 0.0885 | 11 | 0.013 | 0 |

| Analyte Code |  | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBG (%AR) | 010.99 | Other | 142 | 0.0851 | 0.0849 | 0.0844 | 0.0848 | -0.29 | 0.0885 | 11 | 0.013 | 0 |
| CBG (%AR) | 010.99 | Other | 183 | 0.0933 | 0.0905 | 0.0899 | 0.0912 | 0.21 | 0.0885 | 11 | 0.013 | 0 |
| CBG (%AR) | 010.01 | AOAC 2018.10 | 210 | 0.0950 | 0.0930 | 0.0910 | 0.093 | 0.35 | 0.0885 | 11 | 0.013 | 0 |
| CBG (%AR) | 010.02 | AOAC 2018.11, U | 183 | 0.0960 | 0.0948 | 0.0956 | 0.0955 | 0.54 | 0.0885 | 11 | 0.013 | 0 |
| CBG (%AR) | 010.02 | AOAC 2018.11, U | 125 | 0.0920 | 0.1021 | 0.0999 | 0.098 | 0.73 | 0.0885 | 11 | 0.013 | 0 |
| CBG (%AR) | 010.99 | Other | 197 | 0.0857 | 0.1170 | 0.1022 | 0.102 | 1.01 | 0.0885 | 11 | 0.013 | 0 |
| CBG (%AR) | 010.99 | Other | 189 | 0.1095 | 0.1052 | 0.1014 | 0.105 | 1.3 | 0.0885 | 11 | 0.013 | 0 |

Population of Lab Values

| CBGA (%AR) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBGA (%AR) | 011.02 | AOAC 2018.11, U | 183 | <0.0010 | <0.0010 | <0.0010 | <0.0010 | 1 |
| CBGA (%AR) | 011.02 | AOAC 2018.11, U | 125 | <0.0020 | <0.0018 | <0.0018 | <0.0020 | 1 |
| CBGA (%AR) | 011.99 | Other | 142 | <0.0040 | <0.0040 | <0.0040 | <0.0040 | 1 |
| CBGA (%AR) | 011.99 | Other | 196 | <0.0050 | <0.0050 | <0.0050 | <0.0050 | 1 |
| CBGA (%AR) | 011.99 | Other | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| CBGA (%AR) | 011.99 | Other | 197 | <0.0102 | <0.0100 | <0.0980 | <0.0102 | 1 |
| CBGA (%AR) | 011.01 | AOAC 2018.10 | 210 | <0.0240 | <0.0240 | <0.0240 | <0.0240 | 1 |
| CBGA (%AR) | 011.99 | Other | 106 | <0.0300 | <0.0300 | <0.0300 | <0.0300 | 1 |
| CBGA (%AR) | 011.99 | Other | 173 | 0.0000 | 0.0000 | 0.0000 | 0 | 0 |
| CBGA (%AR) | 011.99 | Other | 189 | 0.0119 | 0.0122 | 0.0105 | 0.0115 | 0 |

THCV (%AR)

THCV (%AR)

012.02	AOAC 2018.11, U

# 183

<0.0020  <0.0020	<0.0020	<0.0020	1

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| THCV (%AR) | 012.99 | Other | 142 | <0.0040 | <0.0040 | <0.0040 | <0.0040 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 196 | <0.0050 | <0.0050 | <0.0050 | <0.0050 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 197 | <0.0102 | <0.0100 | <0.0980 | <0.0102 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 183 | <0.0200 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  | 1 |
| THCV (%AR) | 012.02 | AOAC 2018.11, U | 210 | <0.0230 | <0.0230 | <0.0230 | <0.0230 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 106 | <0.0300 | <0.0300 | <0.0300 | <0.0300 |  |  |  |  | 1 |
| THCV (%AR) | 012.99 | Other | 173 | 0.0000 | 0.0000 | 0.0000 | 0 |  |  |  |  | 0 |
| THCV (%AR) | 012.99 | Other | 189 | 0.0340 | 0.0319 | 0.0267 | 0.0309 |  |  |  |  | 0 |
|  |  |  |  | Δ8-THC (%AR) |  |  |  |  |  |  |  |  |
| Δ8-THC (%AR) | 013.02 | AOAC 2018.11, U | 183 | <0.0030 | <0.0030 | <0.0030 | <0.0030 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 142 | <0.0040 | <0.0040 | <0.0040 | <0.0040 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 196 | <0.0050 | <0.0050 | <0.0050 | <0.0050 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.03 | AOAC 2018.11, M | 139 | <0.0125 | <0.0125 | <0.0125 | <0.0125 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 183 | <0.0200 | <0.0200 | <0.0200 | <0.0200 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.01 | AOAC 2018.10 | 210 | <0.0230 | <0.0230 | <0.0230 | <0.0230 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 106 | <0.0300 | <0.0300 | <0.0300 | <0.0300 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 173 | 0.0000 | 0.0000 | 0.0000 | 0 |  |  |  |  | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 197 | <0.0152 | 0.0156 | <0.0144 | 0.0151 |  |  |  |  | 1 |
| Δ8-THC (%AR) | 013.99 | Other | 189 | 0.0194 | 0.0178 | 0.0185 | 0.0186 |  |  |  |  | 0 |
| Δ8-THC (%AR) | 013.99 | Other | 138 | 0.0599 | 0.0607 | 0.0562 | 0.0589 |  |  |  |  | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score Rob Mean |  | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Δ8-THC (%AR) | 013.01 | AOAC 2018.10 | 228 | 0.1084 | 0.1097 | 0.0997 | 0.106 |  |  |  |  | 0 |

Population of Lab Values

| CBC (%AR) |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBC (%AR) | 014.99 | Other | 149 | 0.2900 | 0.2920 | 0.2850 | 0.289 | -2.57 | 0.3309 | 11 | 0.0163 | 0 |
| CBC (%AR) | 014.99 | Other | 196 | 0.3204 | 0.3103 | 0.3169 | 0.316 | -0.92 | 0.3309 | 11 | 0.0163 | 0 |
| CBC (%AR) | 014.99 | Other | 106 | 0.3200 | 0.3200 | 0.3200 | 0.32 | -0.67 | 0.3309 | 11 | 0.0163 | 0 |
| CBC (%AR) | 014.99 | Other | 142 | 0.3220 | 0.3240 | 0.3220 | 0.323 | -0.5 | 0.3309 | 11 | 0.0163 | 0 |
| CBC (%AR) | 014.99 | Other | 197 | 0.3281 | 0.3240 | 0.3225 | 0.325 | -0.37 | 0.3309 | 11 | 0.0163 | 0 |
| CBC (%AR) | 014.99 | Other | 162 | 0.3378 | 0.3266 | 0.3393 | 0.335 | 0.22 | 0.3309 | 11 | 0.0163 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, U | 125 | 0.3273 | 0.3355 | 0.3431 | 0.335 | 0.27 | 0.3309 | 11 | 0.0163 | 0 |
| CBC (%AR) | 014.99 | Other | 173 | 0.3350 | 0.3370 | 0.3360 | 0.336 | 0.31 | 0.3309 | 11 | 0.0163 | 0 |
| CBC (%AR) | 014.99 | Other | 183 | 0.3452 | 0.3418 | 0.3447 | 0.344 | 0.8 | 0.3309 | 11 | 0.0163 | 0 |
| CBC (%AR) | 014.02 | AOAC 2018.11, U | 183 | 0.3477 | 0.3437 | 0.3434 | 0.345 | 0.86 | 0.3309 | 11 | 0.0163 | 0 |
| CBC (%AR) | 014.99 | Other | 189 | 0.3734 | 0.3710 | 0.3719 | 0.372 | 2.52 | 0.3309 | 11 | 0.0163 | 0 |

|  |  |  | Cd (ug/g AR) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cd (ug/g AR) | 101.04 | ICP-MS | 183 | <0.0010 | <0.0010 | <0.0010 | <0.0010 | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 102 | <0.0032 | <0.0032 | <0.0032 | <0.0032 | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 142 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 196 | <0.0482 | <0.0482 | <0.0482 | <0.0482 | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 106 | <0.0800 | <0.0800 | <0.0800 | <0.0800 | 1 |
| Cd (ug/g AR) | 101.04 | ICP-MS | 197 | <0.0980 | <0.0980 | <0.0982 | <0.0980 | 1 |

|  |  |  | Lab |  |  |  | Lab |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Analyte | Code | Method | Num | Result1 Result2 Result3 |  |  | Value | Z score |  | Rob Mean | n | Rob StDev | Flag |
| Cd (ug/g AR) | 101.04 | ICP-MS | 162 | <0.2000 <0.2000 <0.2000 |  |  | <0.2000 | 1 |  |  |  |  |  |
|  |  |  |  | As (ug/g AR) |  |  |  |  |  |  |  |  |  |
| As (ug/g AR) | 102.04 | ICP-MS | 183 | <0.0020 | <0.0020 | <0.0020 | <0.0020 | 1 |  |  |  |  |  |
| As (ug/g AR) | 102.04 | ICP-MS | 102 | <0.0032 | <0.0032 | <0.0032 | <0.0032 | 1 |  |  |  |  |  |
| As (ug/g AR) | 102.04 | ICP-MS | 106 | <0.0800 | <0.0800 | <0.0800 | <0.0800 | 1 |  |  |  |  |  |
| As (ug/g AR) | 102.04 | ICP-MS | 196 | <0.0963 | <0.0963 | <0.0963 | <0.0963 | 1 |  |  |  |  |  |
| As (ug/g AR) | 102.04 | ICP-MS | 197 | <0.0980 | <0.0980 | <0.0982 | <0.0980 | 1 |  |  |  |  |  |
| As (ug/g AR) | 102.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 | <0.2000 | 1 |  |  |  |  |  |
| As (ug/g AR) | 102.04 | ICP-MS | 142 | 0.0117 | 0.0102 | 0.0117 | 0.0112 | 0 |  |  |  |  |  |
|  |  |  |  | Hg (ug/g AR) |  |  |  |  |  |  |  |  |  |
| Hg (ug/g AR) | 103.04 | ICP-MS | 106 | <0.0080 | <0.0080 | <0.0080 | <0.0080 | 1 |  |  |  |  |  |
| Hg (ug/g AR) | 103.04 | ICP-MS | 142 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |  |  |  |  |  |
| Hg (ug/g AR) | 103.04 | ICP-MS | 183 | <0.0120 | <0.0120 | <0.0120 | <0.0120 | 1 |  |  |  |  |  |
| Hg (ug/g AR) | 103.04 | ICP-MS | 196 | <0.0145 | <0.0145 | <0.0145 | <0.0145 | 1 |  |  |  |  |  |
| Hg (ug/g AR) | 103.04 | ICP-MS | 197 | <0.0980 | <0.0980 | <0.0982 | <0.0980 | 1 |  |  |  |  |  |
| Hg (ug/g AR) | 103.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 | <0.2000 | 1 |  |  |  |  |  |
| Hg (ug/g AR) | 103.99 | Other | 102 | 0.0040 | 0.0025 | 0.0030 | 0.0032 | 0 |  |  |  |  |  |
|  |  |  |  | Pb (ug/g AR) |  |  |  |  |  |  |  |  |  |
| Pb (ug/g AR) | 104.04 | ICP-MS | 183 | <0.0020 <0.0020 <0.0020 |  |  | <0.0020 | 1 |  |  |  |  |  |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Pb (ug/g AR) | 104.04 | ICP-MS | 142 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 196 | <0.0482 | <0.0482 | <0.0482 | <0.0482 |  |  |  |  | 1 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 106 | <0.0800 | <0.0800 | <0.0800 | <0.0800 |  |  |  |  | 1 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 197 | <0.0980 | <0.0980 | <0.0982 | <0.0980 |  |  |  |  | 1 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 162 | <0.2000 | <0.2000 | <0.2000 | <0.2000 |  |  |  |  | 1 |
| Pb (ug/g AR) | 104.04 | ICP-MS | 102 | 0.0025 | 0.0043 | 0.0025 | 0.0031 |  |  |  |  | 0 |

Population of Lab Values

Cu (ug/g AR)

Cu (ug/g AR)

105.04	ICP-MS

# 196

<0.9790  <0.9790	<0.9790	<0.9790	1

Ni (ug/g AR)

Ni (ug/g AR)

Ni (ug/g AR)

106.04	ICP-MS

106.04	ICP-MS

# 196

# 183

<0.1445  <0.1445	<0.1445	<0.1445	1

0.0914	0.2610	0.1230	0.158	0

Cr (ug/g AR)

Cr (ug/g AR)

107.04	ICP-MS

# 196

<0.1448  <0.1448	<0.1448	<0.1448	1

|  |  |  | Se (ug/g AR) |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Se (ug/g AR) | 108.04 | ICP-MS | 197 | <0.0980 <0.0980 <0.0982 | <0.0980 | 1 |
| Se (ug/g AR) | 108.04 | ICP-MS | 196 | <0.1927 <0.1927 <0.1927 | <0.1927 | 1 |

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Zn (ug/g AR) |  |  |  |  |  |  |  |  |
| Zn (ug/g AR) | 111.04 | ICP-MS | 196 | 10.0000 10.0000 <10.0000 |  |  | <10.0000 |  |  |  |  | 1 |
|  |  |  |  | Sb (ug/g AR) |  |  |  |  |  |  |  |  |
| Sb (ug/g AR) | 112.04 | ICP-MS | 196 | <0.0482 <0.0482 <0.0482 |  |  | <0.0482 |  |  |  |  | 1 |
|  |  |  |  | (-)-alpha-Bisabolol (%AR) |  |  |  |  |  |  |  |  |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| (-)-alpha-Bisabolol (%A | 201.02 | GC,MS | 196 | 0.0750 | 0.0772 | 0.0727 | 0.075 |  |  |  |  | 0 |
| (-)-alpha-Bisabolol (%A | 201.01 | GC,FID | 142 | 0.0877 | 0.1110 | 0.1080 | 0.102 |  |  |  |  | 0 |
|  |  |  |  | beta-Caryophyllene (%AR) |  |  |  |  |  |  |  |  |
| beta-Caryophyllene (% | 202.02 | GC,MS | 162 | 0.0666 | 0.0602 | 0.0601 | 0.0623 |  |  |  |  | 0 |
| beta-Caryophyllene (% | 202.02 | GC,MS | 196 | 0.1313 | 0.1333 | 0.1277 | 0.131 |  |  |  |  | 0 |
| beta-Caryophyllene (% | 202.01 | GC,FID | 142 | 0.1430 | 0.1810 | 0.1750 | 0.166 |  |  |  |  | 0 |
|  |  |  | (-)-Caryophyllene oxide (%AR) |  |  |  |  |  |  |  |  |  |
| (-)-Caryophyllene oxide | 203.02 | GC,MS | 196 | 0.0332 | 0.0326 | 0.0313 | 0.0324 |  |  |  |  | 0 |
| (-)-Caryophyllene oxide | 203.01 | GC,FID | 142 | 0.0335 | 0.0414 | 0.0394 | 0.0381 |  |  |  |  | 0 |
|  |  |  |  | Geraniol (%AR) |  |  |  |  |  |  |  |  |
| Geraniol (%AR) | 204.01 | GC,FID | 142 | <0.0010 <0.0010 <0.0010 |  |  | <0.0010 |  |  |  |  | 1 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Geraniol (%AR) | 204.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| Geraniol (%AR) | 204.02 | GC,MS | 196 | <0.5040 | <0.5040 | <0.5040 | <0.5040 |  |  |  |  | 1 |
|  |  |  |  | alpha-Humulene (%AR) |  |  |  |  |  |  |  |  |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 162 | 0.0215 | 0.0189 | 0.0192 | 0.0199 |  |  |  |  | 0 |
| alpha-Humulene (%AR) | 205.02 | GC,MS | 196 | 0.0446 | 0.0464 | 0.0437 | 0.0449 |  |  |  |  | 0 |
| alpha-Humulene (%AR) | 205.01 | GC,FID | 142 | 0.0671 | 0.0884 | 0.0852 | 0.0802 |  |  |  |  | 0 |
|  |  |  |  | d-Limonene (%AR) |  |  |  |  |  |  |  |  |
| d-Limonene (%AR) | 206.02 | GC,MS | 196 | 0.9351 | 0.9533 | 0.9289 | 0.939 |  |  |  |  | 0 |
| d-Limonene (%AR) | 206.02 | GC,MS | 162 | 1.0713 | 1.1006 | 1.0400 | 1.07 |  |  |  |  | 0 |
| d-Limonene (%AR) | 206.01 | GC,FID | 142 | 1.0400 | 1.3000 | 1.2800 | 1.21 |  |  |  |  | 0 |
|  |  |  |  | Linalool (%AR) |  |  |  |  |  |  |  |  |
| Linalool (%AR) | 207.02 | GC,MS | 196 | 0.0133 | 0.0128 | 0.0143 | 0.0135 |  |  |  |  | 0 |
| Linalool (%AR) | 207.02 | GC,MS | 162 | 0.0164 | 0.0147 | 0.0143 | 0.0151 |  |  |  |  | 0 |
| Linalool (%AR) | 207.01 | GC,FID | 142 | 0.0160 | 0.0196 | 0.0190 | 0.0182 |  |  |  |  | 0 |
|  |  |  |  | beta-Myrcene (%AR) |  |  |  |  |  |  |  |  |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 196 | 0.0116 | 0.0126 | 0.0124 | 0.0122 |  |  |  |  | 0 |
| beta-Myrcene (%AR) | 208.02 | GC,MS | 162 | 0.0147 | 0.0144 | 0.0145 | 0.0145 |  |  |  |  | 0 |
| beta-Myrcene (%AR) | 208.01 | GC,FID | 142 | 0.0137 | 0.0169 | 0.0172 | 0.0159 |  |  |  |  | 0 |

Population of Lab Values

| Analyte | Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Nopinene (%AR) |  |  |  |  |  |  |  |  |
| Nopinene (%AR) | 209.01 | GC,FID | 142 | 0.0498 | 0.0618 | 0.0609 | 0.0575 |  |  |  |  | 0 |
|  |  |  |  | Ocimene (%AR) |  |  |  |  |  |  |  |  |
| Ocimene (%AR) | 210.01 | GC,FID | 142 | <0.0010 | <0.0010 | <0.0010 | <0.0010 |  |  |  |  | 1 |
| Ocimene (%AR) | 210.02 | GC,MS | 196 | <0.0078 | <0.0078 | <0.0078 | <0.0078 |  |  |  |  | 1 |
| Ocimene (%AR) | 210.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
|  |  |  |  | alpha-Pinene (%AR) |  |  |  |  |  |  |  |  |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |
| alpha-Pinene (%AR) | 211.02 | GC,MS | 196 | 0.0098 | 0.0101 | 0.0099 | 0.0099 |  |  |  |  | 0 |
| alpha-Pinene (%AR) | 211.01 | GC,FID | 142 | 0.0117 | 0.0140 | 0.0139 | 0.0132 |  |  |  |  | 0 |
|  |  |  |  | (-)-beta-Pinene (%AR) |  |  |  |  |  |  |  |  |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 196 | 0.0393 | 0.0399 | 0.0393 | 0.0395 |  |  |  |  | 0 |
| (-)-beta-Pinene (%AR) | 212.02 | GC,MS | 162 | 0.0509 | 0.0507 | 0.0505 | 0.0507 |  |  |  |  | 0 |
|  |  |  |  | alpha-Terpinene (%AR) |  |  |  |  |  |  |  |  |
| alpha-Terpinene (%AR) | 213.01 | GC,FID | 142 | <0.0010 | <0.0010 | <0.0010 | <0.0010 |  |  |  |  | 1 |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 196 | <0.0078 | <0.0078 | <0.0078 | <0.0078 |  |  |  |  | 1 |
| alpha-Terpinene (%AR) | 213.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 |  |  |  |  | 1 |

Population of Lab Values

| Analyte Code | Method | Lab Num | Result1 | Result2 | Result3 | Lab Value | Z score | Rob Mean | n | Rob StDev | Flag |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

Population of Lab Values

|  |  |  | gamma-Terpinene (%AR) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| gamma-Terpinene (%A | 214.02 | GC,MS | 162 | 0.0387 | 0.0372 | 0.0371 | 0.0377 | 0 |
| gamma-Terpinene (%A | 214.02 | GC,MS | 196 | 0.0371 | 0.0380 | 0.0376 | 0.0376 | 0 |

|  |  |  | alpha-Terpinolene (%AR) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 196 | <0.0078 | <0.0078 | <0.0078 | <0.0078 | 1 |
| alpha-Terpinolene (%A | 215.02 | GC,MS | 162 | <0.0100 | <0.0100 | <0.0100 | <0.0100 | 1 |
| alpha-Terpinolene (%A | 215.01 | GC,FID | 142 | 0.0075 | 0.0093 | 0.0092 | 0.0087 | 0 |

|  |  |  | Moisture (%AR) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Moisture (%AR) | 500.99 | Other | 142 | 4.5000 | 4.0000 | 4.7000 | 4.4 | 0 |

- Flag definitions: 1 (Lab Value not used because there were less than 2 numeric lab results), 2 (Lab Value not used because there were less than 2 numeric lab results greater than 0).

- Robust statistics performed to determine robust mean and Z scores only if there were 6 or more numeric lab values.

- AR refers to concentration on as-received basis. DW refers to concentration on dry weight basis.